# Executive Clinical Summary

P020012

Artecoll

Artes Medical, Inc.

TABLE OF CONTENTS

PART I

Product Description……………………………………………………………………….2

Proposed Intended Use……………………………………………………………………2

PART II

Artes Medical Clinical Study……………………………………………………….……..2

Artes Medical Study Conduct……………………………………………………………..7

Artes Medical Study Results…..…………………………………………………………10

Study Objectives versus Study Results ….………………………………………………15

PART III

Rofil Medical Clinical Study…………………………………………………………….17

Rofil Medical Study Conduct ..………………………………………………………….18

Rofil Medical Study Results……………………………………………………………..19

PART IV

Bovine Collagen Immunogenicity Study…………………………………………….…..20

Study Conduct .…………………………………………………………………………..23

Study Results……………………………………………………………………….……24

Study Objectives versus Study Results..…………………………………………………26

PART V …………………………………………………………………………………26

FFA Scale reference photographs………………..……………………Volume 3 p. 12-096

Attachment 2a  (Duration of Adverse Events – Artecoll treated patients)

Attachment 2b  (Duration of Adverse Events – Control group)

Attachment 2c (Duration of Adverse Events – Rofil Study)

Attachment 2d (Duration of Adverse Events – Bovine collagen immunogenicity study)

###### PART I

## Product Description

Artecoll is a suspension of polymethylmethacrylate (PMMA) microspheres in a 3.5% collagen solution.

No changes were made in the final product composition or formulation between or during the Rofil and Artes Medical studies that are described below.

#### Proposed Indications for Use

“Artecoll implant is indicated for the correction of contour deficiencies of soft tissue”.

PART II

## I.   	Artes Medical Clinical Study

  - Design

    - Prospective, randomized, double-blind, controlled clinical trial

    - Compares Artecoll to Zyplast® Implant for correction of:

      - Nasolabial Folds

      - Radial Upper Lip Lines

      - Depressed Mouth Corners

    - Compares Artecoll to Zyderm® Implant for correction of Glabellar Folds

  - Inclusion/Exclusion Criteria

| Inclusion | Exclusion |
| --- | --- |
| Subjects 18 years of age or older, of either sex Subjects having realistic expectations of the benefit and limitation of the augmentation procedure, as determined by a willingness to sign the informed consent form after it has been carefully explained. The informed consent form includes the statement that the subjects will not look younger after the treatment, and that treatment with Artecoll does not replace a facelift or eyelid surgery, nor does it replace the effect of laser treatment, chemical peel or dermabrasion on fine wrinkling. Subjects able and willing to give informed consent Subjects presenting for correction of any of the following four types of dermal contour deformities of the face: - Glabellar Folds (right and/or left side and/or center) Nasolabial Folds (right and/or left side) Radial Upper Lip Lines (right and/or left side) Depressed Corners of the Mouth (right and/or left side) Subjects willing and able to comply with the requirements of the study (e.g. study duration, number of visits, completion of questionnaires) | Subjects who were pregnant Subjects who had been treated with collagen, Botulinum toxin, or other wrinkle/fold therapies within the last six months at any intended implant area Subjects considering additional cosmetic treatments to the treatment area at any later time during the study Subjects who had received chemotherapy agents or corticosteroids within the last 3 months Subjects receiving UV light therapy Subjects who were planning to use substances which reduce coagulation, such as aspirin, non-steroidal anti-inflammatory drugs or Coumadin, within the 4 weeks preceding treatment (this was changed to 72 hours as of January 10th, 2000 per Protocol Revision 1) Subjects presenting with history of autoimmune disorder Subjects presenting with atrophic skin diseases at any intended treatment area Subjects with very thin and flaccid skin at any intended treatment area Subjects with known susceptibility to keloids Subjects with known lidocaine hypersensitivity Subjects with dietary beef allergy, or undergoing or planning to undergo desensitization to meat products Subjects with known allergy to collagen Subjects with severe allergies manifested by a history of anaphylaxis or history or presence of multiple severe allergies Subjects with cellulitis or infection at the treatment area Subjects demonstrating anti-bovine collagen serum IgG levels outside the normal range at baseline (after randomization) Subjects demonstrating a positive skin test or two equivocal skin tests (after randomization) |

  - Treatment Protocol

1.  Patient screened

- Patients randomized to receive either Artecoll or Zyplast/Zyderm treatment in one or more of the eight treatment areas (four for each side).  Subjects were masked regarding the randomization assignment.

- Blood was then drawn for a test of serum IgG level and the subject received the collagen skin test that correlated with the randomization assignment

- Subject underwent a 30-day observation period following the collagen skin test.

- In the case of equivocal test results, the subject underwent a second skin test with a subsequent 30-day observation period.

- Upon a negative skin test and a serum IgG level within normal range, the subject received treatment with Artecoll or Zyplast/Zyderm according to the randomization assignment.

    - Among subjects treated with the Control, Zyplast was used for Nasolabial folds, Upper lip lines and Mouth Corners

    - Zyderm was used for Glabellar Folds on the basis of labeling for Zyderm/Zyplast.

- Before treatment, a topical or deep dermal anesthetic was applied to areas selected for treatment as necessary.

- During each treatment session, the appropriate implant was injected deep dermally, immediately above the subcutaneous tissue, using a tunneling technique.

- To even distribution, the implant was spread and modeled with the fingertips.

- Post-procedure, subjects were encouraged to minimize facial movement for several days following the treatment session.

- The treatment course lasted up to one month and included a total of up to 3 treatment sessions.

  - Follow-up Protocol

1.  Follow-up intervals

- All subjects in both treatment groups were scheduled to follow-up at 1, 3, and 6 months post-treatment.

- Only the Artecoll treatment group patients were seen again at 12 months post-treatment to assess the occurrence of late adverse events

- Follow-up Assessments included

  - Adverse event assessment, documentation and additional treatments required

  - Photographic documentation using a standardized process

- Assessment Tools

| Assessment made by: | Assessment Tool | Details Regarding Assessment Tool |
| --- | --- | --- |
| Masked Observer | Facial Fold Assessment (FFA) Scale (Validated Scale) | A scale created and validated for this study for: Glabellar Folds Nasolabial Folds Radial Upper Lip Lines Marionette Lines And not for mouth corners. Each fold/line is graded on a 0 – 5 point scale based on reference photos where 0 is considered the least severe, optimal condition. _____________________________ Reference photographs used by the masked observers were the same size as the photographs being evaluated. (Reference photographs appear in Volume 3 p. 12-096) _____________________________ Although the scale was not validated for this, non-integral values for Masked Observer FFA scale ratings occurred under 2 circumstances: For any facial location at any time point ratings were averaged across raters and sides MO used ½ point steps for FFA ratings between 2 FFA scale anchor points Then the pre-treatment severity was determined by rounding the FFA scale score to the nearest integer. _____________________________ Interclass correlations for rater reliability of masked observer FFA ratings were acceptable being greater than 0.8 for each treatment area. |
| Unblinded Investigator | Rating of Treatment Success (Scale not validated) | 1) During each follow-up visit, the unblinded investigator rated the success of the implant at each treatment area using the following five point rating scale. These evaluations were based on seeing the patient in person. 1 = Completely Successful (Defined as Optimal Desired Outcome) 2 = Very Successful 3 = Moderately Successful 4 = Somewhat Successful 5 = Not at all Successful 2) During each follow-up visit, the unblinded investigator rated the success of the implant based on seeing the patient in person and using reduced size FFA reference photographs. |
| Blinded Patient | Subject satisfaction of treatment (Scale Not Validated) | During each follow-up visit, subject satisfaction with the result of each treatment area was assessed using a questionnaire at each follow-up period. Subjects rated their personal opinion of each treatment area using the following 5-point rating scale: 1 = Very Satisfied 2 = Satisfied 3 = Somewhat Satisfied 4 = Dissatisfied 5 = Very Dissatisfied |

  - Study Endpoints

      - Primary Endpoints

        - Objective 1:  The cosmetic correction will be assessed using the FFA Scale, related to the Fitzpatrick scale, but adapted to the deep folds that are the subject of this study, rather than the fine wrinkles for which the Fitzpatrick scale was developed.  A comparison will be made between Artecoll and Zyplast/Zyderm by measuring the correction that remains at the evaluation 6 months after completion of the treatment course.  The null hypothesis will be that there is no difference between Artecoll and Zyplast/Zyderm 6 months after completion of the treatment course.  A clinically significant difference will be defined as a difference of at least one FFA Scale classification.

        - Objective 2:  The assessment of safety will be based on the incidence of serious unanticipated adverse events.  Unanticipated adverse events will be characterized by the severity and relation to the implant product.

In addition, the occurrence of all anticipated adverse events will be assessed at each follow-up period and compared between treatment groups.

The level of serum IgG will be monitored pre-treatment and at 1 month post-treatment.  If the IgG level is elevated at the 1 month post-treatment visit, an additional sample will be taken at the 3 month visit.  If the 3 month sample is elevated, a final sample will be taken at the 6 month post-treatment visit.  The occurrence of levels falling outside the normal range will be compared between treatment groups.

#### Secondary Endpoints

- Objective 3:  The quality of the initial treatment result will be characterized using the FFA Scale and will be used to characterize the cosmetic correction during the early period up to 3 months post-treatment.  The percent of folds/lines that exhibit a decrease of at least one classification (improvement) between the pre-implant assessment and the 1 and 3 month post-treatment assessments will be calculated for each treatment group.

- Objective 4:   The physician’s assessment of success will be characterized using a non-parametric 5-point scale, measured at 1, 3, and 6 months post-treatment.

- Objective 5:  The patient’s assessment of satisfaction will be characterized using a non-parametric 5-point scale, measured at 1, 3, and 6 months post-treatment.

#### II.	Artes Medical Study Conduct

a.  Enrollment

1.  Included 285 subjects (141 Artecoll, 144 Control)

2. 111 subjects being available for 1-year safety evaluation

b.  Demographics

|  | Artecoll | Control | p-value |
| --- | --- | --- | --- |
| Gender |
| Male | 11 (8.6%) | 11 (8.9%) | 1.000 |
| Female | 117 (91.4%) | 112 (91.1%) |  |
| Mean Age (Std. dev.) | 53.2 (10.3) | 51.2 (11.3) | 0.157 |
| Ethnicity |
| Caucasian | 102(76.7%) | 107 (82.3%) | 0.277 |
| Hispanic | 24 (18.0%) | 21 (16.2%) |  |
| Asian | 1 (0.8%) | 1 (0.8%) |  |
| Other | 6 (4.5%) | 1 (0.8%) |  |
| Smoking Rate |
| None | 122 (91.0%) | 108 (83.1%) | 0.061 |
| Low | 8 (6.0%) | 18 (13.8%) |  |
| Medium | 4 (3.0%) | 3 (2.3%) |  |
| High | 0 (0.0%) | 1 (0.8%) |  |
| Sun Exposure |
| Low | 69 (51.9%) | 84 (64.6%) | 0.043 |
| Medium | 49 (36.8%) | 35 (26.9%) |  |
| High | 15 (11.3%) | 11 (8.5%) |  |
| Mean pretreatment masked observer FFA Scale Ratings (Std. dev.) |
| Glabellar Folds | 1.42 (0.90) | 1.40 (0.86) | 0.934 |
| Nasolabial Folds | 1.74 (1.06) | 1.45 (1.02) | 0.039 |
| Upper Lip Lines | 1.17 (0.82) | 1.18 (0.86) | 0.886 |
| Mouth Corners | 1.91 (1.15) | 1.87 (1.10) | 0.891 |

Using the Spearman Rank Order Correlation, the sponsor determined that the smoking rates were not associated with treatment effects and thus not a source of bias in the clinical trial.

Using the Spearman Rank Order Correlation, the sponsor found only one correlation significant for sun exposure.  In the Control group, Nasolabial fold improvement was found to be negatively correlated with sun exposure.  Lower sun exposure was associated with greater improvement in this group.

Please note that the mean pretreatment masked observer FFA scale ratings for nasolabial folds indicate that the fold severity is more severe for the Artecoll treatment group versus the control group severity.  This statistically significant difference between groups allows the Artecoll treatment group more room for improvement than the control.

|  | Artecoll | Control | p-value |
| --- | --- | --- | --- |
| Treated folds (treated subjects) |
| Glabellar Folds | 155 (81) | 165 (86) | 0.153 (0.327) |
| Nasolabial Folds | 214 (108) | 206 (104) | 0.939 (1.00) |
| Upper lip lines | 137 (69) | 116 (59) | 0.182 (0.415) |
| Mouth corners | 86 | 87 | 0.210 (0.638) |
| Number of facial fold areas treated |
| 1 area | 20 | 22 | 0.412 |
| 2 areas | 36 | 24 |  |
| 3 areas | 36 | 42 |  |
| 4 areas | 36 | 35 |  |
| Bilateral treatment to: |
| Glabellar Folds | 74 | 79 | 0.312 |
| Nasolabial Folds | 106 | 102 | 0.886 |
| Upper Lip Lines | 68 | 57 | 0.621 |
| Mouth Corners | 85 | 83 | 0.963 |
| Unilateral treatment to: |
| Glabellar Folds | 7 | 7 | 0.578 |
| Nasolabial Folds | 2 | 2 | 0.674 |
| Upper Lip Lines | 1 | 2 | 0.485 |
| Mouth Corners | 1 | 4 | 0.173 |

c. Accountability

| Patient Accounting - Artes medical study – Artecoll patients |
| --- |
|  | Preop | Intraop | 1 month | 3 months | 6 months | 12 months |
| Theoretical1 | 141 |
| Device removals, (cumulative) | 0 | 0 | 0 | 1* | 1* | 1* |
| Expected2 | 141 | 141 | 141 | 140 | 140 | 140 |
| Lost to follow-up (cumulative) | 0 | 0 | 2 | 4 | 6 | 10 |
| Aborted study (includes not treated and voluntary withdrawal) (cumulative) | 13 | 13 | 14 | 15 | 15 | 16 |
| Evaluated3 | 128 | 128 | 125 | 121 | 119 | 114 |
| % Expected Follow-up | 100% 141/141 | 100% 141/141 | 100% 141/141 | 99.3% 140/141 | 99.3% 140/141 | 99.3% 140/141 |
| Actual % Follow-up4 | 90% 128/141 | 90% 128/141 | 89% 125/141 | 86% 121/140 | 85% 119/140 | 81% 114/140 |

* Subject with device removal (due to lumpiness) was seen at all follow-up visits

- Theoretical = Patients enrolled in the study

- Expected = Theoretical –  Device removals

- Evaluated = Theoretical – (Failures+ Lost to Follow-Up + Aborted)

4	As per evaluated3

| Patient Accounting – Artes Medical Study - Control Patients |
| --- |
|  | Preop | Intraop | 1 month | 3 months | 6 months |
| Theoretical1 | 144 |
| Device removals, (cumulative) | 0 | 0 | 0 | 0 | 0 |
| Expected2 | 144 | 144 | 144 | 144 | 144 |
| Lost to follow-up (cumulative) | 0 | 0 | 0 | 2 | 2 |
| Aborted study (includes not treated and voluntary withdrawal) (cumulative) | 21 | 21 | 21 | 22 | 23 |
| Evaluated3 | 123 | 123 | 123 | 120 | 119 |
| % Expected Follow-up | 100% 144/144 | 100% 144/144 | 100% 144/144 | 100% 144/144 | 100% 144/144 |
| Actual % Follow-up4 | 85% 123/144 | 85% 123/144 | 85% 123/144 | 83% 123/144 | 83% 123/144 |

1.	Theoretical = Patients enrolled in the study

2.	Expected = Theoretical –  Device removals

3.	Evaluated = Theoretical – (Failures+ Lost to Follow-Up + Aborted)

4. 	As per evaluated3

d.  Adequacy of Blinding

1)   Patient Satisfaction

1 month – guess rates were not different than chance

3 months – guess rates were 61.3% (p = 0.001) accurate

6 months – guess rates were 73.6% (p < 0.001) accurate

Additional analysis (table 57 p. 12-053) demonstrates that treatment guesses at 6 months were not related to outcome on the masked observer FFA scale ratings.

2)  Masked Observer - The observers were not asked questions regarding their guesses about treatment assignment or follow-up time for the photos.  No conclusions can be drawn regarding the efficacy of the blinding of the masked observers.

#### Artes Medical Study Results

The sponsor has performed a number of post-hoc supplemental analyses of various subsets of patients.  These subset analyses have not been included in this review memo as they do not provide any additional information that are not already addressed in the prospectively defined data analysis.

  - Safety

Onset of Adverse Events – Artecoll patients in the Artes Medical Study

|  | 24-48 hours | 1 week | 1 month | 3 months | 6 months | 12 months |
| --- | --- | --- | --- | --- | --- | --- |
| Number of patients evaluated | 128 | 128 | 112 | 106 | 113 | 111 |
| Total number of patients with adverse events | 12 | 3 | 2 | 3 | 2 | 0 |
| Total number of adverse events | 14 | 4* | 2 | 4 | 2 | 0 |
| Lumpiness at injection area more than one month after injection | 3 | 0 | 0 | 3 | 2 | 0 |
| Persistent swelling or redness | 6 | 1* | 0 | 0 | 0 | 0 |
| Increased sensitivity | 2 | 0 | 2 | 0 | 0 | 0 |
| Rash, itching more than 48 hours after injection | 0 | 1 | 0 | 1 | 0 | 0 |
| Blurred vision | 0 | 1 | 0 | 0 | 0 | 0 |
| Flu-like symptoms | 0 | 1 | 0 | 0 | 0 | 0 |
| Recurrence of existing herpes labialis | 1 | 0 | 0 | 0 | 0 | 0 |
| Sensitization reactions1 | 0 | 0 | 0 | 0 | 0 | 0 |
| Abscess | 0 | 0 | 0 | 0 | 0 | 0 |
| Visibility of the puncture area | 0 | 0 | 0 | 0 | 0 | 0 |
| Granuloma or enlargement of the implant | 0 | 0 | 0 | 0 | 0 | 0 |
| Infection | 0 | 0 | 0 | 0 | 0 | 0 |
| Other local complications | 1 –dry skin | 0 | 0 | 0 | 0 | 0 |
| Other systemic complications | 1 – vasovagal episode, patient briefly lost consciousness during injection | 0 | 0 | 0 | 0 | 0 |
| Severe illness, trauma, death | 0 | 0 | 0 | 0 | 0 | 0 |

* Subject reported that event occurred 3 additional times – no dates for onset of 2nd and 4th times

1. As defined per each investigator.

Onset of Adverse Events – Control subjects in the Artes Medical Study

|  | 24-48 hours | 1 week | 1 month | 3 months | 6 months |
| --- | --- | --- | --- | --- | --- |
| Number of patients evaluated | 123 | 123 | 111 | 109 | 116 |
| Total number of patients with adverse events | 10 | 1 | 3 | 2 | 4 |
| Total number of adverse events | 21 | 3 | 4 | 2 | 6 |
| Lumpiness at injection area more than one month after injection | 2 | 0 | 1 | 0 | 1 |
| Persistent swelling or redness | 9 | 1 | 1 | 2 | 0 |
| Increased sensitivity | 1 | 0 | 0 | 0 | 0 |
| Rash, itching more than 48 hours after injection | 1 | 1 | 0 | 0 | 0 |
| Blurred vision | 0 | 0 | 0 | 0 | 0 |
| Flu-like symptoms | 0 | 0 | 0 | 0 | 1 |
| Recurrence of existing herpes labialis | 0 | 0 | 0 | 0 | 0 |
| Sensitization reactions1 | 3 | 1 | 2 | 0 | 0 |
| Abscess | 2 | 0 | 0 | 0 | 1 |
| Visibility of the puncture area | 2 | 0 | 0 | 0 | 0 |
| Granuloma or enlargement of the implant | 1 | 0 | 0 | 0 | 0 |
| Infection | 0 | 0 | 0 | 0 | 1 |
| Other local complications | 0 | 0 | 0 | 0 | 1 –acneform |
| Other systemic complications | 0 | 0 | 0 | 0 | 0 |
| Severe illness, trauma, death | 0 | 0 | 0 | 0 | 1 - carjacking |

1. As defined per each investigator.

Duration of Adverse Events:

Please see Attachment 2a and 2b for the Duration of adverse events in Artecoll treated patients and Control group patients.

Summary of incidence of adverse events:

|  | Artecoll | Control |
| --- | --- | --- |
| Number of adverse events | 26 | 36 |
| Number of patients experiencing adverse events | 21 | 16 |
| Total number of subjects treated | 128 | 123 |
| % of subjects with adverse events | 16.4% | 13.0% |

The top 3 adverse events occurring in the Artecoll treatment group were lumpiness more than 1 month post-injection, persistent swelling or redness, and increased sensitivity.

For Artecoll treated patients, these events lasted over 3 months in

- 5/7 patients found to have lumpiness more than 1 month post injection

- 3/7 patients found to have persistent swelling or redness

- ¼ patients found to have increased sensitivity

For control treated patients, these events lasted over 3 months in:

- 2/4 patients found to have lumpiness more than 1 month post injection

- 9/12 patients found to have persistent swelling or redness

- 1/1 patient found to have increased sensitivity

  - Effectiveness

The results of the Masked Observer Assessments are shown in the following table:

| Follow-up | Treatment Area | Treatment | N | Mean change in FFA | P-Value |
| --- | --- | --- | --- | --- | --- |
| 1-month | Glabellar Folds | Artecoll | 64 | 0.17 | 0.004 |
|  |  | Control | 77 | 0.49 |  |
|  | Nasolabial Folds | Artecoll | 91 | 0.75 | 0.713 |
|  |  | Control | 91 | 0.74 |  |
|  | Upper Lip Lines | Artecoll | 58 | 0.31 | 0.205 |
|  |  | Control | 53 | 0.48 |  |
|  | Mouth Corners | Artecoll | 71 | 0.46 | 0.179 |
|  |  | Control | 76 | 0.30 |  |
| 3-months | Glabellar Folds | Artecoll | 65 | 0.25 | 0.348 |
|  |  | Control | 75 | 0.35 |  |
|  | Nasolabial Folds | Artecoll | 87 | 0.81 | <0.001 |
|  |  | Control | 88 | 0.15 |  |
|  | Upper Lip Lines | Artecoll | 53 | 0.18 | 0.454 |
|  |  | Control | 51 | 0.25 |  |
|  | Mouth Corners | Artecoll | 64 | 0.45 | 0.001 |
|  |  | Control | 77 | 0.01 |  |
| 6-months | Glabellar Folds | Artecoll | 71 | 0.34 | 0.971 |
|  |  | Control | 79 | 0.32 |  |
|  | Nasolabial Folds | Artecoll | 92 | 0.77 | <0.001 |
|  |  | Control | 91 | 0.00 |  |
|  | Upper Lip Lines | Artecoll | 55 | 0.08 | 0.176 |
|  |  | Control | 50 | 0.22 |  |
|  | Mouth Corners | Artecoll | 69 | 0.26 | 0.316 |
|  |  | Control | 79 | 0.09 |  |

The un-blinded investigator assessments of success (ranging from not at all successful to completely successful; 1 – 5 scale) are shown in the following table:

| Follow-up | Treatment Area | Treatment | N | Mean | Standard Deviation |
| --- | --- | --- | --- | --- | --- |
| 1-month | Glabellar Folds | Artecoll | 68 | 1.93 | 0.94 |
|  |  | Control | 79 | 1.91 | 1.00 |
|  | Nasolabial Folds | Artecoll | 93 | 1.99 | 0.89 |
|  |  | Control | 93 | 2.06 | 1.03 |
|  | Upper Lip Lines | Artecoll | 61 | 1.88 | 0.82 |
|  |  | Control | 54 | 2.02 | 0.88 |
|  | Mouth Corners | Artecoll | 74 | 2.00 | 0.97 |
|  |  | Control | 77 | 2.55 | 1.28 |
| 3-months | Glabellar Folds | Artecoll | 67 | 2.02 | 1.00 |
|  |  | Control | 74 | 2.68 | 1.40 |
|  | Nasolabial Folds | Artecoll | 89 | 1.90 | 0.87 |
|  |  | Control | 89 | 3.07 | 1.41 |
|  | Upper Lip Lines | Artecoll | 58 | 2.04 | 0.90 |
|  |  | Control | 51 | 3.11 | 1.39 |
|  | Mouth Corners | Artecoll | 68 | 2.08 | 0.94 |
|  |  | Control | 76 | 3.38 | 1.34 |
| 6-months | Glabellar Folds | Artecoll | 74 | 2.06 | 1.07 |
|  |  | Control | 82 | 3.43 | 1.51 |
|  | Nasolabial Folds | Artecoll | 97 | 1.73 | 0.69 |
|  |  | Control | 96 | 4.05 | 1.32 |
|  | Upper Lip Lines | Artecoll | 60 | 1.98 | 0.93 |
|  |  | Control | 54 | 4.13 | 1.28 |
|  | Mouth Corners | Artecoll | 73 | 2.19 | 0.98 |
|  |  | Control | 81 | 4.09 | 1.32 |

Patient Satisfaction

| Time (months) | Treatment area | Artecoll | Control |
| --- | --- | --- | --- |
|  |  | N | Mean | Std. Dev. | N | Mean | Std. Dev. |
| 1 | Glabellar Folds | 66 | 1.86 | 0.97 | 78 | 1.94 | 0.94 |
| 1 | Nasolabial Folds | 92 | 1.85 | 1.03 | 93 | 1.92 | 0.89 |
| 1 | Upper Lip Lines | 61 | 1.94 | 1.09 | 54 | 1.98 | 1.04 |
| 1 | Mouth Corners | 74 | 2.01 | 0.94 | 77 | 2.32 | 1.07 |
| 3 | Glabellar Folds | 67 | 2.22 | 1.13 | 74 | 2.70 | 1.34 |
| 3 | Nasolabial Folds | 89 | 2.16 | 1.08 | 89 | 2.78 | 1.30 |
| 3 | Upper Lip Lines | 58 | 2.14 | 1.05 | 51 | 3.07 | 1.30 |
| 3 | Mouth Corners | 68 | 2.29 | 1.08 | 77 | 3.06 | 1.29 |
| 6 | Glabellar Folds | 74 | 2.14 | 1.19 | 82 | 3.44 | 1.35 |
| 6 | Nasolabial Folds | 97 | 2.02 | 0.95 | 96 | 3.52 | 1.37 |
| 6 | Upper Lip Lines | 60 | 2.17 | 1.12 | 54 | 3.94 | 1.28 |
| 6 | Mouth Corners | 73 | 2.31 | 1.04 | 81 | 3.97 | 1.19 |

Both the unmasked investigator assessment of patient success and the blinded patient satisfaction rating demonstrated higher success in the Artecoll treated group versus the control.  However, the potential for bias is present as the investigator was not blinded to treatment group and the assessment of patient blinding demonstrated that over 73.6% of patients accurately guessed their randomization assignment at 6 months.

IV.  Did the Study Meet Its Objectives?

Objective 1 (a primary objective) was to explore whether the cosmetic correction provided by Artecoll at the end of a six-month period following injection is superior to that provided by Zyplast/Zyderm at the same time period.

Although a statistically significant difference was noted at six months for patients in the Artecoll treatment arm versus the control in the area of the NL fold, it is unknown whether this represents a clinically detectable difference as the FFA scale only recognizes full point differences.  The findings may be skewed by the statistically significant difference of the pre-treatment severity score with the Artecoll group that was greater (more severe at baseline) than the control group.  The sponsor states that this issue was addressed by using a non-parametric Mann-Whitney U tests for these comparisons since the severity ratings were positively skewed.   The ability of this test to adequately compensate for this difference should be considered.

The results show that Artecoll efficacy in other 3 treatment areas versus the control has not been demonstrated as per objective 1.

The sponsor states that Artecoll efficacy from the perspective of the unblinded investigator and blinded patient demonstrated higher success and better patient satisfaction than the control.  This is useful information however is biased due to unblinding of the investigator and lack of effective blinding of the patient as demonstrated by the blinding assessment.

Objective 2 (primary objective) To explore whether the safety of Artecoll as an injectable implant for correction of contour deformities of the dermis of the face

Based on the information provided in this executive summary, the panel should address:

- The level of safety required for a cosmetic device

- Whether this objective was met for Artecoll based on the information presented

Objective 3 (a secondary objective) To characterize the initial quality of the cosmetic correction provided by Artecoll and Zyplast/Zyderm.

The study attempted to characterize the initial quality of cosmetic correction by the devices as per the patients’ satisfaction, investigator assessment and masked observer FFA scale rating.  Each of these assessment tools has associated flaws whether it be bias, differences in observation from photos versus in-person assessment, baseline severity, use of the FFA scale in methods that have not been validated, or level of expectation.

Objective 4 (a secondary objective) of the study is to characterize the investigator assessment of success with respect to how closely the treatment met the investigator’s initial expectations for correction.

This is a vague objective as the investigators assessment of success was based on the investigator’s initial expectations for correction.  The investigator’s initial expectations for correction were never defined in the protocol.

Objective 5 (a secondary objective) is to characterize the subject’s assessment of satisfaction with respect to the subject’s personal expectations.

This is a vague objective as the subjects’ assessment of success was based on their initial expectations for correction.  The subjects’ initial expectations for correction were never defined in the protocol.

PART III

I.  	Rofil Clinical Study

a.  	Background – This study was intended to evaluate the safety and effectiveness of Artecoll when used for cosmetic correction.  However, the originally approved protocol did not involve a control group, and the ratings made by the investigators were based on improvement rather than objective measurement.  The protocol changed significantly during the process of obtaining unconditional approval.  Also, the method of capturing efficacy data was too subjective to be adequate, and the number of subjects and areas treated were not sufficient to be considered statistically significant.

For all of these reasons, the primary focus of the analysis of the Rofil study has been to obtain the safety of Artecoll when used for cosmetic correction.

b.  Design  -

- Open label study

- Uncontrolled

  - Inclusion/Exclusion

| Inclusion | Exclusion |
| --- | --- |
| Subjects 18 years of age or older In the opinion of the investigator, the subjects has realistic expectations of the benefit and limitation of the augmentation procedure Subjects presenting for correction of dermal contour deformities of the face: - Glabellar Frowns/Folds Nasolabial Folds Perioral Lines ) Depressed Corners of the Mouth | Subjects who had been treated with collagen, Botulinum toxin, or other wrinkle/fold therapies within the last six months at any intended implant area Subjects who were receiving current treatment with corticosteroids subdermally, intradermally or epiperiostally Subjects receiving UV light therapy Subjects presenting with a personal or family history of autoimmune disorder Subjects presenting with atrophic skin diseases Subjects with thin and flaccid skin Subjects with known susceptibility to keloids Subjects with known allergy to collagen Subjects requiring correction of atrophy or defects in the subcutaneous fat Subjects requiring correction of crow’s feet |

d. 	Treatment Protocol

1.	Treatment areas included:

Glabellar Folds

Nasolabial folds

Perioral lines

Depressed mouth corners

- Each subject received as many as 6 facial treatment areas (3 of the 4 per side) as deemed appropriate by the investigators

e.  	Follow-up Protocol - Subjects had follow-up assessments at 3, 6, and 12 months following the last treatment session.

II.  Rofil Medical Study Conduct

- 157 patients were treated with Artecoll

- One year safety data was obtained for 126 of the 157 patients treated

| Patient Accounting - Rofil Study |
| --- |
|  | Preop | Intraop | 3 months | 6 months | 12 months |
| Theoretical1 | 167 |
| Device removals, (cumulative) | 0 | 0 | 0 | 0 | 0* |
| Expected2 | 167 | 167 | 167 | 167 | 167 |
| Lost to follow-up (cumulative) | 5 | 0 | 24 | 6 | 2 |
| Aborted study (includes not treated and voluntary withdrawal) (cumulative) | 5 | 0 | 0 | 0 | 0 |
| Evaluated3 | 157 | 157 | 133 | 127 | 125 |
| % Expected Follow-up | 100% 167/167 | 100% 167/167 | 100% 167/167 | 100% 167/167 | 100% 167/167 |
| Actual % Follow-up4 | 94% 157/167 | 94% 157/167 | 80% 133/167 | 76% 127/167 | 75% 125/167 |

*2 excisions after 12 month time period

1.	Theoretical = Patients enrolled in the study

2.	Expected = Theoretical –  Device removals

3.	Evaluated = Theoretical – (Failures+ Lost to Follow-Up + Aborted)

4. 	As per evaluated3

III.  Rofil Medical Study Results

Onset of Adverse Events – Rofil Study

|  | 24-48 hours | 1 week | 1 month | 3 months | 6 months | 12 months |
| --- | --- | --- | --- | --- | --- | --- |
| Number of patients evaluated | NA | NA | NA | 133 | 127 | 125 |
| Total number of patients with adverse events | 6 | 2 | 1 | 3 | 5 | 3 |
| Total number of adverse events | 7 | 2 | 1 | 3 | 6 | 6 |
| Lumpiness at injection area more than one month after injection | 2 (2 estimated?) | 2 | 0 | 0 | 1 | 0 |
| Persistent swelling or redness | 1 (1 estimated?) | 0 | 0 | 0 | 2 | 0 |
| Increased sensitivity2 | 1 (1 estimated?) | 0 | 0 | 0 | 0 | 0 |
| Rash, itching more than 48 hours after injection | 0 | 0 | 0 | 2 | 1 (1 estimated?) | 1 (1 estimated?) |
| Blurred vision | 0 | 0 | 0 | 0 | 0 | 0 |
| Flu-like symptoms | 0 | 0 | 0 | 0 | 0 | 0 |
| Recurrence of existing herpes labialis | 0 | 0 | 0 | 0 | 0 | 0 |
| Sensitization reactions1 | 1 (1 estimated?) | 0 | 0 | 0 | 1 | 0 |
| Abscess | 0 | 0 | 0 | 0 | 0 | 0 |
| Visibility of the puncture area | 0 | 0 | 0 | 0 | 0 | 0 |
| Granuloma or enlargement of the implant | 0 | 0 | 0 | 1 | 0 | 2 (1 after 12 months) |
| Infection | 0 | 0 | 0 | 0 | 0 | 0 |
| Other local complications2 | 0 | 0 | 0 | 0 | 1 | 3 (2 estimated after 12 months) |
| Other systemic complications2 | 2 | 0 | 0 | 0 | 0 | 0 |
| Severe illness, trauma, death | 0 | 0 | 1 – breast cancer | 0 | 0 | 0 |

1. As defined per each investigator.

2  Local complications -  “redness and visible capillaries in NF” and “patchy complete alopecia areata on head at 3 mos”.

3  Systemic complications - “mild chest congestion after both treatment sessions.”

Duration of Adverse Events:

See Attachment 2c:  Duration of Adverse Events – Rofil Study

Based on the Duration of Adverse events attachment 2c, the top 3 specific adverse events occurring on a per incident basis in the Rofil treated patients were lumpiness more than 1 month post-injection, persistent swelling or redness, and granuloma or enlargement of the implant.

For Rofil treated patients, these events lasted over 3 months in:

- 4/5 patients found to have lumpiness more than 1 month post injection

- 1/3 patients found to have persistent swelling or redness

- 1/1 patient granulomas or enlargement of the implant

PART IV

I.  Product Description

The test material is a 3.5% purified bovine collagen manufactured by European Medical Contract Manufacturing.  The specific formulation is 3.5% collagen, 2.7% phosphate buffer, 0.3% sodium chloride, and 0.3% lidocaine hydrochloride and 93.2% WFI.  Test material is provided sterile and nonpyrogenic in a prefilled 1.0 cc syringe with 0.1cc of collagen solution.

II.   Bovine Collagen Immunogenicity Study

- Design

- Single investigator, prospective, open-label non-controlled clinical trial

- To determine the safety and immunogenicity of the bovine collagen in Artecoll because collagen source for the proposed commercial product differs from the collagen component used in the large scale clinical study (i.e., the Artes Study).

- Inclusion/Exclusion Criteria

| Inclusion | Exclusion |
| --- | --- |

| Subjects 35 years of age of older, of either sex Subjects wiling and able to comply with study requirements Subjects willing and able to comply study follow-up requirements Subjects willing and able to give informed consent | Subjects who are pregnant Subjects who have had collagen treatment within the last 3 months Subjects who were treated with chemotherapy agents or corticosteroids within the past 3 months Subjects with a history of autoimmune disorder Subjects with known lidocaine hypersensitivity Subjects with known sensitivity to bovine collagen Subjects who have a history of dietary beef, undergoing desensitization to beef products or planning to undergo desensitization within the study evaluation period. Subjects with severe allergies manifested by a history of anaphylaxis Subjects with a current disease state that can effect the immunoresponse (e.g., flu, cancer, HIV) Subjects who are currently treated with immunosuppressive drugs |
| --- | --- |

c.  Treatment Protocol

- Subjects screened

- If entry criteria met then patient had a pre-injection serum sample drawn

- The physician then injected 0.1 ml of a collagen solution intra-dermally in the volar forearm

- The physician then assessed the test site for an acute reaction for 15-30 minutes after the injection

- The subject was permitted to go home and was instructed to observe the test site daily for 30 days and notified the investigator if any effects indicative of a positive response were observed or systemic effects were experienced.  The subject received both written instructions for assessing the test site and a “Patient Skin Response Sheet”

- 30 days after the first injection the subjects returned to the clinic.  If the investigator determined that the subject displayed a positive response, the subject’s participation in the study was completed and a final blood drawn was taken.

- If the subject did not display a positive response to the first collagen injection, they received a second 0.1 ml injection intradermally in the contra lateral volar forearm.

- The physician then assessed the test site for an acute reaction for 15-30 minutes after the injection after which the subject went home and observed the test site daily for 30 days.  The subjects were instructed to notify the investigator if any effects indicative of a positive skin test response or systemic effects were experienced.  The subjects also recorded observations in a “Patient Skin Response Sheet”

d. Follow-up Protocol

- Visit 2 30 days following initial skin test

- Visit 3 following 2nd skin test

- Daily evaluations by patient

e.  Follow-up Assessment Method

1)  Skin test results -  The injection site was evaluated to determine negative, positive or equivocal response:

- Negative = No localized skin reaction and the patient has no systemic reaction

- Positive = Erythema of any degree, induration, tenderness, rash and swelling with or without pruritis, which can appear immediately following implantation and persists for more than 24 hours or appears more than 24 hours following implantation

- Equivocal =No localized skin reaction but the patient does elicit a possible systemic reaction such as rash, arthralgia, or myalgia which occurs at any time during the 30 (+/- 5) day observation period.

2)  Abnormal Serum IgG Test Results  - the number of patients exhibiting a normal serum IgG level before administration of the skin test and the an abnormally high serum IgG at the time of the second blood draw

3) Adverse events

f. Study Purpose and Objective

Purpose  To determine the frequency of positive collagen skin test results for a new Artecoll collagen source.

Objective  To provide information that can be used in labeling to be submitted to the US Food and Drug Administration and other regulatory bodies in connection with applications for marketing approval.

III. Bovine Collagen Study Conduct

a. Enrollment

- 235 subjects were injected with bovine collagen.

- 225 subjects completed the study per protocol.

b. Demographics

| N = 235 |
| --- |
| Gender |
| Male | 78 (32.8%) |
| Female | 157 (67.2%) |
| Mean Age | 48.2 years |
| Ethnicity |
| Caucasian | 188 (80.0 %) |
| Hispanic | 32 (13.6 %) |
| Black | 11 (4.7 %) |
| Asian | 2 (0.085%) |
| Black & Hispanic | 1 (0.042%) |
| Native American | 1 (0.042%) |

c. Accountability

| Patient Accounting |
| --- |
| Time of: |  | Injection 1 | Injection 2 |  |
|  | Screened | Enrolled | Visit 1 | Visit 2 | Visit 3 |
| Expected1 | 244 | 235 |
| Lost to follow-up (cumulative)3 | 0 | 0 | 0 | 7 | 2 |
| Aborted study (includes not treated and voluntary withdrawal) (cumulative)4 | 0 | 0 | 0 | 0 | 1 |
| Evaluated5 | 244 | 235 | 235 | 228 | 225 |
| Experienced skin reaction but were not discontinued (cumulative) | 0 | 0 | 0 | 5 | 8 |
| Actual % Follow-up6 | 100% 244/244 | 100% 235/235 | 100% 235/235 | 97% 228/235 | 97.5% 225/235 |

1.	Expected = Patients enrolled in the study

2.	Of the patients lost to follow-up, 5/9 were contacted and reported no local skin reaction

3.   Patient violated the exclusion criteria, but had no skin test reaction or abnormal elevation in serum IgG levels.

4.   Evaluated = Theoretical – (skin reactions+ Lost to Follow-Up + Aborted)

5.   As per evaluated

IV. Study Results

a.	 Positive Skin Tests Responses –

|  | Within 15-30 minutes Post Injection 1 | Within 30 days Post Injection 1 | Within 15-30 minutes Post Injection 2 | Within 30 days Post Injection 2 |
| --- | --- | --- | --- | --- |
| Number of patients with positive skin test (8 patients of 235 studied = 3.4%) | 0 | 4 | 1 | 3 |

b.	Adverse Events

Time to onset of Adverse Events

| Category | 24-48 hours | 1 week | 1 month |
| --- | --- | --- | --- |
| Number of subjects evaluated | 235 |
| Total number of adverse events | 5 | 22 | 31 |
| Total number of patients with adverse events | 5 | 18 | 26 |
| Arm tingling, warmth, clammy | 1 | 0 | 0 |
| Broken toe | 0 | 1 | 0 |
| Broken wrist | 0 | 0 | 1 |
| Cataract surgery | 0 | 0 | 1 |
| Cellulitis | 0 | 1 | 0 |
| Cough | 0 | 1 | 3 |
| Dental extraction | 0 | 0 | 1 |
| Diarrhea | 0 | 3 | 0 |
| Facial and Hand Edema | 0 | 1 | 0 |
| Flu Symptom | 0 | 0 | 1 |
| Headache | 1 | 0 | 1 |
| Increased perspiration | 1 | 0 | 0 |
| Irritant dermatitis | 1 | 0 | 0 |
| Lymphadenopathy | 1 | 0 | 0 |
| Migraine Headache | 0 | 0 | 1 |
| Muscle Ache | 0 | 1 | 0 |
| Recurrent Herpes simplex virus | 0 | 1 | 0 |
| Recurrent oral herpes simplex virus | 0 | 0 | 1 |
| Rhinitis | 0 | 0 | 1 |
| Sinus congestion | 0 | 1 | 1 |
| Sore throat | 0 | 0 | 2 |
| Stomach Ache | 0 | 0 | 1 |
| Swollen glands right neck | 0 | 1 | 0 |
| Tooth infection | 0 | 0 | 1 |
| Torn ligament right foot | 0 | 0 | 1 |
| Upper respiratory infection | 0 | 8 | 13 |
| Urinary tract infection | 0 | 0 | 1 |
| Vomiting | 0 | 3 | 0 |

Duration of Adverse Events

See attachment 2d:  Bovine Collagen Immunogenicity Study – Duration of Adverse Events

c. 	Serum Levels of Bovine Collagen Antibodies

The Center testing laboratory defined the normal serum IgG range as a titer between 700 and 1,600.  No subjects transitioned from the normal serum IgG level before administration to an abnormally high serum IgG level in the post-treatment blood sample.

V. Did the Study Meet Its Purpose and Objective?

Purpose  To determine the frequency of positive collagen skin test results for a new Artecoll collagen source.

The incidence of a positive skin response was found to be 3.4%.  The incidence of equivocal skin test results as prospectively defined is not well characterized and should be assessed by examination of adverse events.

Objective  To provide information that can be used in labeling to be submitted to the US Food and Drug Administration and other regulatory bodies in connection with applications for marketing approval.

The information provided can be used in device labeling.

Of note, the labeling of ZYDERM collagen implant states that the incidence of positive test site response or similar systemic responses as described here has been approximately 3.0%.