PHENIX TEC OPERATIONS IN THE PEH

PHENIX Procedure No. PP-2.5.2.6-02

Revision: A                                     Date: 5-8-00

Hand Processed Changes

HPC No.			Date			Page Nos.		Initials

______________		______________	______________	______________

______________		______________	______________	______________

______________		______________	______________	______________

Approvals

______________________________	__________________________________________

PHENIX S E & I     Date				Cognizant Scientist/Engineer   Date

/Activity Manager

______________________________	__________________________________________

PHENIX Safety    Date				CA-D  ES&H/SAFETY  Date

______________________________	__________________________________________

# Purpose

The purpose of this document is to define the local emergency plan for operation of the Gas, High Voltage and Low Voltage Systems for the PHENIX Time Expansion Chamber. In addition to the Local Emergency Plan this document specifies the Operating Procedures of the Time Expansion Chamber (TEC). This Local Emergency Plan will ensure:

# the safety of all personnel from risks associated with the operation of the gas, HV or LV systems required for running the Time Expansion Chamber

# the implementation of the appropriate emergency procedures,

# prompt notification of the appropriate C-A and S&EP specialists,

# the maintenance of appropriate C-A emergency status,

# the preservation and protection of the environment, and

# the preservation of BNL facilities and equipment, and

# Responsibilities

During PHENIX operations there will be two levels of responsibility for the oversight of the TEC HV, LV and gas systems: the PHENIX shift crew and the TEC experts.

The first level of responsibility resides with the PHENIX Shift Crew.  During any period when the TEC has flammable gas flowing or HV on, there will be a minimum of two  people on continuous shift in the PHENIX counting house.  Once data taking has begun the number of people on shift  at PHENIX increases to four.  The second level of responsibility resides with the TEC Gas, HV and LVexperts. The TEC experts will be on-call to respond to any alarm or unusual occurrence detected by the PHENIX  shift crew. A record  of the performance of the TEC gas system will be maintained and monitored by the TEC experts.

During watch shifts or data taking, it will be the responsibility of the PHENIX Shift Crew to:

  - monitor the status and alarms for the gas and HV  system.

  - In the event of an alarm or unusual occurrence, contact person expert from the Expert Call List given in Appendix A.

The second level of responsibility is the TEC  experts.  It is the responsibility of the TEC Gas experts to:

# maintain the TEC Gas  System in a safe operating condition.  This includes:

## changing gas cylinders and dewars  when required,

## setting, adjusting, and checking the gas mixture, flow rates and pressures.

## checking the certification of the operating gas, (see details in Precautions, Section 4)

# posting any special instructions or notifications as required, and

# carrying out any emergency actions, as prescribed in the Procedures section of this document.

# maintain the TEC HV Systems in a safe operating condition. This includes

## verifying the readiness of the chamber for HV,

## setting, adjusting and maintaining the TEC HV settings

## turning on the HV according to the operating procedures described below,

# posting any special instructions or notifications as required, and

# carrying out any emergency actions, as prescribed in the Procedures section of this document.

# maintain the TEC LV Systems in a safe operating condition. This includes:

## verifying the readiness of the TEC electronics for LV,

## setting, adjusting and maintaining the TEC LV settings

# posting any special instructions or notifications as required, and

# carrying out any emergency actions, as prescribed in the Procedures section of this document.

## Prerequisites

The TEC Gas, HV and LVexperts shall have read or have training in the following areas:

# PHENIX Local Emergency Plan, RHIC-OPM 3.16,

# BNL Compressed Gas Safety Training Course,

# BNL Electrical Safety I,

# RHIC/PHENIX access training.

# BNL Hazcom

# BNL General Employee Training

## Precautions

# The safety of personnel is of primary importance.  The TEC experts shall take great care to ensure that the TEC Gas, High Voltage and Low Voltage Systems will be operated in a way that does not place personnel or equipment at risk of physical harm.

# Gas, HV, and LV  System Precautions:

# All gas cylinder storage is on the PHENIX Gas pad located just south and east of Building 1008 F, the PHENIX Gas Mixing Hut. All  gas cylinders and dewars are to be changed by designated PHENIX personnel with current BNL Compressed Gas Safety Training.

# All valves and controls associated with the TEC Gas system are to be operated ONLY by  designated TEC gas system experts with current training.

# Primary care should be given to monitoring the internal pressure of the TEC throughout the duration of the start-up procedure, especially when adjusting flow rates. Over-pressurization of the TEC (above 7.0 mm WC) can result in structural damage to the detector.

# The total energy stored in the TEC with High Voltage on is low, ~ 10 mJoules, adjusting the voltage is not dangerous from a personal safety stand point. Additionally,  all HV points are covered on the TEC such that one would have to remove covering material in order to feel any shock. However, there is significant risk of damaging the equipment, the TEC, without proper operation of the HV.

# Before any HV can be turned on, sufficient P-10 gas must have flowed through each of the detectors for 4 volume exchanges.  The total gas volume of all TEC planes on the East carriage is 6000 liters. The TEC HV is to be turned on only by a TEC HV expert with current training. An exception is listed in 4.1.6.

# All HV controls associated with the TEC HV system are to be operated ONLY by designated TEC HV experts or  by  the PHENIX shift leader following specific instructions from TEC experts.

# Any reconfiguration or adjustment to the TEC Gas or HV system in the PHENIX IR is to be performed ONLY by  TEC Gas and HV experts with current training.

# Any reconfiguration of the TEC LV beyond turning the power on/off  is to be performed by TEC LV experts or people specifically designated for the job.

## Emergency Procedures

## In the event of a fire or emergency in Building 1008, members of the PHENIX Shift Crew shall (in order of priority)

## Follow emergency procedures described in C-A OPM 3.16 for the PHENIX area.

### All TEC HV should be turned -off by the PHENIX shift crew or TEC experts..

### Confirm that the TEC Gas System is in safe/purge mode.  The shift leader or TEC Gas system expert can confirm that the alarm/safety system has functioned properly by going to the PHENIX Gas Mixing Hut 1008F, and observing that the AC power to the TEC Gas system is off , the argon flow meter shows flow and that the CH4 flow meter no-flow. This action should be taken once it has been determined   that no hazard exists in 1008F.

### In the event that a TEC gas line ruptures, the Shift Leader or TEC  Gas System expert  shall:

### Turn off all gas flowing to the TEC  at either the TEC Gas system computer control in the Gas Mixing Hut, the gas distribution values in the Mixing Hut or the cylinders and dewars on the PHENIX gas pad.

### Turn off the power to all TEC high voltage (HV).

# Contact a TEC Gas expert . ( See Appendix A).

## Standard Operating Procedures

# There are three elements to these Operating Procedures.  The first element covers running gas to the chambers.  The second covers the high voltage system for the chambers.  The third element covers TEC LV controls. These Procedures shall be carried out by TEC Gas, HV and LV  experts only.

# The procedures for operating the TEC Gas, HV and LV systems are described in general terms in this OPM. A more detailed description of the TEC Gas system and its operation can be found in the PHENIX Central Tracking Gas System Design and Specification Manual (March 2000), especially see the section on the Alarm System in this document.  The TEC Gas system experts will have familiarized themselves with this document. A more detailed description of the PHENIX HV system and its operation can be found in PHENIX HV Control Manual 2.5.3.14-08

# The Gas system for the TEC is composed of a number of components.

# Gas source of cylinders and dewars located on the PHENIX Gas Pad.

# Gas distribution manifolds located in the PHENIX Mixing Hut 1008F

# TEC Gas rack located in the Gas room in 1008F

# TEC Gas Electronics Rack located in the Electronics room of 1008F

# TEC carriage gas  rack located underneath the East carriage in the PHENIX IR.

# TEC sector-flowmeters and flow direction indicators mounted on the TEC sectors in the PHENIX East carriage.

# When the PHENIX IR is closed the status of the TEC Gas system can be monitored in a number of places.

# TEC Gas Control computer located in the electronics room of 1008F shows the status of all electronically readout  flow meters, flow direction indicators, pressure transmitters/indicators, solenoid valves and temperature transmitters.

# TEC Gas Control rack located in the Gas room of 1008F shows status of all manually controlled flow meters and pressure gauges.

# PCs located in PHENIX Counting House allows monitoring of  TEC Gas system logging file created by the TEC Gas Control computer.

# TV  monitor located in PHENIX Counting House allows one to use cameras in IR  to view status of  flow meters and pressure gauges located in TEC carriage gas rack underneath the East carriage.

For details of the gas system, please refer to the TEC Gas System diagram located in Attachment B and the PHENIX Central Tracking Gas System Design and Specification Manual, March 2000.  The labels of valves and meters in the operation instructions below refer to parts identified in the figures in Attachment B.

Note: All the flow rate settings specified in the procedures below are nominal and can be adjusted by the TEC gas system experts as in accordance with the requirements of the TEC detector or gas system.

## Gas System Procedures The Procedures describe starting gas flowing to the TEC after a long shutdown. They assume no gas is currently  flowing, no flow  meters are adjusted and the PHENIX IR is open. If the PHENIX IR is closed it is necessary that the TEC carriage rack and sector flow meters have been previously adjusted prior to any gas flow. Check that  PP-2.5.3.4-09 Check-off List for Ending Access to the PHENIX IR has been completed. If no gas has been flowing in the TEC for 24 hours one must start flowing inert gas first.

## Power on the TEC Gas System

# Note: If system is powered on, skip to step 6.1.2

      - Turn on AC power for Electronic Rack (Receptacles Breaker)

      - Turn on mass flow controller power supplies (Electronic Rack)

      - Turn on Keithley Digital Multi-meter and High Density Switch System

      - Turn on Alarm/Interlock box

      - Turn on and boot the gas system PC.

      - Start the gas system control program.

      - Turn on Power Supplies Breaker

      - Turn on AC power for Gas Rack

## Prepare for Inert Gas Flow (either argon or nitrogen)

Note: All solenoid valves and compressors are controlled with the gas system PC . Green light means SV is open and red light is closed.

# Caution: For  purging with inert gas, make sure  the PID controller bypass valve is turned OFF( Nitrogen supply line to  PID controller  should be shut off with manual valve MV5 in Gas  rack). It is only used for gas re-circulation.

      - Close SV6 (N/O) (red light) and SV7 (N/O) (red light). This closes off manual flow indicator FI3 and flow meter FM1 (see attached schematic)

      - Close FI3, FI4 (Gas Rack), FI12 (East Carriage Gas Rack), FI11(West Carriage Gas Rack).

      - On the gas pad, supply inert gas to Argon supply line at 20PSIG.

      - In the mixing room, purge for short time(several seconds) Argon supply line to the vent line.

      - In the mixing room Gas Rack confirm gas delivery pressure 20PSIG indicated on PI6.

      - Confirm that MV3 (Compressor Bypass, Gas Rack) is closed and MV8(BPCV-2 Bypass, Gas Rack) is  opened.

      - Check  SV21 (N/O) status (Vent, East Carriage Rack) no/light is open/light is closed.

      - Check that SV16(N/O)(no light)(GMH Input, East Carriage Gas Rack)), SV17(N/O)(no light)(GMH Return, East Carriage  Rack)are open .

      - Confirm that  MV10 is closed (Bypass valve for detector maintenance,  East Carriage Rack)

## Begin inert gas flow

      - Open SV6(green light).

      - Slowly open FI3 and set flow to flow rate of 15 SCFH (~ 6 lpm) nominal or prescribed rate.

Caution: After starting or changing flow, wait and observe the pressure rise	on PI1, PI2, PI3, PI8, PT9 and PT11. To maintain a margin of safety, pressure for this procedure should not exceed 2-3mm WC on PI1,  PT8. Structural damage to  the East TEC can result for a pressure >7.0mm WC.  If the pressure  becomes too high, reduce the flow or close FI3.

      - Using PCV1 and BPCV1, adjust pressure of  40”WC on PI2.

      - Slowly closing MV8 and using PCV1 and BPCV2, adjust pressure of  6”WC on PI3 (Gas Rack).

      - Confirm that  FI8 flow rate is 0 and FI2 = 15SCFH nominal or prescribed flow rate.

      - Slowly open FI12(East Carriage Gas Rack), watch PI8 pressure and set  flow to 15SCFH nominal or prescribed flow rate.

      - Adjust TEC sector flow meters until the flow is equalized between each of the 4 TEC sectors.

      - Slowly open FI3 ,watch PI1 and set flow 30SCFH (12 LPM) nominal or prescribed flow rate.

      - Record the start time, flow rates and TEC pressure  in the data sheet ( Attachment E).

    - Initial Set-up for P10 Flowing

Note: If you have been flowing nitrogen through the TEC this is the point where you switch to argon. If you have been flowing argon through the TEC these steps are necessary prior to mixing with methane.

      - Close SV6(red light)

      - Close FI3

      - In Gas Pad connect Argon to Argon Supply Line

      - In Mixing House confirm that delivery pressure is 20 PSIG on PI6

      - Confirm that SV16(n/o)(no light), SV17(n/o)(no light) are open

      - Check  SV21(n/o)status (East Carriage Rack to Vent) no-light is open/light is closed.

      - Confirm that MV10 is closed

      - Confirm that MV3 is closed

      - Using Attachment 6, set/confirm all SV, MV and FI's to their initial states.

This is the point if you had been flowing nitrogen through the TEC you will switch to argon. If you had been flow argon continue with the next steps.

      - Confirm that MV8 is open

      - Open SV7(green light)

      - Set FM1 flow to 3LPM(Electronic Rack)

### CAUTION: Structural damage to the TEC can result for a pressure                > 7.0mmWC If the pressure becomes too high, reduce the flow. After starting or changing flow, wait and observe the pressure rise on PT9 (TEC pressure).

      - Confirm that the flow through FM1 = 3 LPM(Electronic Rack)

      - Slowly close MV8

      - Confirm that PI2 pressure is 40”WC, PI3 and PT11 –not above 6”WC

      - Confirm FI8 Zero flow

      - Slowly increase flow through FM1 to 10 LPM, argon flow.

      - If you have been flowing nitrogen, exchange 3 East TEC volumes (East TEC volume = 6,000 liters). If you have been flowing argon go to the next step. Time in hours for 3 volume exchanges = 18,000 liters/flow*60. For 10 LPM, time = 30 hours.

      - Open MV3

      - Start Big Compressor and give it to run for 5 seconds. Repeat it 3 times.

      - Stop Big Compressor

      - Close MV3

## Prepare for P10 flow

## Notify the MCR that flammable gas is being put into PHENIX so that they will heed the flammable gas alarms.

      - Confirm that FI4 is closed

      - Open SV9(green light)

      - Slowly open FI4 and set 10SCFH flow. Purge for 10 seconds.

      - Close FI4.

      - Close SV9(red light)

      - Open SV8(green light)

      - Set 1LPM flow through FM2. Purge for 10 seconds.

      - Set 0LPM flow through FM2.

      - Close SV8(red light).

    - Start P10 flow

      - Confirm the PHENIX SMCS is operating.

      - Confirm that the control pots (labeled "command") for FM2 (Electronic Rack ) are turned fully counterclockwise.

CAUTION: The mass flow controller pot used in the following step is very sensitive. The pot should be turned in very small increments.

      - Confirm that Methane Bank was connected to Methane supply line on Gas Pad

      - Set 15PSIG  Methane pressure with pressure regulator in the Mixing house. Open MV-M10A on Methane Distribution panel.

      - Confirm that Methane Supply pressure(PI7) is 15PSIG (Gas Rack)

      - Confirm than Nitrogen Purging Line at 15PSIG

      - Open MV6(Gas Rack)

      - Set FI15 flow to purge PI1 case to 1.0LPM, nitrogen flow to photohelic.

      - Open SV8(green light)

      - Set FM2 flow to 1.0 LPM nominal to obtain 90:10 mixture of  argon/methane gas mixture.

NOTE:      At this point, the methane mass flow controller (FM2) is slaved to the argon mass flow controller (FM1).	Increasing the argon flow will automatically increase the methane flow, maintaining the 90:10 ratio.

      - Slowly raise FM1 argon flow to 20LPM

      - Set FM1 flow to appropriate value, 10LPM nominal for long term operation.

      - Enable  Alarms for PC and Alarm Box

      - Record the time and date that gas flow was started for each chamber in the  TEC Gas System logbook

### Allow the chamber to flow Mixed Gas until there has been at least four exchanges of gas inside the chamber. DO NOT TURN ON HV UNTIL FOUR VOLUME EXCHANGES OF P10 GAS HAVE TAKEN PLACE IN THE TEC. The TEC volume is 6.0 cu meters, 6000 liters

### After 4 volume changes,  24,000 liters of P10, it is safe to turn on the high voltage of the TEC. HV to be turned on by TEC HV experts only.

## Gas System Procedures: In order to change a six-pack of  gas cylinders:

### Close the valve on the six-pack in use, either MV-E1A, MV-E1C, MV-E1E, MV-E1G, MV-E1J or MV-E1L.

### Open the valve to the full six-pack.

### Verify that the regulator is set to 40 psig.

### Record the starting cylinder pressure and the time and date in the gas logbook.

  - Shut off flammable gas flow

    - Disable the hardware and PC alarms

    - Make sure HV supplies are off

    - Set FM2 flow to Zero.

    - Close SV8(red light)

    - Close the Manual Valve MV-M10A on the Methane Gas Distribution Panel

    - Open Methane Supply Line Purging Valve to Vent Line

    - Confirm that PI7 pressure is Zero(Gas Rack)

    - Close Methane Supply Line Purging Valve to Vent Line

    - Set FM1 flow to 10LPM

    - Exchange 1 East TEC volume(10 hrs duration)

    - Close FI15(Gas Rack)

    - Close MV6(Gas Rack)

### Record the pressure remaining in the cylinders, and the time and date in the gas logbook.

### Notify the PHENIX shift leader that the TEC flammable gas has been purged.  If  PHENIX has purged all flammable gas the shift leader should notify the MCR of this fact.

  - Switch from argon purge to nitrogen purge. This is an option depending on inert gas supply.

    - Set FM1 flow to 6LPM

    - On Gas Pad connect Nitrogen to Argon Supply line

    - Close Argon

    - Open Nitrogen

    - Confirm that PI6 pressure is 20PSIG(Gas Rack)

    - Set FM1 flow to 10LPM(Electronic Rack)

    - Exchange 3  TEC gas volumes, 18000 liters prior to any long term Gas Shutdown.

    - Set FM1 flow to 6LPM(Electronic Rack)

  - Gas Shutdown

    - Set FM1 flow to Zero

    - Close Supply Line manual valve in Mixing House

    - Close inert gas supply on Gas Pad

    - Stop the gas system control program on the PC.

    - Log off the PC

    - Turn off the AC power for the PC.

    - Turn off the AC power for the alarm/interlock box

    - Turn off the AC power for Power supplies and Gas Rack

### HV System Procedures:  Turning on HV:

### If the HV is being turned on the first time, verify by checking with TEC Gas experts that operating gas has been flowing to the TEC for  a minimum of 4 volume turnovers . The TEC can be damaged if the HV is turned on without the chambers containing the proper gas mixture.

### Check that the appropriate current limits are in place for the power supply.  These limits are given in Attachment C.  The TEC HV experts shall maintain a HV logbook where the operating parameters of the HV settings are recorded.  This shall include the current limits, target voltages, ramp rates, operating voltages and currents, and trip tolerances.

### Go to the HV control  computer in the PHENIX Counting House.

### Check that the current limits, ramp rates and trip tolerances are set in accordance with the values in the TEC HV logbook.

### Check that the target voltage for each HV output line is appropriate (50 volts).  The first stage of bringing on the HV shall be a single increment in the ramp up.  This is because the current trips are disabled during ramping, and in order to locate a short in the system, it is necessary to halt the ramping and check the current at the earliest possible stage.

### Check that the ramp up rate for each HV supply is appropriate (50 volts per step).

### Ramp up the HV to 50V to test for electrical shorts.

### If HV current stays low,  < 1 microamp,  set initial voltages on computer control.

### Begin ramping Drift Window and Anode HV. Set Drift Window HV to logbook setting. Set Anode Voltages to 200V below logbook settings. Leave WC2 and BW off for now.

### Inspect current draw on HV computer screen.  If any channel has 5 microamps current draw after 5 minutes reduce the HV setting on that channel by 100V and wait for the current draw to drop. Any channels  exempted from the 5 microamp limit will be identified by the TEC HV experts and noted in the TEC HV logbook.

### Leave the HV at these settings for 1 hour.

### If current draw is stable after 1 hour  increase the anode voltage by 50 volts and repeat steps 6.6.8, 6.6.9 until reaching the HV values noted in the log book.

### If at any point during this procedure there is a HV trip determine the reason for the trip and return to step 6.6.3 for that channel. It may be necessary to reduce all 4 bias voltages on a particular TEC chamber prior to restarting the HV ramp-up procedure on that chamber.

### If  DW and Anode voltages are set to logbook values and  the current draw, especially on the anode wire , has been stable for a minimum of 1 hour begin setting the Wire Cathode 2 (WC2) and Back Window (BW) voltages to the logbook settings in 50V increments in a manner similar to steps 6.6.8 – 6.6.11

### If HV trips occur during normal TEC operation determine the cause of the trip and restore the tripped channel to 50V below its most recent operating value. Wait a sufficient amount of time to allow the current to stabilize. Restore the channel to its logbook value.  If the channel continues to trip at the logbook value reduce that channels HV operation setting by 50V and make a note in the TEC HV logbook.

### Once all TEC HV channels have been set to the appropriate TEC HV logbook values or their new modified values, monitor all channels for the operating current and HV trips. Keep a record of all TEC HV actions, especially unusual occurrences in the TEC HV logbook

### The TEC is ready for HV operation.

### HV System Procedures: Turning off High Voltage to a chamber:

### Begin ramping down the HV.

### Verify by the read back that the HV is off the system.

### In the event of irregularities, call a TEC HV expert.

### LV System Procedures: Turning on Low Voltage:

### Go to LV control computer in the PHENIX Counting House

### Go to PHENIX  Rack Control and Monitor page and turn on the appropriate TEC racks.

### Verify that the AC is on to the LV crate.

### Turn on one rack of LV supplies at a time powering the G-P and F supplies in the proper sequence as specified by the TEC experts. The sequence the TEC LV supplies are turned on is very important for the proper operation of the TEC electronics.

### Verify that the LV supplies have come on by monitoring  their status on the TEC LV control page. If the LV will not come on in some particular channel call a TEC expert.

### LV System Procedures: Turning off Low Voltage:

### Click on the appropriate button on the TEC LV computer page to turn LV power off to the channels required.

### Verify that the power has turned off by monitoring its status on the TEC LV computer page.

## References

## RHIC-OPM	3.0, "Local Emergency Plan for the Relativistic Heavy Ion Collider Project.”

## BNL ES&H Health Standard, Section 1.4.0, "Compressed Gas Cylinder Safety", December 18, 1991.

## BNL Occupational Health and Safety Guide (Interim), Section 4.11.0, "Installation of Flammable Gas Systems (Experimental & Temporary Installations)", June 21, 1989.

## Appendix A

## Attachment A - Call list for the TEC Gas, HV and LV System experts.

## Attachment B – TEC Gas system diagram.

## Attachment C – TEC HV settings and limits

  - Attachment D --TEC Gas System Acronym Glossary

  - Attachment E  --Log Sheet for Inert Gas Purge

  - Attachment F – TEC Gas System Check List

  - Attachment G – TEC Gas System Computer Alarm Conditions

  - Attachment H – TEC Gas System Hardware Alarm Conditions

  - Attachment  I – TEC P10 Operations Initial Checklist

  - Attachment J – TEC HV Log Sheet

Attachment A: TEC Gas System Experts

The following people have been trained to operate the PHENIX TEC gas system within the scope described in section 1.0 above. They have completed the prerequisite BNL training courses (see 3.1).

Leonid Kotchenda	x5795

Peter Kravtsov		x5795

Ed O’Brien		x4318

Carter Biggs		x7821	beeper 0227

Additional qualified users are to be listed below and posted in the

TEC gas system mixing house:

| Name | Phone | Approved by | Date |
| --- | --- | --- | --- |

Attachment A continued: TEC HV and LV System Experts

The following people have been trained to operate the PHENIX HV and LV  system. They have current training in the required BNL courses.

Ed O’Brien		x4318

Robert Pisani		x5301

Saskia Moiduszewski	x8428

Suzana Botelho		x2074

Additional qualified users are to be listed below and posted in the

TEC gas system mixing house:

| Name | Phone | Approved by | Date |
| --- | --- | --- | --- |

Attachment B: 		TEC Gas System Diagrams

Attachment B: 		TEC Gas System Diagrams

Attachment C:		TEC HV Settings

Each TEC plane has 4 separate channels of HV on the North and South side. The 4 bias voltages are Anode, Drift Window, Wire Cathode 2, and Back Window.  The nominal voltages and currents  for each chamber are:

Anode				+1600V

Drift Window 		-3000V

Wire Cathode 2		- 150V

Back Window		+ 220V

Current trip setting/channel	20 microamps

##### Ramping voltage steps	50 V

Note: For the actual voltage setting for each TEC HV channel refer to the TEC HV Logbook

Attachment  D .  TEC GAS SYSTEM ACRONYM GLOSSARY

BPCV	Back pressure control valve

CV	Check valve

F	Filter

FI	Flow indicator

FM	Mass flow controller

MV	Manual valve

PCV	Pressure control valve

PI	Pressure indicator

PS	Pressure switch

PSV	Pressure safety valve

PT	Pressure transmitter

SV	Solenoid valve

TT	Temperature transmitter

Attachment E :	Log Sheet for Inert Gas Purge

# Fill out and affix to TEC Gas Logbook

1._____	All TEC viewing windows in place

2.____ Gas supply lines intact from mixing rack to small racks and TEC  (both ends)

3.____ Gas Return lines intact from small racks  to vent line and to Mixing house Gas  Rack.

| Time/Date | Gas Flow Rate(FI3) | East TEC Pressure (PI8) | East TEC Pressure (PT9) | Supply Pressure (PI2) | Bypass Pressure (PT11) |
| --- | --- | --- | --- | --- | --- |

Operator __________________________

Attachment F:   PHENIX TEC GAS SYSTEM CHECK LIST   (To be filled out every 12 hours and placed in Gas System Binder)

Computer Monitored Sensors

| Sensor | Function | Value | Range | Comments |
| --- | --- | --- | --- | --- |
| PT1 | Leakage Pressure |  | 0.0mbar |  |
| PT2 | Input of Big Compressor |  | 1.0 - 2.0mmWC |  |
| PT5 | Input of East TEC pressure |  | 0.5 – 1.5”WC |  |
| PT6 | Input of West TEC |  | 0 .0mmWC |  |
| PT7 | West CO2 pressure |  | 0.0mmWC |  |
| PT8 | Output of West TEC |  | 0.0mmWC |  |
| PT9 | Output of East TEC |  | 0.5-2.0mmWC |  |
| PT11 | Return Bypass pressure |  | 2.0-6.0”WC |  |
| PT13 | East CO2 pressure |  | 0.0mmWC |  |
| PI2 | Output of Big Compressor |  | 35-45”WC |  |
| PI3 | Supply pressure |  | 2.0-6.0”WC |  |
| PI6 | Argon Delivery pressure |  | 15-21PSIG |  |
| PI7 | Methane Delivery pressure |  | 12-16PSIG |  |
| FI1 | Flow indicator |  | 9-11LPM |  |
| FM1 | Argon mass flow controller |  | 9.8-10.1LPM |  |
| FM2 | Methane mass flow controller |  | 0.98-1.1 LPM |  |
| H20 | Water analyzer |  | 0.0ppm |  |
| CH4 | Methane analyzer |  | 0.0% |  |
| O2 | Oxygen analyzer |  | 0.0ppm |  |
| FDI1 | Flow direction indicator East |  |  |  |
| FDI2 | Flow direction indicator East |  |  |  |
| FDI3 | Flow direction indicator East |  |  |  |
| FDI4 | Flow direction indicator East |  |  |  |
| TT1 | Temperature Sensor Input |  | 290-310K |  |
| TT2 | Temperature Sensor Output East |  | 290-310K |  |
| TT3 | Temperature Sensor Output West |  | 0.0 |  |
| PID | PID Controller |  | 0.0 |  |
| CH4/FM | Methane content |  | 9.8-10.1% |  |
| LEAK | Leakage |  | 0.0LPM |  |

Gas Rack

| PI1 | Input of Big Compressor |  | 1.0-2.0mmWC |  |
| --- | --- | --- | --- | --- |
| PI2 | Output of Big Compressor |  | 35-45”WC |  |
| PI3 | Supply pressure |  | 2.0-6.0”WC |  |
| PI6 | Argon Delivery pressure |  | 15-21PSIG |  |
| PI7 | Methane Delivery pressure |  | 12-16PSIG |  |
| FI15 | Purging flow for PI1 case |  | 1.0LPM |  |

Gas Pad

| Methane Bank |  |  | 1 or 2 |  |
| --- | --- | --- | --- | --- |
| Argon Bank |  |  | 0 - 2000 PSIG |  |

Electronic Rack

| Sensor | Function | Value | Range | Comments |
| --- | --- | --- | --- | --- |
| FM1 | Argon mass flow controller |  | 9.8-10.1LPM |  |
| FM2 | Methane mass flow controller |  | 0.98-1.1 LPM |  |
| FI1 | Flow indicator |  | 9-11LPM |  |

Mixing Room Manifold

| Sensor | Function | Value | Range | Comments |
| --- | --- | --- | --- | --- |
|  | Nitrogen Supply Pressure |  | 15 PSIG |  |
|  | Argon Delivery Pressure |  | 20 PSIG |  |
|  | Methane Delivery Pressure |  | 15 PSIG |  |

Operator__________________________		Date & Time__________________

Attachment G:  Computer Alarm Conditions

| Alarm Event | Problem | Level | Computer Action | Possible Causes |
| --- | --- | --- | --- | --- |
| CH4 /FM Low Alarm | Fresh gas methane content low | < 9.0% | Sound, Light | Methane bottle empty; FM2 set wrong |
| CH4 /FM High Alarm | Fresh gas methane content high | > 12.0 % | Off HV,SV6,SV8 | FM2 set wrong. |
| PT9 High Alarm | East TEC pressure too high | > 2.5 mmWC | Sound, Light | Makeup flow too high; No outlet flow |
| PT9 Low alarm | East TEC Pressure too low | < 1.0 mmWC | Sound, Light | Makeup flow stopped; Argon empty, Leak |
| PT2 Low alarm | Return line pressure too low | < 0.3mmWC | Off HV,SV6, Sound, Light | Big leak, Makeup flow stopped |
| PT2 Low Warning Alarm | Return line pressure low | <0.5mmWC | Sound, Light | Leak; Makeup gas flow low; |
| PT2 High Warning Alarm | Return line pressure high | >1.8mmWC | Sound, Light | Makeup gas flow high, Restricted outlet flow |
| PT2 High Alarm | Return line pressure too high | > 2.5mmWC | Off HV,SV8 On SV7 Sound, Light | Makeup gas flow too high, PCV1 fails |
| PT11 Low Alarm | Return Bypass pressure low | < 2”WC | Sound, Light | Leak, Makeup flow low. |
| PT5 High alarm | East TEC input pressure high | > 1.8”WC | Sound, Light | Input restriction to flow |
| PI2 Low Alarm | Makeup pressure low | < 20”WC | Sound, Light | PCV1 fails, Argon Bank empty |
| PI6 Low Alarm | Argon supply pressure low | <10PSIG | Sound, Light | Argon Bank empty, leak |
| PI7 Low Alarm | Methane supply pressure low | <7PSIG | Sound, Light | Methane Bank empty, leak |

Note:	High Warning = alert by the computer  before alarm condition                                      High Warning Return = alarm flag resets when sensor value returns to this number

Attachment H:   Hardware Alarm Conditions

| Alarm Event | Problem | Level | Action | Possible Causes |
| --- | --- | --- | --- | --- |
| PT9 High Alarm | East TEC pressure too high | > 2.5 mmWC | Sound, Light | Makeup flow too high; No outlet flow |
| PT9 Low alarm | East TEC Pressure too low | < 1.0 mmWC | Sound, Light | Makeup flow stopped; Argon bank empty, Leak |
| PT2 Low alarm | Return line pressure too low | < 0.3mmWC | Off HV,SV6, Sound, Light | Big leak, Makeup flow stopped |
| PT2 Low Warning Alarm | Return line pressure low | <0.5mmWC | Sound, Light | Leak; Makeup gas flow low; |
| PT2 High Warning Alarm | Return line pressure high | >1.8mmWC | Sound, Light | Makeup gas flow high, Restricted outlet flow |
| PT2 High Alarm | Return line pressure too high | > 2.5mmWC | Off HV,SV8 On SV7 Sound, Light | Makeup gas flow too high, PCV1 fails |
| PI2 Low Alarm | Makeup pressure low | < 20”WC | Sound, Light | PCV1 fails, Argon Bank empty |

Alarm = Bell and red flashing light in mixing house. There is also an alarm condition notification through slow controls (EPICS) to run control.

Attachment I.		P10 Operations Initial Checklist

MV1-closed(Dryer In)

MV2-closed(Dryer Out)

MV3-closed(Big Compressor Bypass)

MV4-open(Small Compressor Bypass)

MV5-closed(PID Controller N2 Supply Line)

MV6-closed(N2 Supply Line for  Purging)

MV8-open(Return Bypass Line)

MV10-closed(East TEC In)

MV11-closed(West TEC In)

SV1-closed(Purifier In)                                             SV14-closed(CH4 analyzer In)

SV2-closed(Dryer Out)                                             SV15-closed(Big Compressor In)

SV4-closed(Mixer Out)                                             SV16-open(East TEC In)

SV5-closed(Dryer Out)                                              SV17-open(East TEC Out)

SV6-open(Ar supply line)                                          SV18-open(West TEC In)

SV7-open(FM1 In)                                                     SV19-open(West TEC Out)

SV8-closed(FM2 In)                                                  SV20-closed(West CO2 Line )

SV9-closed(CH4 Supply Line)                                  SV21-status(East TEC Vent)

SV10-closed(Return bypass Line)                             SV22-open(West TEC Vent)

SV11-closed(Reference Mixture Line)                      SV23-closed(East CO2 Line)

SV12-closed(O2 analyzer In)

SV13-closed(H2O analyzer In)

FI3-closed(O2 analyzer In)                                                  FI11-closed(West TEC In)

FI4-closed(H20 analyzer In)                                        FI12-closed(East TEC In)

FI6-open(Ar supply Line)                                            FI13-closed(East DC In)

FI7-closed(CH4 supply Line)                                      FI14-closed(West CO2 Line)

FI10-closed(West CO2 Line)                                       FI15-closed(PI1 case)

Attachment  J.		TEC HV Log Sheet

Record all Voltage Variances and Enter New Values in the Table.

| TEC Sector/plane | Anode V(volts) | Anode I(mamps) | Drift Window V(volts) | Drift Window I(mamps) | Wire Cathode2 V(volts) | Wire Cathode2 I(mamps) | Back Window V(volts) | Back Window I(mamps) | Comments |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| E1/PL1 – S | -1600V |  | +3000V |  | +150V |  | -220V |  |  |
| E1/PL2 – S | -1600V |  | +3000V |  | +150V |  | -220V |  |  |
| E1/PL3 – S | -1600V |  | +3000V |  | +150V |  | -220V |  |  |
| E1/PL4 – S | -1600V |  | +3000V |  | +150V |  | -220V |  |  |
| E2/PL1 – S | -1600V |  | +3000V |  | +150V |  | -220V |  |  |
| E2/PL2 – S | -1600V |  | +3000V |  | +150V |  | -220V |  |  |
| E2/PL3 – S | -1600V |  | +3000V |  | +150V |  | -220V |  |  |
| E2/PL4 – S | -1600V |  | +3000V |  | +150V |  | -220V |  |  |
| E1/PL1 – N | -1600V |  | +3000V |  | +150V |  | -220V |  |  |
| E1/PL2 – N | -1600V |  | +3000V |  | +150V |  | -220V |  |  |
| E1/PL3 – N | -1600V |  | +3000V |  | +150V |  | -220V |  |  |
| E1/PL4 – N | -1600V |  | +3000V |  | +150V |  | -220V |  |  |
| E2/PL1 – N | -1600V |  | +3000V |  | +150V |  | -220V |  |  |
| E2/PL2 – N | -1600V |  | +3000V |  | +150V |  | -220V |  |  |
| E2/PL3 – N | -1600V |  | +3000V |  | +150V |  | -220V |  |  |
| E2/PL4 – N | -1600V |  | +3000V |  | +150V |  | -220V |  |  |

Operator__________________________		Date & Time__________________