MSFC PRACA : 2003-02-12 08:41

MSFC Problem Reporting and Corrective Action (PRACA) System

WHOLE RECORD REPORT( + ADDENDUM)

| MSFC Record # A16988 | In-Flight Anomaly Number -- | Contractor Report Number E-162 | JSC# -- | KSC# -- |
| --- | --- | --- | --- | --- |
| Problem Title LH2, ULLAGE PRESSURE TRANSDUCER OUTPUT VOLTAGE LOW |
| EICN# -- | ELEMENT ET | Contractor LMMSS | FSCM# -- | FCRIT 1R |
| HCRIT -- | Sys_Lvl -- | Misc Codes A B C D E F G H I J K L M N O |
| HARDWARE EIM | NOMENCLATURE -- | PART# -- | SER/LOT# -- | MANUFACTURER -- |
| HARDWARE LRU | NOMENCLATURE -- | PART# -- | SER/LOT# -- | MANUFACTURER -- |
| HARDWARE NCA | NOMENCLATURE H2 ULLAGE PRESSURE T | PART# 098-089 | SER/LOT# 1800 | MANUFACTURER GULTON-STATH |
| Test/Operation AI - ATI | Prevailing Condtion FC - CALIBRATION | F / U UC | Fail Mode EQ - OUTPUT ERROR | Cause U - UNKNOWN |
| System ELECTRICAL | Defect EP - ELVAL | Material CA - DEVICE | Work Contact J. ADAMS | Fail Date 12/06/1996 |
| Received at MSFC 12/10/1996 | Date Isolated 12/06/1996 | FMEA Reference 3.4.1.2 | IFA: Mission Phase -- | Mission Elapsed Time -- |
| Location GULTON-STA | Symptom EQ - OUTPUT ERROR | Time Cycle -- |
| Effectivity Text NONE |
| Vehicle Effectivity Codes |
| Vehicle 1 -- | Vehicle 2 -- | Vehicle 3 -- | Vehicle 4 -- | Vehicle 5 -- |
| Mission Effectivity Codes |
| Mssn 1 -- | Mssn 2 -- | Mssn 3 -- | Mssn 4 -- | Mssn 5 -- |
| Estimated Completion Dates |
| MSFC Approved Defer Until Date 06/18/1997 | Contractor Req Defer Until Date 02/18/1997 | LVL 3 Close -- | Remark / Action 02/18/1997 |
| Investigation / Resolution Summary |
| Last MSFC Update 02/15/2001 | CN RSLV SBMT 03/04/1997 | Defer Date 12/18/1996 | Add Date 12/10/1996 | R/C Codes 0 - EXPL -- -- |
| Assignee |
| Design J. RATLEY | Chief Engineer M. PESSINN | S & MA M. SMILES | Project -- | Project MGR P. COUNTS |
| Approval |
| Design J. RATLEY | Chief Engineer M. PESSIN | S & MA M.D. SMILES | Project -- | Project MGR P. COUNTS |
| PAC Assignee R. TUCKER | PAC Review Complete RT | MSFC Closure Date 03/11/1997 | Status C - CLOSED | F/A Completion -- |
| Problem Type -- | SEV -- | Program Name -- | REVL -- | OPRINC -- |
| FUNC MOD -- | Software Effectivity -- -- -- -- -- | Software Fail CD -- | SUBTYPE -- | Software Closure CD -- |
| RES PERSON L2 -- | Approval Signature L3 -- |
| Related Document Type -- | Related Document ID N008269 |
| Related Document Title NONE |
| Related Document Type -- | Related Document ID -- |
| Related Document Title -- |
| Related Document Type -- | Related Document ID -- |
| Related Document Title -- |
| Contractor Status Summary |
| Reliability/Quality Assurance Concerns, Recommendations: |
| Problem Description A TRANSDUCER FAILED THE CALIBRATION TEST SEGMENT OF THE SUPPLIERS ACCEPTANCE TEST PROCEDURE. THE TRANSDUCER'S OUTPUT VOLTAGE WAS LESS THAN THE MINIMUM ACCEPTANCE REQUIREMENT BETWEEN THE APPLIED PRESSURES OF 32 TO 34 PSIA CRITICALITY: THE PRESSURE TRANSDUCER IS LISTED AS CRIT. 1R: FMEA ITEM CODE 3.4.1.2; "FAILS WITH LOW READING" NOTE: THE TRANSDUCER IS LIMITED LIFE CONTROLLED * REVISION "B" - UPDATES AND CLOSES TASK I, II, AND III. SEE ASTERISKS |
| Contractor Investigation/Resolution GENERAL: AN LH2 ULLAGE PRESSURE TRANSDUCER WAS BEING TESTED AT THE SUPPLIER, GULTON-STATHAM TRANSDUCERS, PER ACCEPTANCE TEST PLAN (ATP) 3031-13803. THE TRANSDUCER FAILED THE CALIBRATION ACCURACY TEST OF ATP PARAGRAPH 5.2, SEQUENCE 3. PARTIAL TEST RESULTS ARE LISTED IN TABLE 1 BELOW TABLE 1 APPLIED PRESS VOLTAGE MINIMUM PSIA ASCENDING PRESS. DESCENDING PRESS. ACCEPTABLE VOLT 32 2.402 2.436 2.450 33 2.529 2.551 2.575 34 2.654 2.679 2.700 THE TRANSDUCER IS USED TO MEASURE THE ULLAGE PRESSURE IN THE LIQUID HYDROGEN PROPELLANT TANK DURING BOTH GROUND PROPELLANT LOADING AND IN FLIGHT. THE TRANSDUCER IS ATTACHED TO A SHOCK MOUNT ON THE FORWARD DOME OF THE HYDROGEN TANK, WITHIN THE INTERTANK. THE TRANSDUCER COVERS THE RANGE OF 12 TO 52 PSIA THE ORBITER PROVIDES 5 VDC TO THE TRANSDUCER AND RECEIVES A VOLTAGE FROM THE TRANSDUCER PROPORTIONAL TO THE ULLAGE PRESSURE THE PRESSURE MEASUREMENT IS USED TO CONTROL THE SPACE SHUTTLE MAIN ENGINE (SSME) GASEOUS HYDROGEN FLOW CONTROL VALVE TASK I. FAILURE INVESTIGATION A. THE TRANSDUCER FAILURE WAS DOCUMENTED ON NCD N-8269. THE NCD WILL BE DISPOSITIONED FOR FAILURE ANALYSIS TO BE PERFORMED AT GULTON RESPONSIBILITY: J. ADAMS/3741 - D. WESTPHAL/3740 L. COLON/4120 - W. SIGUR/4100 * COMPLETE: FEBRUARY 18, 1997 * CLOSURE STATEMENT FAILURE ANALYSIS OF THE TRANSDUCER ISOLATED THE FAILURE TO AN EXTREMELY SMALL GAS LEAK ON THE CIRCUMFERENTIAL WELD OF THE ANEROID CAPSULE. THE MEASURED LEAK RATE VARIED BETWEEN 2.2 (10-9 STANDARD CUBIC CENTIMETERS PER SECOND (CC/SEC) OF HELIUM AT 15 PSIG AND 5.4 (10-7 CC/SEC) AT 45 PSIG. THE LEAK WAS ONLY PRESENT WHEN THE PRESSURE DIFFERENTIAL WAS APPLIED WITH THE HIGHER PRESSURE INTERNAL TO THE ANEROID CAPSULE. TESTING ISOLATED THE LEAK TO AN APPROXIMATELY 3/8 INCH LENGTH OF THE WELD. OPTICAL EXAMINATION WITH UP TO 40 DIAMETERS MAGNIFICATION COULD NOT DETECT ANY VISIBLE FLAW. THE DELICATE NATURE OF THE ANEROID AND THE THIN METAL OF WHICH IT IS FABRICATED PREVENTED ANY FURTHER ATTEMPTS TO LOCATE THE WELD IMPERFECTION SINCE THE IMPERFECTION COULD NOT BE IDENTIFIED, DETERMINATION OF THE ROOT CAUSE OF THE LEAK WAS NOT POSSIBLE TASK CLOSED CAUSE: SUPPLIER ERROR - THE FAILURE WAS ASSOCIATED WITH A GAS LEAK AT A WELD ON THE INTERNAL ANEROID CAPSULE THE ROOT CAUSE OF THE GAS LEAK COULD NOT BE DETERMINED SEE TASK I TASK II. CORRECTIVE ACTION * THE ROOT CAUSE OF THE FAILURE COULD NOT BE DETERMINED. THIS PREVENTS THE DEVELOPMENT OF APPROPRIATE CORRECTIVE ACTION THE MANUFACTURING TRAVELER FOR THE WELD OF THE ANEROID CAPSULE WAS REVISED AS A PROCESS ENHANCEMENT WHICH SHOULD REDUCE THE LIKELIHOOD OF FUTURE SIMILAR FAILURES. REVISION 10 TO GULTON TRAVELER 121-601 ADDED STEP 065, "WELD SHOP GROUP LEADER INSPECT ALL HEAT SINKS AND REFURBISH AS REQUIRED, WITH THE EXCEPTION OF THE O.D. WELD HALF AND SOLID HEAT SINKS WHICH MUST BE REPLACED FOR EACH LOT." TASK CLOSED TASK III. CLEARANCE OF EFFECTIVITIES * THIS IS AN EXPLAINED CLOSURE OF THE CAPS ALL ETS CLEARED. THE TRANSDUCER FAILED DURING THE SUPPLIER ACCEPTANCE TESTS AND ALL DELIVERED TRANSDUCERS HAVE PASSED THE TESTS NOTE: THIS IS ALSO THE DEFERRAL RATIONALE TASK CLOSED TASK IV. CAPS CLOSURE SUMMARY * THE FAILURE OF THE TRANSDUCER TO MEET THE SUPPLIER ATP REQUIREMENTS FOR CALIBRATION ACCURACY WAS THE RESULT OF AN EXTREMELY SMALL GAS LEAK IN/ADJACENT TO THE CIRCUMFERENTIAL WELD OF THE INTERNAL ANEROID CAPSULE. IT WAS NOT POSSIBLE TO DETERMINE THE ROOT CAUSE OF THE DEFECT. A WELD PROCESS ENHANCEMENT WAS MADE TO REDUCE THE LIKELIHOOD OF FUTURE OCCURRENCES. THE SUPPLIER ACCEPTANCE TEST PLAN (ATP) AND THE TESTS PERFORMED ON THE TRANSDUCERS AT BOTH MAF AND KSC AFTER INSTALLATION ON AN ET ARE ALL EFFECTIVE SCREENS FOR THIS TYPE OF PROBLEM TASK CLOSED * THIS AN EXPLAINED CAPS * THIS CAPS IS CLOSED. NO FURTHER ACTION IS REQUIRED |
| MSFC Response/Concurrence |

MSFC Problem Reporting and Corrective Action (PRACA) System

ASSESSMENT ADDENDUM REPORT

| MSFC Report# A16988 | IFA# -- | Contractor RPT# E-162 | JSC# -- | KSC# -- | EICN# -- |
| --- | --- | --- | --- | --- | --- |
| Asmnt Part# PD740098-089 | Asmnt Part Name LH2,ULLGE PRESS TRAN | Asmnt Serial/Lot# 1800 |
| HCRIT CD -- | FCRIT CD 1R | CAUSE CD U - UNKNOWN | FAIL MODE EQ - OUTPUT ERROR |
| Asmnt FMEA 3.4.1.2 | Asmnt FM -- | FMEA CSE -- | FMEA SCSE -- |
| Asmnt FMEA -- | Asmnt FM -- | FMEA CSE -- | FMEA SCSE -- |
| Asmnt FMEA -- | Asmnt FM -- | FMEA CSE -- | FMEA SCSE -- |
| Correlated Part# -- | Correlated Part# -- | Correlated Part# -- |
| Associated LRU# -- | Associated LRU# -- | Associated LRU# -- |
| MAJOR DESIGN CHANGES |
| APRV DATE -- | DESCRIPTION OF CHANGES -- |
| ASSESSMENT TEXT |