Results of vibration versus water-flow measurements at BL7.2 on Tuesday, January 21, 2004:

# Tom Scarvie

The vibration measurements were performed using seismic accelerometers (model 731A) and power amplifiers (model P31) from Wilcoxon Research connected to a spectrum analyzer (HP-3563A).  The analysis was then performed using Matlab. The bandwidth and noise specifications for the accelerometers, as provided by the manufacturer, are the following (see Wilcoxon Research Accelerometer manual for more details):

# Bandwidth:			0.1 – 300 Hz

Noise Power Spectrum:	0.03 g/(Hz at 2 Hz

0.01 g/(Hz at 10 Hz

0.004 g/(Hz at 100 Hz, where g is the gravitational acceleration

1% coupling between horizontal and vertical motion (not including mounting).

The accelerometers were mounted on a “T” aluminum block that was securely clamped to the alignment monument fixtures on top of the pinhole array flange and to the bolt heads of the visible light port flange. The horizontal accelerometer in measurements 1. and 2. was parallel to the beam axis; measurement 3. is a comparison of the parallel horizontal orientation to the one perpendicular to the beam axis. It should be noted that these measurements are of the motion of the housings of the elements, and not the elements themselves. Since the elements are securely mounted to the housings, the data should give us a good picture of the magnitude of the motion, and a very good relative measurement of vibration for different flow settings.

The frequency range of interest for these measurements was from the lowest possible frequencies of about 0.3 Hz (somewhere between 0.3 and 1.0 Hz the noise floor of the accelerometers begins to dominate) up through to about 400 Hz. In order to obtain higher resolution power spectrum data, three data sets were taken for each measurement---one from 0-20 Hz, one from 0-100 Hz, and one from 0-400 Hz. For all the data a Hanning window was used and 20 power spectrums were averaged for each measurement (inside the HP-analyzer), to reduce the variance of the estimated spectrum. The HP analyzer uses a 2048 point FFT (1024 points are kept in the power spectrum). Anti-alias filtering is already performed inside the analyzer.

The spectrum data read from the HP-analyzer is then post-processed inside Matlab. Integration of the acceleration power spectral density provides displacement information and ultimately a cumulative rms-amplitude (Parseval’s theorem). We usually integrate the spectra from high to low frequency. Each figure contains the power spectral density as the upper plot and the cumulative rms-amplitude of the vibration as the lower plot. Sensitivity of the accelerometers was set to 1000 V/g for all measurements, and the low pass filter in the amplifier of the accelerometers was set to 450 Hz.

The raw measurement data is available and can be sent on request (contact Tom Scarvie) or downloaded from \\Als-filer\physdata\matlab\srdata\BL72\vibration_vs_h2oflow.

Let’s just skip to the interesting part first and look at the plots that really matter:

1. Below is a comparison of the vibrations of the pinhole housing for “high” H2O flows of 1.5GPM to the pinholes and 3GPM to the optical mirror, “low” flow of 0.6GPM pinhole, 2.0GPM mirror, and no-flow (0GPM to both).

As one might expect, the vibrations got smaller as the flow was reduced. However, even in the high flow case, the RMS displacement is only about 1 µm at 1 Hz. The shutter is operated at 1/60th sec, where the vibrational motion is only in the 10s of nanometers. There was a noticeable improvement when the flow was reduced, so even if we don’t see a practical improvement in the image, at least we’re “living clean”, with the flows set as low as possible while still protecting the optical elements. The flows are currently set to the low-flow values.

Now let’s look at some other interesting data:

2. Comparison of vibrations of visible light port and pinhole assembly with “high” H2O flows of 3GPM for the visible port mirror, and 1.5GPM for the pinhole array:

There are different vibrational modes, as one might expect, but the magnitudes are roughly the same and small. At 1Hz, we have less than 1 µm horizontal motion, and 1-2 µms vertically. Above 10Hz, both are below 0.2 µms.

3. Comparison of horizontal pinhole motion parallel to the beam axis and perpendicular to it for the low-flow settings.

Different modes of vibration again, but still the amplitudes are very small.