July 22, 2002

Item:  	SSCOR Incorporated, Suction Pumps, Models 2314, 2314B and 2315, all units.

Specific Incident:	These pumps are designed for short-term suction, primarily during resuscitation efforts, yet one medical center used them for extended procedures, such as wound drainage and endoscopic procedures, resulting in pump failures. Use of these pumps for extended high vacuum/high occlusion can lead to overheating and premature pump failure, making the equipment unavailable for use and limiting needed suction.

Recommendations:	1.Check the medical equipment inventory to determine if you have the specified pumps.  Model numbers for the SSCOR suction units can be found on a silver label on the back of the units.  Ensure any units on hand are being used as intended, and not placed in areas where they can be used for long term or continuous suction.

2. Refer to the revised page 3 of the operator and service manuals (attached), copy and insert the revised pages in all SSCOR 2314, 2314B and 2315 operator and service manuals.  If manuals are not available they can be obtained from the manufacturer or from the Center for Engineering and Occupational Safety & Health (CEOSH) at (314) 543-6700 or at http://vaww.ceosh.med.va.gov.

3. Label pumps to reflect intended use: “CAUTION: Not for continuous suction; for short term use only.”

4. Train users on the intended use of these pumps and the possible risk associated with use beyond the equipment’s design.  If necessary, purchase appropriate suction systems for wound or procedural drainage

Source:		CEOSH and manufacturer

Contact:	For additional information, please contact Mr. Paul Sherman at CEOSH (314) 543 6700.

URGENT FIELD CORRECTION

July 1, 2002

Subject:	Field Correction Notice for S-SCORT new-Duet and new-Sentinel Suction Units sold exclusively through Armstrong Medical Industries.

To Whom It May Concern:

SSCOR needs your prompt cooperation in carrying out an urgent field correction.  Based on reports from the field, our company has become aware of instances where some of the suction units referenced below have been used in medical procedures for which they were not designed or intended.  In evaluating these reports we have determined it would be appropriate to make the intended use of our devices even more specific.  These events have been reported to United States Food and Drug Administration (FDA), and FDA is aware of our corrective action.

The corrective action specifically affects all SSCOR suction units sold through Armstrong Medical Industries with the following model numbers:

SSCOR Suction Unit Model 2314 (AE 6975 or 6975S)

SSCOR Suction Unit Model 2314B (AE 6976or 6976S)

SSCOR Suction Unit Model 2315 (AE 6905)

Model numbers for these SSCOR suction units can be found on a silver label on the back of the units.

Because of what we have learned, we are adding the following cautionary language to page 3 of the General Description portion of the Operating Instructions and Maintenance Manual:

"CAUTION:  S-SCORT suction units are designed for hospital crash cart suction, patient transport and emergency medical service.  They are not designed or intended for use in extended procedures that require prolonged high vacuum/low airflow applications, as is the case in wound drainage or endoscopic use or in any other procedure that produces high vacuum levels within an occluded system for an extended period of time.  Turn the S-SCORT suction unit off when it is not in use."

Attached to this letter is a revised page 3 for your Operating Instructions and Maintenance Manual.  Please make certain that a copy of this new page is inserted into the Manual for each suction unit.  If necessary, SSCOR will provide you with additional copies of the revised manual page.  In addition, we ask that you make the revised portion of the Manual part of a notice alerting all users of these devices in your facility to the appropriate intended use for our products.

If you have any questions regarding the contents of this letter or any questions as to how to carry out this field correction, please contact our Customer Service at 1-800-434-5211.

We appreciate your full cooperation in this matter.

Sincerely Yours,

Scott Eamer

Product Manager

SSCOR, Inc.

WWW.SSCOR.com

Table of Contents

S-SCORT newDuet Model 2314

| Caution-Notice | 2 |
| --- | --- |
| General Description and Symbols | 3 |
| Getting Acquainted- Basic Instructions | 4 |
| General Specifications-Trouble Shooting | 5 |
| Warranty-Maintenance | 6 |
| Internal Access-Internal Components | 7 |
| Conformity | 8 |

General Description

S-SCORT newDuet Model 2314

S-SCORT newDuet is a portable, 100V-240V AC constant suction device, with a

12V DC battery back-up.  The battery is charged by an internal, DC Dual Mode battery charger. A fully charged battery will power the unit for 45 minutes (+-10%).  Battery condition is automatically monitored and visualized by indicator lights on the control panel.  AC charging should bring the battery to a dependable working charge in 6 to 8 hours.

Suction power can be regulated when full power may be considered harmful to the patient.

CAUTION:  S-SCORT suction units are designed for hospital crash cart suction, patient transport and emergency medical service.  They are not designed or intended for use in extended procedures that require prolonged high vacuum/low airflow applications, as is the case in wound drainage or endoscopic use or in any other procedure that produces high vacuum levels within an occluded system for an extended period of time.  Turn the S-SCORT suction unit off when it is not in use.

Description of Symbols

S-SCORT newDuet Model 2314

Symbol                                Location				Meaning

Table of Contents

S-SCORT new-Sentinel® Model 2315

| Caution-Notice | 2 |
| --- | --- |
| General Description and Symbols | 3 |
| Getting Acquainted- Basic Instructions | 4 |
| General Specifications-Trouble Shooting | 5 |
| Warranty-Maintenance | 6 |
| Internal Access-Internal Components | 7 |
| Conformity | 8 |

General Description

S-SCORT new-Sentinel® Model 2315

The S-SCORT new-Sentinel™ is a portable, 12V DC battery operated suction pump.  It is powered by sealed lead acid batteries, capable of driving the unit at full power for 30 minutes ( 10%.  The batteries are recharged by a 115V AC to 13.8V DC Dual Mode Automatic Charger (included with the unit).

CAUTION:  S-SCORT suction units are designed for hospital crash cart suction, patient transport and emergency medical service.  They are not designed or intended for use in extended procedures that require prolonged high vacuum/low airflow applications, as is the case in wound drainage or endoscopic use or in any other procedure that produces high vacuum levels within an occluded system for an extended period of time.  Turn the S-SCORT suction unit off when it is not in use.

The new-Sentinel™ is designed to provide instant, effective suctioning, independent of external sources of power and can be pre-set to be activated immediately upon reaching the distressed patient.

Suction power is controllable for those instances when full power would be considered harmful to the patient.

The S-SCORT new-Sentinel™ is equipped with a disposable collection canister which features a bacterial filter to screen airborne particulates and a mechanical shut-off valve to prevent fluid overflow.

Description of Symbols

S-SCORT new-Sentinel® Model 2315

# Symbol                                Meaning				Location