NASA GLENN SPECIFICATION FOR AN X-RAY DIFFRACTION SYSTEM, PR# 199480

Proposed systems shall contain the following components and meet or exceed the given specifications. Vendors shall provide descriptive literature in the form of sales brochures and system information verifying existing capabilities.

- X-RAY - ENCLOSURE:

- The system x-ray enclosure shall have a large opening providing easy access for sample loading and goniometer configuration changes.

- The system shall have large leaded Plexiglas doors for optimum visibility and accessibility.

- The system shall have full radiation protection.

- The system shall have two independent security circuits.

- The system shall have x-ray radiation warning lights.

- The system when operational shall pass a radiation leakage test performed by the NASA Radiation

Safety Committee to verify radiation protection.

- GONIOMETER:

The goniometer shall be a vertical theta/theta design with the following capabilities:

- The controlling mechanism shall have both stepper motors and optical encoders for optimum scanning speed and positioning precision.

- The minimum step size shall be 0.0001 degree.

- The reproducibility of the system shall be plus or minus 0.0001 degree.

- The goniometer shall have high precision dovetail tracks for variable positioning of all components.

- Shall accommodate scintillation, solid-state and position-sensitive detectors with quick and easy exchange without realignment.

- Shall accommodate monochromators for Cr, Fe, Co, Cu, and Mo, radiation.

- 2( arm for detector shall allow for a variable sample to detector distance.

- Collimators and beam stop for use with K( filter or monochromator.

- X-RAY GENERATOR:

A.	X-ray generator shall be of solid-state design.

B.	X-ray generator shall have a continuous output power of 3.0 kW minimum.

C.	X-ray generator shall have a voltage range of 10 to 60 kV, tube current of 2 to 60 mA, with a stability of tube voltage and current of less than 0.01 percent with 10 percent line fluctuation.

D.	Computer controlled or manual operation.

- Permits remote software control of tube current and voltage.

- The kV and mA settings must be continuously adjustable, with digital readout of measured units.

- Transfers generator status and error messages to PC.

- Permits selection and locking of generator operating parameters and limit values.

- Permits current or default values to be used for tube voltage and current.

- Has diagnostics and operation time counter.

- The X-ray generator shall include a microprocessor with automatic startup and burn-in cycle for new X-ray tubes.

- X-RAY TUBE:

- A 2.2 kW Cu long fine focus tube (0.4 x 12 mm) that shall allow the tube to be rotated for point source without realignment.

- SAMPLE STAGE:

- All sample stages shall have a pre-aligned interface to allow interchange without the need for realignment.

- As a minimum the system shall have a multi-purpose sample stage with capability for a minimum of 20 mm sample height and method for rapid alignment.

- Sample stage shall be capable of future installation of a 1200oC sample chamber with ability to change in less than one hour.

VI.    INCIDENT BEAM OPTICS SHALL INCLUDE:

A.   A software controlled variable divergence slit.

- Appropriate Soller slits to limit the lateral divergence of the beam (along the goniometer axis) on the sample.

- Incident beam anti scatter slit to improve peak to background ratio.

VII.  DIFFRACTED BEAM OPTICS SHALL INCLUDE:

- A beam knife to improve peak to background ratio for 2 theta values < 10 degrees.

- Appropriate Soller slit with a slot available for placement of a K-Beta filter or manual beam attenuator.

- A flat graphite diffracted beam monochromator to reduce the effect of fluorescent radiation, also used with a higher speed detector system, shall also use optics with or without monochromator and not require realignment.

VIII.  DETECTOR SHALL BE CAPABLE OF:

- Possessing a detector system that significantly performs faster measurements. Measurements that are a minimum of 15 times faster when compared to conventional electronics using a proportional counter, scintillating counter or equivalent, while maintaining or improving signal to noise characteristics and peak resolution.

IX.   SOFTWARE, THE SYSTEM SOFTWARE SHALL INCLUDE THE FOLLOWING:

- The operating system shall be Microsoft®Windows 2000.

- Microsoft®Office 2000 or better, including Microsoft Access.

- Norton Systemworks.

- The system operation software shall be able to perform the following

  - Handle and maintain data collection and storage

  - Contain software for alignment and system performance evaluation

  - Search-match algorithm with peak identification and peak profile determination, including automatic phase identification

  - ICDD reference database, including user license and four (4) years of annual updates

  - Routines for crystallite size and micro-strain evaluation

  - Indexing of powder patterns

  - Structure determinations

  - Stress analysis software

  - Analytical software for crystallographic and Rietveld analysis

  - Software that can implement a batch program, optimize calibration based on standards, customize the user interface, and allows customized presentation of results

X.    COMPUTER SYSTEM AND PRINTER AS A MINIMUM SHALL BE:

- A 2.0 GHz Pentium 4 processor, 512 MB DRAM, 80 GB hard drive, 3Com ethernet card 100 Mbps, Iomega 250 MB Zip drive with five (5) 250 MB disks, CD ROM with 40x read speed, CD ROM recordable/rewritable/read with m 20x/10x/40x speeds, 18.1 inch LCD flat panel/screen monitor, HP color ink jet printer with cable.

XI.   OTHER INCLUSIONS THAT SHALL BE MADE A PART OF THIS PURCHASE:

- Shipping and delivery of entire system, including any and all optional equipment and accessories.

  - Installation and set up of entire system to operational status.

  - Provide pre-installation guide and milestone with bid detailing our facility requirements.

  - After on-site operational status is achieved, a demonstration is required to show that the XRD is operating and providing data that meets the purchased specifications

B.    System shall include a two (2) year warranty for parts, labor, emergency and preventive maintenance.

C.    Documentation shall include but not be limited to the following:

1.    Two (2) complete sets of printed manuals and tutorials for all system software.

2.    Data file structure to be provided.

3.    On-line Windows Help for system software.

4.    Two (2) complete sets of printed manuals for the system hardware

D.     Training including travel, the following training shall be included in the bid price:

  - Minimum of two (2) customer/user training/class/instruction at manufacturer’s site, preferably prior to system delivery to customer.

  - Minimum of two (2) customer/user training/class/instruction on-site, 4 days.

XII.  OPTIONS SHALL BE AVAILABLE BUT ARE NOT LIMITED TO THE FOLLOWING:

- Monocapillary collimator that converges and collimates beam size to at least 0.3mm, includes microscope to    view the area of excitation.

- A high temperature stage capable of 1200oC operation, including automated software control of temperature and vacuum system.

- Texture cradle to perform pole figure analysis, including any additional required hardware and software.