Goal

A list of unresolved issues and a compilation of test concepts that could potentially address these issues.

Table 6

The following table lists unresolved issues and test concepts that may need to be developed to address these critical parameters/failure modes.  These comments were provided by participants in response to the work assignment and were discussed during the workshop.  Comments generated at the workshop are highlighted in the table.

Table 6 Additional Testing Needs

| Unresolved issue | Test concept |
| --- | --- |
| Metal fractures, fabric erosion, wear holes |
| Mechanical failures (metal fracture, fabric erosion). | Incorporate design-dependent effects of angulation and tortuosity into tests. May not need new test concepts, just new protocols that recognize these effects (on both loads and device response). |
| Evaluation of effects of predicted failure locations (fixation modes), for example barbs or support bars. | Workshop comments: Same as for number 6. |
| Graft wear holes | New test that simulates movement between the graft material and wireform in a simulated AAA model. Also to be assessed is the functional significance of any wear holes observed. In other words, if wear holes are generated, what is their impact on WEP, IWP, etc.? Given the porosity of Dacron fabrics, it is not unrealistic to expect that wear holes below a certain size will have minimal adverse impact on the device’s ability to exclude an aneurysm (i.e., there is probably some tolerance around allowable minimal type III endoleak rate). Workshop comments: This is relevant to all graft materials, not just textiles. There is a need to model the attachment sites and the part of the graft that is in the sac, and at interaction points (junctions), in an appropriate model. |
| Clinically-significant hole size | Workshop comments: Captured in number 3. Characterization of hole size (any hole), and holes identified in number 3 is important. Need to consider distortion in weave as a sign of hole development. |
| Stent fracture | Transition zones & flexion site forces. Workshop comments: See earlier discussion. |
| Effects of Fractured Stent Element(s) on Mechanical Integrity | Evaluate the impact of stent-wire fracture(s) on the mechanical attributes of the deployed / implanted endoprosthesis (e.g., hoop strength, migration resistance, sealing efficiency, etc.) to answer the question- Does stent-wire fracture necessarily constitute a device failure? The location and number of fractures could be identified from the high strain regions identified by the FEA analysis, anticipated to be ‘weak-spots’. If it could be demonstrated that the mechanical attributes of an endoprosthesis are not significantly compromised by wire fracture(s), this information could be a strong mitigator of our interpretation of the mechanical significance of fatigue fracture of the stent. Workshop comments: This is a form of prospective failure analysis. The question is: is this device failure amenable to corrective intervention? The presence of hooks or barbs represents a new source of metal-against-fabric wear. A risk analysis for failure modes and the potential for failure of components should be conducted. The potential need for repair versus surgical conversion should be considered. |
| Testing of device to vessel wall attachment components such as barbs. May also apply to device to device attachment components for multi-component devices. | New test or new vessel model that accounts for simulation of actions upon the attachment component of the device. Workshop comments: It may be possible to group this with number 5. This concern is covered by other tests, e.g., pull tests for modular components, tensile tests, durability test. In response to a question of the definition of a hook versus a barb, one definition was that a hook is curved and a barb is straight. For all components of the device, there is a need to evaluate the loads and justify why the components (e.g., stents, anchoring systems) will not fail. |
| Effects of Fractured Stent Element(s) on Graft Integrity | Evaluate the impact of stent-wire fracture(s) on the integrity of the deployed / implanted endoprosthesis (e.g., Bond/peel strength, delamination propagation, WEP, IWP, etc.) to answer the question- Does stent-wire fracture necessarily constitute a device failure? The location and number of fractures could be identified from the high strain regions identified by the FEA analysis, anticipated to be ‘weak-spots’. If it could be demonstrated that the graft integrity of an endoprosthesis is not significantly compromised by wire fracture(s), this information could be a strong mitigator of our interpretation of the mechanical significance of fatigue fracture of the stent. Workshop comments: Same comments as for number 6. |
| Effects of Fractured Stent Element(s) on Biological Response | Evaluate the impact of stent-wire fracture(s) on the biological response to a deployed / implanted endoprosthesis (e.g., fistula formation, vessel wall erosion, failure propagation by redistribution of loads, inflammatory response, etc.) to answer the question- Does stent-wire fracture necessarily constitute a device failure? The location and number of fractures could be identified from the high strain regions identified by the FEA analysis, anticipated to be ‘weak-spots’. If it could be demonstrated that the biological response of an endoprosthesis is not significantly compromised by wire fracture(s), this information could be a strong mitigator of our interpretation of the mechanical significance of fatigue fracture of the stent. Workshop comments: Same comments as for number 6. |
| Device damage | Test forced dilation and over-dilation in narrow segments. Workshop comments: Need to evaluate tolerance to balloon dilatation due to re-intervention or treatment of occlusions. This includes more than stent and suture breaks, that is, it includes any and all device damage. |
| Separation of Modular Components |
| Modular component joints. | Evaluate effects of remodeling and other anatomic changes (longitudinal shrinkage, changes in angulation and tortuosity). These changes may produce tension, compression, or bending loads on joints. Workshop comments: Qualitative assessment of modular separations under worst-case conditions (i.e., extreme angulation and tortuosity) is needed. Possible bending or kinking of the device due to changes post implant should be assessed. |
| Axial dislocations, separations. | New test that simulates axial loads on the implant– if determined to be significant. These may be found to have several causes such as vessel motion and hemodynamic components (shear). |
| Separation of overlapping component in supported and unsupported region. | Unidirectional flow in durability tester. |
| Reproduction of kinking of grafts, suture break, endoleak, endotension. | Reproduction of anatomical conditions including development of synthetic material to resemble natural artery. Workshop comments: Delete based upon discussion. Covered under other tests. |
| Aneurysm remodeling effects |
| AAA remodeling | A relevant, realistic test does not exist; perhaps only clinical trials with good long-term follow-up will enable understanding of device response. |
| Iliac leg kinking. | Evaluate aneurysm longitudinal shrinkage effects on implant. Workshop comments: Delete based upon discussion. |
| How the grafts null adapt to morphologic changes – i.e. fatigue testing in multiple or changing conditions. |  |
| Stent/device _________ | Angles & flexion sites with changing morphology. Workshop comments: Delete based upon discussion. |
| Failure resulting from changes in aneurysm over time | New type of durability test Workshop comments: Delete based upon discussion. |
| Fixation post-neck dilation | Wall contact evaluation in growing neck model at various minimum and maximum oversizings and engagement lengths (can use tensile loads as well for migration resistance). Workshop comments: Delete based upon discussion. |
| The effect of the biologic milieu on the materials (corrosion, etc). | Workshop comments: Delete based upon discussion. |
| Other |
| Endotension / aneurysm sac reduction | Flow simulation of active versus passive diffusion. Workshop comments: Delete based upon discussion. |
| Type IV endoleak | Determine integral water leakage specification using pure water that correlates to no Type IV endoleaks in vivo. |
| Proximal seal in short neck | Pulsatile dye penetration in progressively angled necks. Workshop comments: Delete based upon discussion. Hopkinson’s testing is a useful reference. Covered under vessel wall attachment. |
| Migration | Unidirectional flow model. |
| Columnar Strength / Buckle Resistance | Harris (1999) showed that aneurysmal shrinkage leading to endoprosthesis buckling is a significant source of early and late-term complications. A number of authors (Carpenter, 20001; Resch, 2000; Kalliafas, 2000) have demonstrated that kinking, buckling, or twisting of the deployed endoprosthesis subsequent to device migration, hemodynamic forces, or aneurysmal shrinkage constitutes a significant failure mode necessitating reintervention. Characterization of the columnar strength /buckle resistance of the deployed device to in response to anticipated hemodynamic forces and/or longitudinal displacements may be a way to gain understanding with this clinical problem. Workshop comments: Numbers 26, 27 and 28 are of academic interest. |
| Revision Methodologies | Given the relatively high rate (26.6%) of reintervention reported for AAA endoprostheses (Holzenbien, 2001), the compatibility of the deployed endoprostheses with convention revision techniques (e.g., adjunctive stenting, coil embolization, balloon dilatation, etc.) should be established empirically or through rationales. Workshop comments: Numbers 26, 27 and 28 are of academic interest. |
| Rupture failure mode consequence | Are the consequences of aneurysm rupture as severe when an endovascular prosthesis is implanted (containing the rupture) as when a rupture occurs without the device? Workshop comments: Numbers 26, 27 and 28 are of academic interest. This may be relevant in defining successful clinical outcomes. |
| Aneurysm size and shape | Workshop comments: A standard size and shape of aneurysm is difficult to define for in vitro testing purposes, yet can have impact on test relevance and outcome. |

## References / Abstracts

Carpenter (2001)

Harris (1999)

Holzenbien (2001)

Kalliafas (2000)

Resch (2000)