WHOLE RECORD REPORT( + ADDENDUM)

| MSFC Record # A11809 | In-Flight Anomaly Number -- | Contractor Report Number E-118-1 | JSC# -- | KSC# -- |
| --- | --- | --- | --- | --- |
| Problem Title LOX PRESSURE TRANSDUCER FAILED THE CALIBRATION TEST |
| EICN# -- | ELEMENT ET | Contractor MMMSS | FSCM# -- | FCRIT 1R |
| HCRIT 1R | Sys_Lvl N | Misc Codes A B C D E F G H I J K L M N O |
| HARDWARE EIM | NOMENCLATURE LOX PRES TRANSDCR | PART# 80931003730-010 | SER/LOT# 260 | MANUFACTURER MMC |
| HARDWARE LRU | NOMENCLATURE N/A | PART# N/A | SER/LOT# N/A | MANUFACTURER N/A |
| HARDWARE NCA | NOMENCLATURE LOX PRES TRANSDCR | PART# PD7400098-079 | SER/LOT# 260 | MANUFACTURER GULTON |
| Test/Operation A - ATP | Prevailing Condtion F - FUNCTIONAL | F / U F | Fail Mode EG - SIG HI OR LO | Cause ETW - EI-TEST-WORK |
| System ELECTRICAL | Defect XN - NA | Material C - EEE | Work Contact J. ADAMS | Fail Date 10/26/1988 |
| Received at MSFC 10/27/1988 | Date Isolated -- | FMEA Reference 3.2.13.1 | IFA: Mission Phase -- | Mission Elapsed Time -- |
| Location MAF | Symptom ET - MEAS ANOMALY | Time Cycle -- |
| Effectivity Text NONE |
| Vehicle Effectivity Codes |
| Vehicle 1 -- | Vehicle 2 -- | Vehicle 3 -- | Vehicle 4 -- | Vehicle 5 -- |
| Mission Effectivity Codes |
| Mssn 1 -- | Mssn 2 -- | Mssn 3 -- | Mssn 4 -- | Mssn 5 -- |
| Estimated Completion Dates |
| MSFC Approved Defer Until Date -- | Contractor Req Defer Until Date -- | LVL 3 Close -- | Remark / Action -- |
| Investigation / Resolution Summary |
| Last MSFC Update 02/09/1995 | CN RSLV SBMT 12/09/1988 | Defer Date -- | Add Date 10/28/1988 | R/C Codes 3 - F/TE -- -- |
| Assignee |
| Design A. JACKMAN | Chief Engineer J. NICHOLS | S & MA R. JACKSON | Project M. PESSIN | Project MGR P. BRIDWELL |
| Approval |
| Design A. JACKMAN | Chief Engineer J. NICHOLS | S & MA R. JACKSON | Project M. PESSIN | Project MGR J. SMELSER |
| PAC Assignee J.EL-IBRAHIM | PAC Review Complete JE | MSFC Closure Date 01/30/1989 | Status C - CLOSED | F/A Completion -- |
| Problem Type -- | SEV -- | Program Name -- | REVL -- | OPRINC -- |
| FUNC MOD -- | Software Effectivity -- -- -- -- -- | Software Fail CD -- | SUBTYPE -- | Software Closure CD -- |
| RES PERSON L2 -- | Approval Signature L3 -- |
| Related Document Type -- | Related Document ID -- |
| Related Document Title -- |
| Related Document Type -- | Related Document ID -- |
| Related Document Title -- |
| Related Document Type -- | Related Document ID -- |
| Related Document Title -- |
| Contractor Status Summary |
| Reliability/Quality Assurance Concerns, Recommendations: |
| Problem Description FOUR LOX ULLAGE PRESSURE TRANSDUCERS PROVIDED TO KSC FOR INSTALLATION ON LWT-29 WERE NOT PROPERLY TESTED PRIOR TO SHIPMENT. TEST PROCEDURE 7R101-FT, WHICH CHECKS THE ACCURACY OF THE TRANSDUCER OUTPUT VOLTAGE VERSUS THE APPLIED PRESSURE, WAS IMPROPERLY PERFORMED DUE TO PERSONNEL ERROR. THE PROBLEM WAS DISCOVERED DURING A REVIEW OF THE MANUFACTURING HISTORY OF ONE OF THE TRANSDUCERS WHICH FAILED A POST-INSTALLATION CONFIDENCE TEST ON LWT-29. (SEE CAPS E-117). REF. MARS T-102801 |
| Contractor Investigation/Resolution CAUSE: PERSONNEL ERROR R/C: TESTING TOOL HAS BEEN DEVELOPED TO REDUCE THE NUMBER OF OCCURRENCES THIS PROBLEM REPORT HAS BEEN DEFERRED FOR STS-27 BASED ON THE FOLLOWING RATIONALE: FAILURE INVESTIGATION HAS DETERMINED THE CAUSE TO BE AN IMPROPER APPLI- CATION OF POWER TO THE TRANSDUCER DURING OR PRIOR TO, TP-7R101 NOISE TEST PERFORMED AT MAF. THE INSTALLED TRANSDUCERS ON STS-27 WERE NOT EXPOSTED TO THE SAME TEST PROCEDURE. THEREFORE, THIS PROBLEM CONDITION DOES NOT EXIST ON STS-27. THE ABOVE RATIONALE MEETS THE RATIONALE MEETS THE DEFERRAL REQUIREMENTS OF NSTS 07700 VOL XI AND NSTS 08126 PARA 3.4.1, ITEM C THIS DEFERRAL HAS BEEN APPROVED BY THE ET PROJECT MANAGER, J. SMELSER ____JERRY W. SMELSER (11/21/88)______ 11/2/88 GENERAL: TEST PROCEDURE 7R101-FT CHECKS THE TRANSDUCER FOR CALIBRATION ACCURACY, ELECTRICAL NOISE/DROPOUTS, AND STICKING/STALLING. THE TEST IS PERFORMED TO VERIFY THE TRANSDUCER OPERATION PRIOR TO INSTALLATION ON A FLIGHT VEHICLE. THE TEST USES AN ARRAY OF "5-WAY BINDING POSTS" MOUNTED ON A METAL BOX AS A MEANS OF TYING TOGETHER THE VARIOUS WIRES FROM THE TRANSDUCERS, POWER SUPPLY, METERS, AND STRIP CHART RECORDER. MANY ELECTRICAL CONNECTIONS BETWEEN THE BINDING POSTS AND OTHER COMPONENTS ARE MADE BY "POMONA" BANANA PLUG JUMPERS. THE GAS PRESSURE TO OPERATE THE TRANSDUCERS IS PROVIDED BY THE FACILITY NITROGEN SUPPLY AS REGULATED/CONTROLLED BY A PRESSURE REGULATOR AND A HIGH ACCURACY GAUGE FOUR TRANSDUCERS WHICH HAD BEEN THROUGH THIS TEST IMMEDIATELY PRIOR TO SHIPMENT TO KSC WERE INSTALLED ON LWT-29. UPON COMPLETION OF SPLICING THE TRANSDUCER WIRES INTO THE LWT-29 HARNESS, A RESISTANCE TEST WAS PERFORMED FROM THE ET/ORBITER ELECTRICAL DISCONNECT SO AS TO VERIFY THE SPLICING. DURING THIS TEST ONE TRANSDUCER WAS FOUND TO HAVE AN OPEN CIRCUIT BETWEEN THE WIPER AND THE "RETURN" END OF THE POTENTIOMETER. THE TRANSDUCER WAS REMOVED FROM LWT-29, RETURNED TO MAF, DOCUMENTED ON MARS T-101465, AND CAPS E-117 WAS OPEN TO TRACK THE PROBLEM AS PART OF THE FAILURE ANALYSIS ON MARS T-101465, THE MANUFACTURING HISTORY OF THE TRANSDUCER WAS REVIEWED. SUCH FAILURES ARE OFTEN THE THE RESULT OF IMPROPER APPLICATION OF POWER TO THE TRANSDUCER, I.E APPLYING 5 VOLTS ACROSS THE WIPER TO RETURN CONNECTIONS. DURING THIS REVIEW IT WAS FOUND THAT, DUE TO PERSONNEL ERROR, TEST 7R101 HAD NOT BEEN PROPERLY PERFORMED ON THE FOUR TRANSDUCERS FOR LWT-29 THE OUTPUT VOLTAGE WAS 10 MILLIVOLTS HIGHER THEN ALLOWABLE AT 6 PSIG ON ALL FOUR UNITS AND 70 MILLIVOLTS HIGHER THAN ALLOWABLE AT 21 PSIG AT THIS POINT MARS T-102801 WAS WRITTEN AGAINST THE THREE UNITS STILL INSTALLED ON LWT-29 AND THIS CAPS WAS OPENED. FURTHER REVIEW OF THE TEST RESULTS FROM THE LWT-29 UNITS REVEALED THAT ALL FOUR TRANSDUCER OUTPUTS HAD BEEN GANGED TOGETHER DURING THE NOISE TEST, THEREBY MASKING ANY SIGNAL DROPOUTS TASK I FAILURE INVESTIGATION A. THE RECORDS OF ALL USES OF TEST PROCEDURE 7R101-FT WILL BE REVIEWED TO ASSURE THAT NO OTHER ERRORS IN THE CALIBRATION DATA HAVE OCCURRED CLOSURE STATEMENT TP-7R101-FT HAD BEEN USED A TOTAL OF SEVEN TIMES. A REVIEW OF THE TEST RESULTS FOUND THAT, WITH THE EXCEPTION OF THE TEST FOR THE PARTS USED ON LWT-29, ALL TESTS WERE PROPERLY PERFORMED SPECIFICALLY, THE NOISE TEST STRIP CHARTS AND THE TWO POINT CALIBRATIONS WERE IN AGREEMENT WITH THE VENDOR ATP'S OF THE TRANSDUCERS B. ANALYSIS/SIMULATION OF THE ORIGINAL TEST FAILURE WAS PERFORMED IN AN EFFORT TO DETERMINE THE CAUSES OF THE ERRONEOUS DATA FOR THE LWT-29 TRANSDUCERS CLOSURE STATEMENT ANALYSIS AND TESTING FOUND THAT: 1) THE TRANSDUCERS WITH THE OPEN ELEMENT (CAPS E-117) WAS DAMAGED AT THE START OF, OR PRIOR TO, THE TEST; 2) ALL FOUR TRANSDUCER OUTPUTS WERE GANGED TOGETHER TASK II CORRECTIVE ACTION THE TEST EQUIPMENT USED TO PERFORM TP-7R101-FT WILL BE REDESIGNED IN ORDER TO PRECLUDE SUCH FAILURES IN THE FUTURE. THE TEST PROCEDURE WILL BE SIMILARITY REVISED CLOSURE STATEMENT A NEW TEST TOOL, F78-2478, WAS DEVELOPED AND RELEASED FOR USE IN THE CALIBRATION/NOISE TEST OF THE TRANSDUCERS. THE TOOL GREATLY REDUCED THE NUMBER OF JUMPERS USED AND THEREBY REDUCED THE LIKELIHOOD OF PERSONNEL ERROR. THE TEST PROCEDURES, TP-7R101-FT AND TP-7R103-FT, FOR THE BENCH TEST WERE REWRITTEN FOR CLARITY AND THE INCLUSION OF A FINAL RESISTANCE CHECK OF THE TRANSDUCERS AT COMPLETION OF THE TEST (REFERENCE: MEMO 3613-88-315) TASK CLOSED TASK III CLEARANCE OF EFFECTIVITIES ALL ET'S CLEARED. ONLY LWT-29 WAS AFFECTED BY THE TEST ERROR. THE THREE TRANSDUCERS ON LWT-29 WERE REMOVED IN ACCORDANCE WITH THE ENGINEERING DISPOSITION ON MARS T-102801 TASK CLOSED TASK IV CAPS CLOSURE SUMMARY THE TRANSDUCERS WERE IMPROPERLY TESTED DUE TO PERSONNEL ERROR. THE TEST EQUIPMENT AND TEST PROCEDURE WERE IMPROVED SO AS TO REDUCE THE LIKELIHOOD OF FUTURE, SIMILAR FAILURES. THE TEST PROCEDURE INCORPORATES A CHECK OF THE TRANSDUCER RESISTANCE, AT THE COMPLETION OF ALL OTHER TESTING, SO AS TO DETECT DAMAGED TRANSDUCERS TASK CLOSED THIS PROBLEM IS CONSIDERED CLOSED |
| MSFC Response/Concurrence |

ASSESSMENT ADDENDUM REPORT

| MSFC Report# A11809 | IFA# -- | Contractor RPT# E-118-1 | JSC# -- | KSC# -- | EICN# -- |
| --- | --- | --- | --- | --- | --- |
| Asmnt Part# PD7400098-079 | Asmnt Part Name LOX PRESS TRANSDUCER | Asmnt Serial/Lot# 260 |
| HCRIT CD -- | FCRIT CD 1R | CAUSE CD ETW - EI-TEST-WORK | FAIL MODE EG - SIG HI OR LO |
| Asmnt FMEA 3.2.1.1 | Asmnt FM 1 | FMEA CSE N/A | FMEA SCSE N/A |
| Asmnt FMEA -- | Asmnt FM -- | FMEA CSE -- | FMEA SCSE -- |
| Asmnt FMEA -- | Asmnt FM -- | FMEA CSE -- | FMEA SCSE -- |
| Correlated Part# -- | Correlated Part# -- | Correlated Part# -- |
| Associated LRU# -- | Associated LRU# -- | Associated LRU# -- |
| MAJOR DESIGN CHANGES |
| APRV DATE -- | DESCRIPTION OF CHANGES -- |
| ASSESSMENT TEXT |