## Readme
INSTRUCTIONS FOR CALIBRATION WORKBOOK
Purpose
This worksheet supports linear, quadratic, cubic, and quartic regression models:
| Linear: | y = b0 + b1x+ e
| Quadratic: | y = b0 + b1x + b2x2 + e
| Cubic: | y = b0 + b1x + b2x2 + b3x3 +e
| Quartic: | y = b0 + b1x + b2x2 + b3x3 +b4x4 + e
y = response
x = concentration
The worksheet estimates the coefficients (bs) and the variance of the error term, e. The workbook
then helps the user perform the following functions:
- | determine which model (linear, quadratic, etc.) is best.
- | determine the replication of unknowns needed for uncertainty control.
- | determine whether zero and span responses are acceptable.
- | estimate the concentration and 95% uncertainty of candidate standards
analyzed on the same day as the initial calibration or a subsequent day
Organization
The workbook consists of several worksheets, which are displayed as tabs at the bottom of
the screen. The functions of these worksheets are described below:
ReadMe (this sheet) - describes the workbook, explaining how to use the worksheets
Measurement Data - allows for user input of calibration and other analytical data
and includes statistical calculations for polynomial regression
Curves 1 - displays the calibration data, the best-fit line, and its confidence
bands.
Residuals 1 - displays the difference between the observed responses and those
estimated by the best-fit calibration line
Curves 2 - displays the calibration data, the best-fit quadratic curve, and its
confidence bands.
Residuals 2 - displays the difference between the observed responses and those
estimated by the best-fit quadratic regression line.
Curves 3 - displays the calibration data, the best-fit cubic curve, and its
confidence bands.
Residuals 3 - displays the difference between the observed responses and those
estimated by the best fit cubic regression line.
Curves 4 - displays the calibration data, the best-fit quartic curve, and its
confidence bands.
Residuals 4 - displays the difference between the observed responses and those
estimated by the best-fit quartic regression line.
Chart Data - includes the data used to create the Curves and Residuals charts
Conventions
The Measurement Data worksheet includes instructions that guide the user through the
steps in its use. The worksheet is also color coded to simplify use. Shaded cells that are
bordered in blue lines are for input of data. These cells are unprotected, but all other cells of
the Measurement Data worksheet are protected. The only other unprotected cell in the
workbook is cell F4 of the Chart Data worksheet. That cell controls the width of the
confidence bands plotted in the "Curves" charts.
Derived values and statements are colored red. These cells contain formulas and are
protected to prevent alteration.
Spreadsheet background colors indicate the order of the polynomial supported by the
calculations in the area.
Light green is used for the linear model.
Yellow is used for the quadratic model.
Gray color is used for the cubic model.
Light blue is used for the quartic model.
Use
The Measurement Data worksheet guides the user through six steps.
STEP 1 |  |  | Enter Calibration Data
In this step, up to 50 calibration points may be entered. Each calibration point has two parts, the
certified concentration of the calibration gas standard and the instrument response when testing
the standard. These values are entered in two columns. The spreadsheet performs
computations in cloumns I through P (linear), Q through X (quadratic), Y through AZ (cubic), and BA
and above (quartic).
STEP 2 |  |  | Review the Parameter Estimates
In this step, the user reviews the estimates of the intercepts (b0), slopes (b1) and other
coefficients (b2, b3, and b4) for the four models, examines their confidence intervals and the
residual error variances (s2). The result of an F test indicates which of the models is best. The
linear model is recommended unless the quadratic or higher order model significantly reduces
the residual error.
STEP 3 |  | Review the Charts
In this step, the user reviews the charts named Curves 1, Residuals 1, Curves 2, Residuals 2, etc.
These charts help the user understand why one model performs better than the other.
STEP 4 |  | Assess Magnitude of Uncertainty
In this step, the user enters the assumed concentration of a candidate standard and selects a
replication number, r. Based on the calibration results, the worksheet estimates the 95%
uncertainty that would result from measuring such a standard r times. The user can use this as a
guide for deciding whether to proceed with analysis, producing additional calibration points, or
taking some corrective action.
STEP 5 |  | Assay Candidate Standard on Same Day
In this step, the user enters the responses to a candidate standard that is tested on the same day
as the calibration of STEP 1. The worksheet provides an estimate of the candidate's
concentration and its 95% uncertainty. The worksheet also indicates whether the variability in
responses is larger than expected (unacceptable.)
STEP 6 |  | Assay Candidate Standard on
Different day from Initial Calibration
In this step, the user enters the responses to a candidate standard that is tested on a different day
from the calibration of STEP 1. The worksheet first assesses the zero and span responses. If
the zero and span responses are acceptable, the user proceeds to enter the results from testing
a candidate standard. The results include those for zero and non-zero reference standards. (The
quadratic, cubic, and quartic models requires the use of two, three, and four different non-zero
standards, respectively.)
The spreadsheet determines whether the regression curve has changed since the initial
calibration. The data are corrected for any change and the estimated concentration of the
candidate standard is provided together with its 95% uncertainty.
The spreadsheet also determines whether the standard error of the mean response is acceptable
(less than 1% of the mean response). This additional check is meant to guard against hysteresis
or other errors that are not corrected by the spreadsheet's adjustments.
Reference: Appendix A, EPA Traceability Protocol Assay and Certification of Gaseous
Calibration Standards, USEPA, 1997.

## Measurement Data
STEP 1 |  | Enter Calibration Data | [5 empty] | Calculations - Linear Model | [5 empty] | Calculations - Quadratic Model | [7 empty] | Calculations - Cubic Model | [6 empty] | Calculations - Cubic Model | [6 empty] | Calculations - Cubic Model | [6 empty] | Calculations - Cubic Model | [6 empty] | Calculations - Quartic Model | [6 empty] | Calculations - Quartic Model | [5 empty] | Calculations - Quartic Model | [6 empty] | Calculations - Quartic Model | [6 empty] | Calculations - Quartic Model | [6 empty] | Calculations - Quartic Model
Enter the calibration data in the shaded spaces below. The first column (i)
simply counts the calibration points that you enter. The second column (X) is
for the certified concentrations of the calibration gas standards. The third
column (Y) is for the instrument responses corresponding to the calibration
standards. The number of points cannot exceed 50.
i | Xi | Yi |  | Color Code |  |  |  | X'X | [5 empty] | X2 |  | X'X | [5 empty] | X3 |  | X'X | [25 empty] | X4 |  | X'X
1 | 0 | 0.2194 | [5 empty] | 19 | 85.5 |  |  |  |  | 0 |  | 19 | 85.5 | 527.25 |  |  |  | 0 |  | 19 | 85.5 | 527.25 | 3655.125 | [22 empty] | 0 |  | 19 | 85.5 | 527.25 | 3655.125 | 27021.5625
2 | 0.5 | 0.7141 |  | red = derived value (protected) |  |  |  | 85.5 | 527.25 |  |  |  |  | 0.25 |  | 85.5 | 527.25 | 3655.125 |  |  |  | 0.125 |  | 85.5 | 527.25 | 3655.125 | 27021.5625 | [22 empty] | 0.0625 |  | 85.5 | 527.25 | 3655.125 | 27021.5625 | 208037.5312
3 | 1 | 1.2885 | [11 empty] | 1 |  | 527.25 | 3655.125 | 27021.5625 |  |  |  | 1 |  | 527.25 | 3655.125 | 27021.5625 | 208037.5312 | [22 empty] | 1 |  | 527.25 | 3655.125 | 27021.5625 | 208037.5312 | 1647030.1406
4 | 1.5 | 1.9132 |  | blue = entered value (unprotected) |  |  |  | Det(X'X) | [5 empty] | 2.25 | [7 empty] | 3.375 |  | 3655.125 | 27021.5625 | 208037.5312 | 1647030.1406 | [22 empty] | 5.0625 |  | 3655.125 | 27021.5625 | 208037.5312 | 1647030.1406 | 13307929.3828
5 | 2 | 2.591 | [5 empty] | 2707.5 | [5 empty] | 4 |  | Det(X'X) | [5 empty] | 8 |  | Det(X'X) | [25 empty] | 16 |  | 27021.5625 | 208037.5312 | 1647030.1406 | 13307929.3828 | 109207686.0352
6 | 2.5 | 3.2866 |  | black = fixed text (protected) | [9 empty] | 6.25 |  | 2295621.5625 | [5 empty] | 15.625 |  | 11011063605.6095 | [25 empty] | 39.0625 | 1 | 2 | 3 | 4 | 5
7 | 3 | 4.1078 | [5 empty] | X'Xinverse | [5 empty] | 9 | [7 empty] | 27 | [27 empty] | 81 | 2 | X'Xinverse
8 | 3.5 | 4.9446 | [5 empty] | 0.19474 | -0.031579 |  |  |  |  | 12.25 |  | X'Xinverse | [5 empty] | 42.875 |  | X'Xinverse | [25 empty] | 150.0625 | 3 | 0.74537 | -0.94931 | 0.35677 | -0.051314 | 0.0024964
9 | 4 | 5.8145 | [5 empty] | -0.031579 | 0.0070175 |  |  |  |  | 16 |  | 0.38647 | -0.16692 | 0.015038 |  |  |  | 64 |  | 0.58168 | -0.46867 | 0.10116 | -0.0063796 | [22 empty] | 256 | 4 | -0.94931 | 1.9803 | -0.89428 | 0.1418 | -0.00733
10 | 4.5 | 6.723 | [11 empty] | 20.25 |  | -0.16692 | 0.10255 | -0.010615 |  |  |  | 91.125 |  | -0.46867 | 0.56899 | -0.14374 | 0.0098613 | [22 empty] | 410.0625 | 5 | 0.35677 | -0.89428 | 0.43832 | -0.072982 | 0.0038982
11 | 5 | 7.7284 | [5 empty] | X'Y | [5 empty] | 25 |  | 0.015038 | -0.010615 | 0.0011794 |  |  |  | 125 |  | 0.10116 | -0.14374 | 0.039175 | -0.0028145 | [22 empty] | 625 |  | -0.051314 | 0.1418 | -0.072982 | 0.012543 | -0.00068528
12 | 5.5 | 8.7566 | [5 empty] | 142.0161 | [5 empty] | 30.25 | [7 empty] | 166.375 |  | -0.0063796 | 0.0098613 | -0.0028145 | 0.00020848 | [22 empty] | 915.0625 |  | 0.0024964 | -0.00733 | 0.0038982 | -0.00068528 | 3.8071e-05
13 | 6 | 9.8013 | [5 empty] | 909.8948 | [5 empty] | 36 |  | X'Y | [5 empty] | 216 |  | X'Y | [25 empty] | 1296
14 | 6.5 | 10.8818 | [11 empty] | 42.25 |  | 142.0161 | [5 empty] | 274.625 |  | 142.0161 | [25 empty] | 1785.0625 |  | X'Y |  | b
15 | 7 | 12.0982 | [5 empty] | b |  | b' |  |  |  | 49 |  | 909.8948 | [5 empty] | 343 |  | 909.8948 | [25 empty] | 2401 |  | 142.0161 |  | 0.22056
16 | 7.5 | 13.3122 | [5 empty] | -1.0778 |  | -1.0778 | 1.9005 |  |  | 56.25 |  | 6463.08 | [5 empty] | 421.875 |  | 6463.08 | [25 empty] | 3164.0625 |  | 909.8948 |  | 0.92846
17 | 8 | 14.584 | [5 empty] | 1.9005 | [5 empty] | 64 | [7 empty] | 512 |  | 48559.4466 | [25 empty] | 4096 |  | 6463.08 |  | 0.13904
18 | 8.5 | 15.9238 | [11 empty] | 72.25 |  | b |  | b' |  |  |  | 614.125 |  | b | [25 empty] | 5220.0625 |  | 48559.4466 |  | -0.0069299
19 | 9 | 17.3271 | [5 empty] | Y'Y | [5 empty] | 81 |  | 0.19642 |  | 0.19642 | 1.0011 | 0.099935 |  | 729 |  | 0.19517 | [25 empty] | 6561 |  | 378182.0241 |  | 0.00038726
[8 empty] | 1584.6791 | [7 empty] | 1.0011 | [7 empty] | 1.003
[16 empty] | 0.099935 | [7 empty] | 0.099384
[8 empty] | df | [15 empty] | 4.0861e-05
[8 empty] | 17 | [7 empty] | t(0.95) | [7 empty] | t(0.95) | [27 empty] | t(0.95)
[16 empty] | 2.1199 | [7 empty] | 2.1314 | [27 empty] | 2.1448
t(0.95,df)
[8 empty] | 2.1098 | [7 empty] | b'X'Y = SSmodel | [7 empty] | b'X'Y = SSmodel | [27 empty] | b'X'Y = SSmodel
[16 empty] | 1584.6714 | [7 empty] | 1584.6714 | [27 empty] | 1584.6753
b'X'Y = SSmodel
[8 empty] | 1576.2036 | [7 empty] | SSregression | [7 empty] | SSregression | [27 empty] | SSregression
[16 empty] | 523.1675 | [7 empty] | 523.1675 | [27 empty] | 523.1715
[8 empty] | SSregression | [17 empty] | b' | [27 empty] | b'
[8 empty] | 514.6998 | [7 empty] | SSresidual | [7 empty] | SSresidual |  | 0.19517 | 1.003 | 0.099384 | 4.0861e-05 | [22 empty] | SSresidual |  | 0.22056 | 0.92846 | 0.13904 | -0.0069299 | 0.00038726
[16 empty] | 0.007769 | [7 empty] | 0.007761 | [27 empty] | 0.0038217
SSresidual
[8 empty] | 8.4755 | [7 empty] | s2 | [7 empty] | s2 | [27 empty] | s2
[16 empty] | 0.00048556 | [7 empty] | 0.0005174 | [27 empty] | 0.00027298
s2
[8 empty] | 0.49856 | [7 empty] | Fcrit(0.05) | [7 empty] | Fcrit(0.05) | [27 empty] | Fcrit(0.05)
[16 empty] | 2.3167 | [7 empty] | 2.3849 | [27 empty] | 2.463
Var-Covar
[8 empty] | 0.097088 | -0.015744 | [6 empty] | F | [7 empty] | F | [27 empty] | F
[8 empty] | -0.015744 | 0.0034987 | [6 empty] | 1026.7642 | [7 empty] | 0.93847 | [27 empty] | 1.8954
[8 empty] | n | [7 empty] | Preferred Model | [7 empty] | Preferred Model | [27 empty] | Preferred Model
[8 empty] | 19 | [7 empty] | Q | [7 empty] | Quadratic | [27 empty] | Cubic
[16 empty] | Var-Covar | [7 empty] | Var-Covar | [27 empty] | Var-Covar
[16 empty] | 0.00018765 | -8.1049e-05 | 7.3017e-06 | [5 empty] | 0.00028244 | -0.00022757 | 4.9121e-05 | -3.0977e-06 | [24 empty] | 0.00020347 | -0.00025914 | 9.7392e-05 | -1.4008e-05 | 6.8146e-07
[16 empty] | -8.1049e-05 | 4.9795e-05 | -5.1542e-06 | [5 empty] | -0.00022757 | 0.00027628 | -6.9796e-05 | 4.7883e-06 | [24 empty] | -0.00025914 | 0.00054058 | -0.00024412 | 3.8709e-05 | -2.0009e-06
[16 empty] | 7.3017e-06 | -5.1542e-06 | 5.7268e-07 | [5 empty] | 4.9121e-05 | -6.9796e-05 | 1.9022e-05 | -1.3666e-06 | [24 empty] | 9.7392e-05 | -0.00024412 | 0.00011965 | -1.9923e-05 | 1.0641e-06
[24 empty] | -3.0977e-06 | 4.7883e-06 | -1.3666e-06 | 1.0123e-07 | [24 empty] | -1.4008e-05 | 3.8709e-05 | -1.9923e-05 | 3.4241e-06 | -1.8707e-07
STEP 2 |  | Review the Parameter Estimates | [49 empty] | 6.8146e-07 | -2.0009e-06 | 1.0641e-06 | -1.8707e-07 | 1.0393e-08
Review the estimates of the coefficients (b0, b1, …)
coefficient for the linear and quadratic models, their confidence
limits and the residual error variances (s2).
Linear Model |  | 95% Conf. Limits
parameter | estimate | lower | upper
b0 = | -1.0778 | -1.7351 | -0.42036
b1 = | 1.9005 | 1.7757 | 2.0253
s2 = | 0.49856 | 0.28073 | 1.1205
s = | 0.70609 | 0.52984 | 1.0585
df = | 17
t = | 2.1098
Quadratic Model |  | 95% Conf. Limits
parameter | estimate | lower | upper
b0 = | 0.19642 | 0.19602 | 0.19682
b1 = | 1.0011 | 1.001 | 1.0012
b2 = | 0.099935 | 0.099934 | 0.099936
s2 = | 0.00048556 | 0.00026799 | 0.0011247
s = | 0.022036 | 0.01637 | 0.033537
df = | 16
t = | 2.1199
Comparing the two models:
F ratio = | 1026.7642
F critical = | 2.3167 | (5% significance level)
The quadratic model produces a significantly smaller error variance.
The quadratic model appears to be the better choice.
If cubic or quartic models are supported by compelling
scientific theory or data, then review the following estimates
for those models. Otherwise, go to Step 3.
Cubic Model |  | 95% Conf. Limits
parameter | estimate | lower | upper
b0 = | 0.19517 | 0.15935 | 0.23099
b1 = | 1.003 | 0.9676 | 1.0385
b2 = | 0.099384 | 0.090087 | 0.10868
b3 = | 4.0861e-05 | -0.0006373 | 0.00071902
s2 = | 0.0005174 | 0.00028089 | 0.0012394
s = | 0.022746 | 0.01676 | 0.035204
df = | 15
t = | 2.1314
Comparing quadratic and cubic models:
F ratio = | 0.93847
F critical = | 2.3849 | (5% significance level)
The error variances are not significantly different at the 5% level.
The quadratic model appears to be a better choice than cubic.
Quartic Model |  | 95% Conf. Limits
parameter | estimate | lower | upper
b0 = | 0.22056 | 0.18997 | 0.25116
b1 = | 0.92846 | 0.8786 | 0.97833
b2 = | 0.13904 | 0.11558 | 0.1625
b3 = | -0.0069299 | -0.010899 | -0.0029611
b4 = | 0.00038726 | 0.00016861 | 0.00060591
s2 = | 0.00027298 | 0.00014555 | 0.00067897
s = | 0.016522 | 0.012064 | 0.026057
df = | 14
t = | 2.1448
Comparing quadratic and cubic models:
F ratio = | 1.8954
F critical = | 2.463 | (5% significance level)
The error variances are not significantly different at the 5% level.
The cubic model appears to be a better choice than quartic.
STEP 3 |  | Review the Charts
View the charts named Curves 1 and Residuals 1. Curves 1 shows
confidence bands for the estimated regression. Compare these bands with
those of the quadratic regression, Curves 2.
(Note: You can change the width of the confidence band interval by
changing the "p-value" in cell F4 of the worksheet named Chart Data.)
Residuals 1 shows how the calibration points deviated from the calibration
line. Look for a simple pattern (such as a quadratic curve) in the chart. If
such a pattern appears, the quadratic model may be better. View Residual 2,
the deviations from the best- fit quadratic curve. If Residual 2 effectively
removes the simple pattern observed in Residual 1 and if the magnitude of
the deviations has been significantly reduced (as evidenced by a reduction
in the estimate s2), then the quadratic model is superior. An F-test can be
run to determine if the two error variances are significantly different.
F = | 1026.7642 |  | Prob. of greater F = | 4.5104e-21
The quadratic model produces a significantly smaller error variance.
The quadratic model appears to be the better choice.
If cubic or quartic models are supported by compelling
scientific theory or data, then view their Curves and Residuals charts.
STEP 4 |  | Assess Magnitude of Uncertainty | [5 empty] | STEP 4 |  | Assess Magnitude of Uncertainty |  |  |  | STEP 4 |  | Assess Magnitude of Uncertainty | [5 empty] | STEP 4 |  | Assess Magnitude of Uncertainty | [25 empty] | STEP 4 |  | Assess Magnitude of Uncertainty
Enter the concentration at which you would like to evaluate the uncertainty of | [7 empty] | Calculations - Linear Model | [5 empty] | Calculations - Quadratic Model | [7 empty] | Calculations - Cubic Model | [27 empty] | Calculations - Quartic Model
estimation and prediction. Also enter r, the number of assays to be performed | [49 empty] | x | xb | x'V | x'Vx
Increasing r tends to reduce the prediction uncertainty, but with diminishing | [7 empty] | x | xb | x'V | x'Vx |  |  | x | xb | x'V | x'Vx |  |  |  |  | x | xb | x'V | x'Vx | [24 empty] | 1 | 9.8017 | 1.2209e-05 | 4.3825e-05
effect. | [7 empty] | 1 | 10.3253 | 0.002624 | 0.034112 |  |  | 1 | 9.8006 | -3.5778e-05 | 4.9007e-05 |  |  |  |  | 1 | 9.7999 | 1.6263e-05 | 7.7579e-05 | [24 empty] | 6 |  | -3.6143e-05
[8 empty] | 6 |  | 0.005248 |  |  |  | 6 |  | 3.2171e-05 | [5 empty] | 6 |  | -4.8272e-05 | [25 empty] | 36 |  | 1.6006e-05
| Concentration = | 6 | [11 empty] | 36 |  | -3.0066e-06 | [5 empty] | 36 |  | 1.9953e-05 | [25 empty] | 216 |  | -1.7979e-06
[22 empty] | 216 |  | -1.7007e-06 | [25 empty] | 1296 |  | 4.6767e-08
| r = | 3
[8 empty] | 95% Conf. Limits for Response | [5 empty] | 95% Conf. Limits for Response
Review the estimated mean response (estimate that only takes | [7 empty] | Estimate | Lower | Upper |  |  |  | Estimate | Lower | Upper | [5 empty] | 95% Conf. Limits for Response | [27 empty] | 95% Conf. Limits for Response
into account the calibration uncertainty), and the confidence intervals. | [7 empty] | 10.3253 | 9.9356 | 10.715 |  |  |  | 9.8006 | 9.7858 | 9.8155 | [5 empty] | Estimate | Lower | Upper | [25 empty] | Estimate | Lower | Upper
Review the predicted mean response and its confidence intervals. | [21 empty] | 9.7999 | 9.7812 | 9.8187 | [25 empty] | 9.8017 | 9.8017 | 9.8017
To satisfy the EPA protocol requirements, the "95% uncertainty" | [7 empty] | 95% Conf. Limits for Response. | [5 empty] | 95% Conf. Limits for Response.
should be less than 1% in absolute value. | [7 empty] | Prediction | Lower | Upper |  |  |  | Prediction | Lower | Upper | Slope |  |  |  |  | 95% Conf. Limits for Response. | [27 empty] | 95% Conf. Limits for Response.
[8 empty] | 10.3253 | 9.3811 | 11.2695 |  |  |  | 9.8006 | 9.7698 | 9.8314 | 1.6007 |  |  |  |  | Prediction | Lower | Upper | Slope (near Conc.) | [24 empty] | Prediction | Lower | Upper | Slope (near Conc.)
Estimates below are based on the quadratic model. | [7 empty] | Conc. | Lower | Upper |  |  |  | Conc. | Lower | Upper | Check (Assumes Linearity in Neighborhood) |  |  |  |  | 9.7999 | 9.7662 | 9.8337 | 2.2 | [24 empty] | 9.8017 | 9.7768 | 9.8266 | 2.1831
Tab-Right to view estimates based on the other models. | [7 empty] | 6 | 5.5032 | 6.4968 |  |  |  | 6 | 5.986 | 6.014 | 5.9808 | 6.0192 |  |  |  | Conc. | Lower | Upper | [25 empty] | Conc. | Lower | Upper
[9 empty] | -0.082806 | 0.082806 |  |  |  |  | -0.0023332 | 0.0023303 | [5 empty] | 6 | 5.9847 | 6.0153 | [25 empty] | 6 | 5.9886 | 6.0114
|  |  |  | 95% Conf. Limits | [18 empty] | -0.0025533 | 0.0025533 | [26 empty] | -0.0019013 | 0.0019013
|  |  | Estimate | Lower | Upper
|  | Instrument Response = | 9.8006 | 9.7858 | 9.8155 |  |  | x | xb | x'V | x'Vx |  |  | x | xb | x'V | x'Vx
[8 empty] | 1 | 4.5 | 0.050881 | 0.03481 |  |  | 1 | 4.5 | 1.4818e-06 | 5.1041e-05 |  |  |  |  | x | xb | x'V | x'Vx | [24 empty] | x | xb | x'V | x'Vx
|  |  |  | 95% Conf. Limits |  |  |  | 2.9349 |  | -0.0054759 |  |  |  | 3.2466 |  | 2.6289e-05 | [5 empty] | 1 | 4.5 | -4.4637e-05 | 7.3488e-05 | [24 empty] | 1 | 4.5 | -1.5045e-05 | 4.3603e-05
|  |  | Prediction | Lower | Upper | [8 empty] | 10.5407 |  | -3.3955e-06 | [5 empty] | 3.2463 |  | 9.7589e-05 | [7 empty] | p | q | r | a | b | m | q | Solution 1 | Solution 2 | Solution 3 |  |  |  | Best Solution | Indicator | Unique? | A | B | 3.2428 |  | 2.5444e-05 | [9 empty] | quartic equation coefficients (r.h.s.=0) |  |  |  | resolvent cubic equation coefficients
|  | Instrument Response = | 9.8006 | 9.7698 | 9.8314 | [16 empty] | 10.5383 |  | -2.3747e-05 | [7 empty] | 2432.2567 | 24547.4475 | -105353.9846 | -1947410.1165 | 1045838750.7879 | 1611.3804 | 1.0412 | 3.2463 | -2422.104 | -13.399 | 3.2463 | 999999 | 999999 | 3.2463 | -87553701220288 | #VALUE! | #VALUE! | #VALUE! | 10.5159 |  | 2.3015e-06 | [9 empty] | a | b | c | d | pp | qq | rr | aa | bb | Indicator | mm |  | Conc1 | Conc2 | Conc3 | A | B | y y | R | D | E | Conc1 | Conc2 | Conc3 | Conc4 | In calibration range? |  |  |  | Best
|  | Concentration = | 6 | 5.986 | 6.014 | [16 empty] | 34.2103 |  | 1.5075e-06 | [25 empty] | 34.1011 |  | -1.9058e-06 | [9 empty] | -17.8945 | 359.024 | 2397.5035 | -11050.4909 | -359.024 | 1299.8641 | -18079082.8874 | -41666.2218 | -21351490.4092 | 111292437673780 | 235.701 | #NUM! | #NUM! | #NUM! | #NUM! | 276.8754 | 50.1624 | 446.7125 | 12.9515 | #VALUE! | 10.4899 | #VALUE! | #VALUE! | 3.2428 | -7.2471 | 0 | 0 | 3.2428 | 0 | 3.2428
|  |  | 95% uncertainty in prediction = | 0.2332% | [45 empty] | 110.5838 |  | 1.5306e-07
[23 empty] | Mean Response = | 4.5 | [26 empty] | Mean Response = | 4.5
STEP 5 |  | Assay Candidate Standard on Same Day | [5 empty] | STEP 5 |  | Assay Candidate Standard on Same Day |  |  |  | STEP 5 |  | Assay Candidate Standard on Same Day | [5 empty] | STEP 5 |  | Assay Candidate Standard on Same Day | [25 empty] | STEP 5 |  | Assay Candidate Standard on Same Day
Proceed with the analysis of candidate standards if their 95% uncertainties, | [7 empty] | Calculations - Linear Model | [5 empty] | Calculations - Quadratic Model | [7 empty] | Calculations - Cubic Model | [27 empty] | Calculations - Quartic Model | [11 empty] | quartic equation coefficients (r.h.s.=0) |  |  |  | resolvent cubic equation coefficients
as estimated above, are less than 1%. Enter the responses from the | [31 empty] | p | q | r | a | b | m | q | Solution 1 | Solution 2 | Solution 3 |  |  |  | Best Solution | Indicator | Unique? | A | B | [12 empty] | a | b | c | d | pp | qq | rr | aa | bb | Indicator | mm |  | Conc1 | Conc2 | Conc3 | A | B | yy | R | D | E | Conc1 | Conc2 | Conc3 | Conc4 | In calibration range? |  |  |  | Best
repeated analyses of an individual candidate standard in the spaces | [7 empty] | 2.9349 |  | 95% Conf. Limits for Response |  |  |  | 3.2466 |  | 95% Conf. Limits for Response | [5 empty] | 3.2463 |  | 95% Conf. Limits for Response | [7 empty] | 2432.2567 | 24547.4475 | -105353.9846 | -1947410.1165 | 1045838750.7879 | 1611.3804 | 1.0412 | 3.2463 | -2422.104 | -13.399 | 3.2463 | 999999 | 999999 | 3.2463 | -87553701220288 | #VALUE! | #VALUE! | #VALUE! | 3.2428 | 4.4943 | 95% Conf. Limits for Response | [9 empty] | -17.8945 | 359.024 | 2397.5035 | -11050.4909 | -359.024 | 1299.8641 | -18079082.8874 | -41666.2218 | -21351490.4092 | 111292437673780 | 235.701 | #NUM! | #NUM! | #NUM! | #NUM! | 276.8754 | 50.1624 | 446.7125 | 12.9515 | #VALUE! | 10.4899 | #VALUE! | #VALUE! | 3.2428 | -7.2471 | 0 | 0 | 3.2428 | 0 | 3.2428
provided below. | [7 empty] | 2.9354 |  | Estimate | Lower | Upper |  | 3.2473 |  | Estimate | Lower | Upper |  |  |  | 3.2469 |  | Estimate | Lower | Upper | [5 empty] | 2432.2567 | 24547.4475 | -105378.4581 | -1947410.1165 | 1045838726.3144 | 1611.3804 | 1.0412 | 3.2469 | -2422.104 | -13.3996 | 3.2469 | 999999 | 999999 | 3.2469 | -87566498857504 | #VALUE! | #VALUE! | #VALUE! | 3.2434 | 4.4953 | Estimate | Lower | Upper | [7 empty] | -17.8945 | 359.024 | 2397.5035 | -11053.0731 | -359.024 | 1310.1931 | -18081964.3561 | -41655.8929 | -21353135.7672 | 111311995699435 | 235.6718 | #NUM! | #NUM! | #NUM! | #NUM! | 276.883 | 50.1486 | 446.7063 | 12.9513 | #VALUE! | 10.4909 | #VALUE! | #VALUE! | 3.2434 | -7.2474 | 0 | 0 | 3.2434 | 0 | 3.2434
[8 empty] | 2.9343 |  | 4.5 | 4.1064 | 4.8936 |  | 3.246 |  | 4.5 | 4.4849 | 4.5151 |  |  |  | 3.2457 |  | 4.5 | 4.4817 | 4.5183 | [5 empty] | 2432.2567 | 24547.4475 | -105329.5112 | -1947410.1165 | 1045838775.2613 | 1611.3804 | 1.0412 | 3.2457 | -2422.104 | -13.3984 | 3.2457 | 999999 | 999999 | 3.2457 | -87540903582912 | #VALUE! | #VALUE! | #VALUE! | 3.2422 | 4.4933 | 4.5 | 4.4858 | 4.5142 | [7 empty] | -17.8945 | 359.024 | 2397.5035 | -11047.9087 | -359.024 | 1289.5352 | -18076201.4188 | -41676.5507 | -21349845.0512 | 111272880013896 | 235.7302 | #NUM! | #NUM! | #NUM! | #NUM! | 276.8678 | 50.1762 | 446.7187 | 12.9518 | #VALUE! | 10.4889 | #VALUE! | #VALUE! | 3.2422 | -7.2467 | 0 | 0 | 3.2422 | 0 | 3.2422
Note: This step applies only to candidate standards that are assayed on the | [31 empty] | 2432.2567 | 24547.4475 | 4776.5067 | -1947410.1165 | 1045948881.2792 | 1611.3804 | 1.0437 | -0.19849 | -2422.1228 | -9.9354 | 999999 | 999999 | 999999 | 999999 | -29961301290976 | #VALUE! | #VALUE! | #VALUE! | [12 empty] | -17.8945 | 359.024 | 2397.5035 | 569.5481 | -359.024 | -45180.2917 | -5112474.0131 | -88146.3777 | -13947379.2427 | 23266483739341 | 342.8243 | #NUM! | #NUM! | #NUM! | #NUM! | 227.646 | 129.0694 | 476.39 | 14.0506 | #VALUE! | 4.6098 | #VALUE! | #VALUE! | -0.24679 | -4.8566 | 0 | 0 | 0 | 0 | 0
same day as the calibration. | [9 empty] | 95% Conf. Limits for Response. | [5 empty] | 95% Conf. Limits for Response. | [7 empty] | 95% Conf. Limits for Response. | [7 empty] | 2432.2567 | 24547.4475 | 4776.5067 | -1947410.1165 | 1045948881.2792 | 1611.3804 | 1.0437 | -0.19849 | -2422.1228 | -9.9354 | 999999 | 999999 | 999999 | 999999 | -29961301290976 | #VALUE! | #VALUE! | #VALUE! |  |  | 95% Conf. Limits for Response. | [9 empty] | -17.8945 | 359.024 | 2397.5035 | 569.5481 | -359.024 | -45180.2917 | -5112474.0131 | -88146.3777 | -13947379.2427 | 23266483739341 | 342.8243 | #NUM! | #NUM! | #NUM! | #NUM! | 227.646 | 129.0694 | 476.39 | 14.0506 | #VALUE! | 4.6098 | #VALUE! | #VALUE! | -0.24679 | -4.8566 | 0 | 0 | 0 | 0 | 0
[10 empty] | Prediction | Lower | Upper |  |  |  | Prediction | Lower | Upper | [5 empty] | Prediction | Lower | Upper | Slope (near Conc.) |  |  |  |  | 2432.2567 | 24547.4475 | 4776.5067 | -1947410.1165 | 1045948881.2792 | 1611.3804 | 1.0437 | -0.19849 | -2422.1228 | -9.9354 | 999999 | 999999 | 999999 | 999999 | -29961301290976 | #VALUE! | #VALUE! | #VALUE! |  |  | Prediction | Lower | Upper | Slope (near Conc.) | [6 empty] | -17.8945 | 359.024 | 2397.5035 | 569.5481 | -359.024 | -45180.2917 | -5112474.0131 | -88146.3777 | -13947379.2427 | 23266483739341 | 342.8243 | #NUM! | #NUM! | #NUM! | #NUM! | 227.646 | 129.0694 | 476.39 | 14.0506 | #VALUE! | 4.6098 | #VALUE! | #VALUE! | -0.24679 | -4.8566 | 0 | 0 | 0 | 0 | 0
Below, enter the instrument responses for up to 10 repeated analyses of | [9 empty] | 4.5 | 3.5541 | 5.4459 |  |  |  | 4.5 | 4.4691 | 4.5309 | [5 empty] | 4.5 | 4.4666 | 4.5334 | 1.6496 |  |  |  |  | 2432.2567 | 24547.4475 | 4776.5067 | -1947410.1165 | 1045948881.2792 | 1611.3804 | 1.0437 | -0.19849 | -2422.1228 | -9.9354 | 999999 | 999999 | 999999 | 999999 | -29961301290976 | #VALUE! | #VALUE! | #VALUE! |  |  | 4.5 | 4.4751 | 4.5249 | 1.6116 | [6 empty] | -17.8945 | 359.024 | 2397.5035 | 569.5481 | -359.024 | -45180.2917 | -5112474.0131 | -88146.3777 | -13947379.2427 | 23266483739341 | 342.8243 | #NUM! | #NUM! | #NUM! | #NUM! | 227.646 | 129.0694 | 476.39 | 14.0506 | #VALUE! | 4.6098 | #VALUE! | #VALUE! | -0.24679 | -4.8566 | 0 | 0 | 0 | 0 | 0
a single candidate standard. | [9 empty] | Conc. | Lower | Upper |  |  |  | Conc. | Lower | Upper | [5 empty] | Conc. | Lower | Upper | [5 empty] | 2432.2567 | 24547.4475 | 4776.5067 | -1947410.1165 | 1045948881.2792 | 1611.3804 | 1.0437 | -0.19849 | -2422.1228 | -9.9354 | 999999 | 999999 | 999999 | 999999 | -29961301290976 | #VALUE! | #VALUE! | #VALUE! |  |  | Conc. | Lower | Upper | [7 empty] | -17.8945 | 359.024 | 2397.5035 | 569.5481 | -359.024 | -45180.2917 | -5112474.0131 | -88146.3777 | -13947379.2427 | 23266483739341 | 342.8243 | #NUM! | #NUM! | #NUM! | #NUM! | 227.646 | 129.0694 | 476.39 | 14.0506 | #VALUE! | 4.6098 | #VALUE! | #VALUE! | -0.24679 | -4.8566 | 0 | 0 | 0 | 0 | 0
[10 empty] | 2.9349 | 2.4372 | 3.4326 |  |  |  | 3.2466 | 3.2279 | 3.2654 | [5 empty] | 3.2463 | 3.226 | 3.2665 | [5 empty] | 2432.2567 | 24547.4475 | 4776.5067 | -1947410.1165 | 1045948881.2792 | 1611.3804 | 1.0437 | -0.19849 | -2422.1228 | -9.9354 | 999999 | 999999 | 999999 | 999999 | -29961301290976 | #VALUE! | #VALUE! | #VALUE! |  |  | 3.2428 | 3.2274 | 3.2583 | [7 empty] | -17.8945 | 359.024 | 2397.5035 | 569.5481 | -359.024 | -45180.2917 | -5112474.0131 | -88146.3777 | -13947379.2427 | 23266483739341 | 342.8243 | #NUM! | #NUM! | #NUM! | #NUM! | 227.646 | 129.0694 | 476.39 | 14.0506 | #VALUE! | 4.6098 | #VALUE! | #VALUE! | -0.24679 | -4.8566 | 0 | 0 | 0 | 0 | 0
Analysis |  | Estimated | [8 empty] | -0.16958 | 0.16958 |  |  |  |  | -0.0057806 | 0.0057806 | [6 empty] | -0.0062423 | 0.0062423 | [5 empty] | 2432.2567 | 24547.4475 | 4776.5067 | -1947410.1165 | 1045948881.2792 | 1611.3804 | 1.0437 | -0.19849 | -2422.1228 | -9.9354 | 999999 | 999999 | 999999 | 999999 | -29961301290976 | #VALUE! | #VALUE! | #VALUE! |  |  |  | -0.0047613 | 0.0047613 | [7 empty] | -17.8945 | 359.024 | 2397.5035 | 569.5481 | -359.024 | -45180.2917 | -5112474.0131 | -88146.3777 | -13947379.2427 | 23266483739341 | 342.8243 | #NUM! | #NUM! | #NUM! | #NUM! | 227.646 | 129.0694 | 476.39 | 14.0506 | #VALUE! | 4.6098 | #VALUE! | #VALUE! | -0.24679 | -4.8566 | 0 | 0 | 0 | 0 | 0
Number | Response | Conc. | [5 empty] | mean = | 2.9349 |  |  |  |  | mean = | 3.2466 | [6 empty] | mean = | 3.2463 | [26 empty] | mean = | 3.2428
1 | 4.5 | 3.2466 |  | NOTE: For Cubic and |  |  |  | stdev = | 0.00052618 |  |  |  |  | stdev = | 0.00060606 | [6 empty] | stdev = | 0.00060622 | [26 empty] | stdev = | 0.00060082
2 | 4.501 | 3.2473 |  | Quartic Model estimates, |  |  |  | df = | 2 |  |  |  |  | df = | 2 | [6 empty] | df = | 2 | [26 empty] | df = | 2
3 | 4.499 | 3.246 |  | view the Calculations in |  |  |  | F = | 5.5532e-07 |  |  |  |  | F = | 0.00075646 | [6 empty] | F = | 0.00071028 | [26 empty] | F = | 0.0013224
4 |  |  |  | the shaded regions to |  |  |  | F sig? = | 0 |  |  |  |  | F sig? = | 0 | [6 empty] | F sig? = | 0 | [26 empty] | F sig? = | 0
5 |  |  |  | the right. |  |  |  | Pr{>F} = | 1 |  |  |  |  | Pr{>F} = | 0.99924 | [6 empty] | Pr{>F} = | 0.99929 | [26 empty] | Pr{>F} = | 0.99868
6 | [7 empty] | 95% U = | 16.96% |  |  |  |  | 95% U = | 0.5781% | [6 empty] | 95% U = | 0.6242% | [26 empty] | 95% U = | 0.4761%
7 | [7 empty] | 0.17319 | = Portion of Uncertainty2 Due to Calibration |  |  |  |  | 0.23975 | = Portion of Uncertainty2 Due to Calibration | [6 empty] | 0.29879 | = Portion of Uncertainty2 Due to Calibration | [26 empty] | 0.32396 | = Portion of Uncertainty2 Due to Calibration
8
9
10
| mean = | 3.2466
| stdev = | 0.00060606
| df = | 2
| F = | 0.00075646
| F sig? = | 0 | (The sample variance is acceptable.)
| Pr{>F} = | 0.99924
| 95% Uncertainty = | 0.5781%
| 23.97% | = Portion of Uncertainty2 Due to Calibration
STEP 6 |  | Assay Candidate Standard on | [5 empty] | STEP 6 |  | Assay Candidate Standard on |  |  |  | STEP 6 |  | Assay Candidate Standard on | [5 empty] | STEP 6 |  | Assay Candidate Standard on | [25 empty] | STEP 6 |  | Assay Candidate Standard on
|  | Different day from Initial Calibration | [7 empty] | Different day from Initial Calibration | [5 empty] | Different day from Initial Calibration | [7 empty] | Different day from Initial Calibration | [27 empty] | Different day from Initial Calibration
[8 empty] | X'X | [5 empty] | X'X | [7 empty] | X'X | [27 empty] | X'X
This step applies to candidate standards that are assayed on a different day | [7 empty] | 19 | 85.5 | 0 | 0 |  |  | 19 | 85.5 | 527.25 | 0 | 0 | 0 |  |  | 19 | 85.5 | 527.25 | 3655.125 | 0 | 0 | 0 | 0 | [20 empty] | 19 | 85.5 | 527.25 | 3655.125 | 27021.5625 | 0 | 0 | 0 | 0 | 0
than the initial calibration. Before candidate standards are run, the | [7 empty] | 85.5 | 527.25 | 0 | 0 |  |  | 85.5 | 527.25 | 3655.125 | 0 | 0 | 0 |  |  | 85.5 | 527.25 | 3655.125 | 27021.5625 | 0 | 0 | 0 | 0 | [20 empty] | 85.5 | 527.25 | 3655.125 | 27021.5625 | 208037.5312 | 0 | 0 | 0 | 0 | 0
measurement system is challenged with zero and span checks. Three or | [7 empty] | 0 | 0 | 9 | 40.5 |  |  | 527.25 | 3655.125 | 27021.5625 | 0 | 0 | 0 |  |  | 527.25 | 3655.125 | 27021.5625 | 208037.5312 | 0 | 0 | 0 | 0 | [20 empty] | 527.25 | 3655.125 | 27021.5625 | 208037.5312 | 1647030.1406 | 0 | 0 | 0 | 0 | 0
more discrete checks of the zero gas and three or more checks of the span | [7 empty] | 0 | 0 | 40.5 | 303.75 |  |  | 0 | 0 | 0 | 9 | 40.5 | 303.75 |  |  | 3655.125 | 27021.5625 | 208037.5312 | 1647030.1406 | 0 | 0 | 0 | 0 | [20 empty] | 3655.125 | 27021.5625 | 208037.5312 | 1647030.1406 | 13307929.3828 | 0 | 0 | 0 | 0 | 0
gas are made. Enter the results below: | [13 empty] | 0 | 0 | 0 | 40.5 | 303.75 | 2460.375 |  |  | 0 | 0 | 0 | 0 | 9 | 40.5 | 303.75 | 2460.375 | [20 empty] | 27021.5625 | 208037.5312 | 1647030.1406 | 13307929.3828 | 109207686.0352 | 0 | 0 | 0 | 0 | 0
[8 empty] | Det(X'X) | UpLeft | DnRight |  |  |  | 0 | 0 | 0 | 303.75 | 2460.375 | 20913.1875 |  |  | 0 | 0 | 0 | 0 | 40.5 | 303.75 | 2460.375 | 20913.1875 | [20 empty] | 0 | 0 | 0 | 0 | 0 | 9 | 40.5 | 303.75 | 2460.375 | 20913.1875
| Response to |  | Span | Response |  |  |  | 2960651.25 | 2707.5 | 1093.5 | [11 empty] | 0 | 0 | 0 | 0 | 303.75 | 2460.375 | 20913.1875 | 182682.8438 | [20 empty] | 0 | 0 | 0 | 0 | 0 | 40.5 | 303.75 | 2460.375 | 20913.1875 | 182682.8438
| Zero gas |  | Conc. | to Span | [9 empty] | Det(X'X) | UpLeft | DnRight | [5 empty] | 0 | 0 | 0 | 0 | 2460.375 | 20913.1875 | 182682.8438 | 1619234.2969 | [20 empty] | 0 | 0 | 0 | 0 | 0 | 303.75 | 2460.375 | 20913.1875 | 182682.8438 | 1619234.2969
| 0 |  | 9 | 16.01 |  |  |  | X'Xinverse | [5 empty] | 2.0587e+12 | 2295621.5625 | 896806.6875 | [33 empty] | 0 | 0 | 0 | 0 | 0 | 2460.375 | 20913.1875 | 182682.8438 | 1619234.2969 | 14461007.8359
| 0.001 |  | 9 | 16 |  |  |  | 0.19474 | -0.031579 | 0 | 0 | [10 empty] | X'Xinverse | [27 empty] | 0 | 0 | 0 | 0 | 0 | 20913.1875 | 182682.8438 | 1619234.2969 | 14461007.8359 | 129644616.7617
| -0.001 |  | 9 | 15.99 |  |  |  | -0.031579 | 0.0070175 | 0 | 0 |  |  | X'Xinverse | [7 empty] | 0.58168 | -0.46867 | 0.10116 | -0.0063796 | 0 | 0 | 0 | 0
[8 empty] | 0 | 0 | 0.27778 | -0.037037 |  |  | 0.38647 | -0.16692 | 0.015038 | 0 | 0 | 0 |  |  | -0.46867 | 0.56899 | -0.14374 | 0.0098613 | 0 | 0 | 0 | 0 | [20 empty] | X'Xinverse |  |  |  |  | 1 | 2 | 3 | 4 | 5
[8 empty] | 0 | 0 | -0.037037 | 0.0082305 |  |  | -0.16692 | 0.10255 | -0.010615 | 0 | 0 | 0 |  |  | 0.10116 | -0.14374 | 0.039175 | -0.0028145 | 0 | 0 | 0 | 0 | [20 empty] | 0.74537 | -0.94931 | 0.35677 | -0.051314 | 0.0024964 | 0 | 0 | 0 | 0 | 0
[14 empty] | 0.015038 | -0.010615 | 0.0011794 | 0 | 0 | 0 |  |  | -0.0063796 | 0.0098613 | -0.0028145 | 0.00020848 | 0 | 0 | 0 | 0 | [20 empty] | -0.94931 | 1.9803 | -0.89428 | 0.1418 | -0.00733 | 0 | 0 | 0 | 0 | 0
[8 empty] | X'Y | b | Y'Y | df |  | Estimate | 0 | 0 | 0 | 0.33333 | -0.11111 | 0.0082305 |  | Estimate | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | [20 empty] | 0.35677 | -0.89428 | 0.43832 | -0.072982 | 0.0038982 | 0 | 0 | 0 | 0 | 0 |  |  | quartic equation coefficients (r.h.s.=0) |  |  |  | resolvent cubic equation coefficients
[8 empty] | 142.0161 | -1.0778 | 2621.3774 | 24 |  | 2.3525 | 0 | 0 | 0 | -0.11111 | 0.107 | -0.010974 |  | 2.94 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  | Estimate | p | q | r | a | b | m | q | Solution 1 | Solution 2 | Solution 3 |  |  |  | Best Solution | Indicator | Unique? | A | B | -0.051314 | 0.1418 | -0.072982 | 0.012543 | -0.00068528 | 0 | 0 | 0 | 0 | 0 |  | Estimate | a | b | c | d | pp | qq | rr | aa | bb | Indicator | mm |  | Conc1 | Conc2 | Conc3 | A | B | yy | R | D | E | Conc1 | Conc2 | Conc3 | Conc4 | In calibration range? |  |  |  | Best
[8 empty] | 909.8948 | 1.9005 |  |  |  | 2.3473 | 0 | 0 | 0 | 0.0082305 | -0.010974 | 0.0012193 |  | 2.9337 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0.0024964 | -0.00733 | 0.0038982 | -0.00068528 | 3.8071e-05 | 0 | 0 | 0 | 0 | 0 |  | 2.9432 | -17.8945 | 359.024 | 2397.5035 | -9785.1978 | -359.024 | -3761.3084 | -16667163.2544 | -46727.3943 | -20545264.9711 | 101748204099675 | 249.6061 | #NUM! | #NUM! | #NUM! | #NUM! | 273.0592 | 57.0418 | 449.7757 | 13.0692 | #VALUE! | 10.0084 | #VALUE! | #VALUE! | 2.9432 | -7.0652 | 0 | 0 | 2.9432 | 0 | 2.9432
[8 empty] | 72.8285 | -0.46179 | t | b'X'Y |  | 2.342 | [7 empty] | 2.9274 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 1 | 2.937 | -17.8945 | 359.024 | 2397.5035 | -9759.3755 | -359.024 | -3864.5976 | -16638348.568 | -46830.6836 | -20528811.3907 | 101554136228515 | 249.8818 | #NUM! | #NUM! | #NUM! | #NUM! | 272.9794 | 57.1846 | 449.8387 | 13.0716 | #VALUE! | 9.9984 | #VALUE! | #VALUE! | 2.937 | -7.0614 | 0 | 0 | 2.937 | 0 | 2.937
n = | 3 |  | n = | 3 |  |  |  | 558.6822 | 1.9009 | 2.0639 | 2604.546 |  |  | X'Y | b | Y'Y | df | [13 empty] | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 2 | 2.9307 | -17.8945 | 359.024 | 2397.5035 | -9733.5532 | -359.024 | -3967.8869 | -16609533.8817 | -46933.9728 | -20512357.8103 | 101360092690636 | 250.1572 | #NUM! | #NUM! | #NUM! | #NUM! | 272.8995 | 57.3275 | 449.9017 | 13.0741 | #VALUE! | 9.9883 | #VALUE! | #VALUE! | 2.9307 | -7.0576 | 0 | 0 | 2.9307 | 0 | 2.9307
mean = | 0 |  | mean = | 16 | [9 empty] | 142.0161 | 0.19642 | 2621.3774 | 22 |  |  |  |  | X'Y | b | Y'Y | df | [6 empty] | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 3 |  | -17.8945 | 359.024 | 2397.5035 | 569.5481 | -359.024 | -45180.2917 | -5112474.0131 | -88146.3777 | -13947379.2427 | 23266483739341 | 342.8243 | #NUM! | #NUM! | #NUM! | #NUM! | 227.646 | 129.0694 | 476.39 | 14.0506 | #VALUE! | 4.6098 | #VALUE! | #VALUE! | -0.24679 | -4.8566 | 0 | 0 | 0 | 0 | 0
s = | 0.001 |  | s = | 0.01 |  |  |  | SSregr | SSresid | s2 | s |  |  | 909.8948 | 1.0011 | [6 empty] | 142.0161 | 0.19517 | 2621.3774 | 20 | [6 empty] | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 4 |  | -17.8945 | 359.024 | 2397.5035 | 569.5481 | -359.024 | -45180.2917 | -5112474.0131 | -88146.3777 | -13947379.2427 | 23266483739341 | 342.8243 | #NUM! | #NUM! | #NUM! | #NUM! | 227.646 | 129.0694 | 476.39 | 14.0506 | #VALUE! | 4.6098 | #VALUE! | #VALUE! | -0.24679 | -4.8566 | 0 | 0 | 0 | 0 | 0
Cal. Resp. = | 0.19642 |  | Cal. Response = | 17.301 |  |  |  | 956.0387 | 16.8315 | 0.70131 | 0.83744 |  |  | 6463.08 | 0.099935 | [6 empty] | 909.8948 | 1.003 | [8 empty] | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0 | 0 | 0 | 0 | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 5 |  | -17.8945 | 359.024 | 2397.5035 | 569.5481 | -359.024 | -45180.2917 | -5112474.0131 | -88146.3777 | -13947379.2427 | 23266483739341 | 342.8243 | #NUM! | #NUM! | #NUM! | #NUM! | 227.646 | 129.0694 | 476.39 | 14.0506 | #VALUE! | 4.6098 | #VALUE! | #VALUE! | -0.24679 | -4.8566 | 0 | 0 | 0 | 0 | 0
[14 empty] | 72.8285 | 0.2194 | t | b'X'Y |  |  |  |  | 6463.08 | 0.099384 | [8 empty] | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | [12 empty] | -17.8945 | 359.024 | 2397.5035 | 569.5481 | -359.024 | -45180.2917 | -5112474.0131 | -88146.3777 | -13947379.2427 | 23266483739341 | 342.8243 | #NUM! | #NUM! | #NUM! | #NUM! | 227.646 | 129.0694 | 476.39 | 14.0506 | #VALUE! | 4.6098 | #VALUE! | #VALUE! | -0.24679 | -4.8566 | 0 | 0 | 0 | 0 | 0
|  | Zero Gas Results |  | Span Gas Results | [9 empty] | 558.6822 | 0.9926 | 2.0739 | 2621.3662 |  |  |  |  | 48559.4466 | 4.0861e-05 | t | b'X'Y |  | x | x'V |  |  |  | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | X'Y | b | Y'Y | df | [8 empty] | -17.8945 | 359.024 | 2397.5035 | 569.5481 | -359.024 | -45180.2917 | -5112474.0131 | -88146.3777 | -13947379.2427 | 23266483739341 | 342.8243 | #NUM! | #NUM! | #NUM! | #NUM! | 227.646 | 129.0694 | 476.39 | 14.0506 | #VALUE! | 4.6098 | #VALUE! | #VALUE! | -0.24679 | -4.8566 | 0 | 0 | 0 | 0 | 0
| Std. Error = s/sqrt(n) = | 0.00057735 |  | 0.0057735 | [9 empty] | 4619.3125 | 0.10092 | [6 empty] | 72.8285 | #VALUE! | 2.086 | #VALUE! |  | 1 | #VALUE! |  |  |  | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 142.0161 | 0.22056 | 2621.3774 | 18 | [8 empty] | -17.8945 | 359.024 | 2397.5035 | 569.5481 | -359.024 | -45180.2917 | -5112474.0131 | -88146.3777 | -13947379.2427 | 23266483739341 | 342.8243 | #NUM! | #NUM! | #NUM! | #NUM! | 227.646 | 129.0694 | 476.39 | 14.0506 | #VALUE! | 4.6098 | #VALUE! | #VALUE! | -0.24679 | -4.8566 | 0 | 0 | 0 | 0 | 0
| Rrs/100 = | 0.16 |  | 0.16 | [7 empty] | nnn = | 3 | [6 empty] | nnn = | 3 | 558.6822 | #VALUE! |  |  |  | #VALUE! | #VALUE! |  |  |  | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 909.8948 | 0.92846 | [10 empty] | -17.8945 | 359.024 | 2397.5035 | 569.5481 | -359.024 | -45180.2917 | -5112474.0131 | -88146.3777 | -13947379.2427 | 23266483739341 | 342.8243 | #NUM! | #NUM! | #NUM! | #NUM! | 227.646 | 129.0694 | 476.39 | 14.0506 | #VALUE! | 4.6098 | #VALUE! | #VALUE! | -0.24679 | -4.8566 | 0 | 0 | 0 | 0 | 0
|  | Std. Error is okay. |  | Std. Error is okay. | [7 empty] | mean = | 2.3473 | SSregr | SSresid | s2 | s |  |  | mean = | 2.9337 | 4619.3125 | #VALUE! |  |  |  | #VALUE! | #VALUE! |  | nnn = | 0 | [18 empty] | 6463.08 | 0.13904 | [8 empty] | nnn = | 3
| Relative Difference (RD)= | 1.135% |  | -7.52% | [7 empty] | stdev = | 0.0052608 | 972.859 | 0.011238 | 0.0005108 | 0.022601 |  |  | stdev = | 0.0063102 | 39734.0903 | #VALUE! |  |  |  | #VALUE! | #VALUE! |  | mean = | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 48559.4466 | -0.0069299 | t | b'X'Y |  | x | x'V |  |  |  | mean = | 2.937 | -17.8945 | 359.024 | 2397.5035 | -9759.3755 | -359.024 | -3864.5976 | -16638348.568 | -46830.6836 | -20528811.3907 | 101554136228515 | 249.8818 | #NUM! | #NUM! | #NUM! | #NUM! | 272.9794 | 57.1846 | 449.8387 | 13.0716 | #VALUE! | 9.9984 | #VALUE! | #VALUE! | 2.937 | -7.0614 | 0 | 0 | 2.937 | 0 | 2.937
|  | RD is okay. |  | RD is excessive. | [7 empty] | df = | 2 | [6 empty] | df = | 2 | [8 empty] | stdev = | #VALUE! | [18 empty] | 378182.0241 | 0.00038726 | 2.1009 | #VALUE! |  | 1 | #VALUE! |  |  |  | stdev = | 0.0062303
[12 empty] | F = | 3.9463e-05 | [6 empty] | F = | 0.077954 | SSregr | SSresid | s2 | s |  |  |  |  | df = | -1 | [18 empty] | 72.8285 | #VALUE! |  |  |  | 2.937 | #VALUE! |  |  |  | df = | 2
Following successful completion of the zero and span checks, the candidate | [11 empty] | F sig? = | 0 | [6 empty] | F sig? = | 0 | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  |  |  |  | F = | #VALUE! | [18 empty] | 558.6822 | #VALUE! |  |  |  | 8.6259 | #VALUE! |  |  |  | F = | #VALUE!
standard is measured together with reference standards. While the candidate | [11 empty] | Pr{>F} = | 0.99996 | [6 empty] | Pr{>F} = | 0.92526 | [8 empty] | F sig? = | #VALUE! | [18 empty] | 4619.3125 | #VALUE! |  |  |  | 25.334 | #VALUE! |  |  |  | F sig? = | #VALUE!
standard is normally interspersed with the reference standards, the analysis | [29 empty] | Pr{>F} = | #VALUE! | [18 empty] | 39734.0903 | #VALUE! |  |  |  | 74.4053 | #VALUE! |  |  |  | Pr{>F} = | #VALUE!
conducted in this sheet requires that the results be entered separately. There are | [49 empty] | 349328.0611 | #VALUE!
two ways to do this. One way is to enter an analysis set (one candidate standard
response and the responses from its zero and nonzero standard analyses) | [49 empty] | SSregr | SSresid | s2 | s
below. Another approach is to enter all of the responses (multiple sets) below. | [49 empty] | #VALUE! | #VALUE! | #VALUE! | #VALUE!
Enter zero and reference standard responses in the area to the left and enter the
responses to a single candidate standard in the are to the right, below.
STEP 6 | (Continued) | [6 empty] | STEP 6 |  | Assay Candidate Standard on |  |  |  | STEP 6 |  | Assay Candidate Standard on | [5 empty] | STEP 6 |  | Assay Candidate Standard on | [25 empty] | STEP 6 |  | Assay Candidate Standard on
[10 empty] | Different day from Initial Calibration, continued | [5 empty] | Different day from Initial Calibration, continued | [7 empty] | Different day from Initial Calibration, continued | [27 empty] | Different day from Initial Calibration, continued
More than one nonzero reference standard is needed | [7 empty] | Calculations - Linear Model | [5 empty] | Calculations - Quadratic Model | [7 empty] | Calculations - Cubic Model | [27 empty] | Calculations - Quartic Model
for the quadratic and higher-ordered models.
[8 empty] | Var-Covar | [5 empty] | Var-Covar | [7 empty] | Var-Covar | [27 empty] | Var-Covar
Estimates below are based on the quadratic model. | [7 empty] | 0.13657 | -0.022147 | 0 | 0 |  |  | 0.00019741 | -8.5262e-05 | 7.6813e-06 | 0 | 0 | 0 |  |  | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | [20 empty] | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
Tab-Right to view estimates based on the other models. | [7 empty] | -0.022147 | 0.0049215 | 0 | 0 |  |  | -8.5262e-05 | 5.2383e-05 | -5.4221e-06 | 0 | 0 | 0 |  |  | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | [20 empty] | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
Reference Standards | [5 empty] | Candidate Standard |  | 0 | 0 | 0.19481 | -0.025974 |  |  | 7.6813e-06 | -5.4221e-06 | 6.0245e-07 | 0 | 0 | 0 |  |  | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | [20 empty] | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
(Enter 0 for zero Conc.) | [5 empty] | Observed | Estimated | 0 | 0 | -0.025974 | 0.0057721 |  |  | 0 | 0 | 0 | 0.00017027 | -5.6756e-05 | 4.2042e-06 |  |  | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | [20 empty] | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
Conc. | Response | Conc.^2 | Conc.^3 | Conc.^4 |  | Response | Conc. | [6 empty] | 0 | 0 | 0 | -5.6756e-05 | 5.4654e-05 | -5.6055e-06 |  |  | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | [20 empty] | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
0 | 0.2184 | 0 | 0 | 0 |  | 4.01 | 2.94 | Have Coefficients Remained the Same? | [5 empty] | 0 | 0 | 0 | 4.2042e-06 | -5.6055e-06 | 6.2284e-07 |  |  | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | [20 empty] | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
0 | 0.2194 | 0 | 0 | 0 |  | 4 | 2.9337 | b0-b2 | -1.8041 | 0.57214 | 1 | [10 empty] | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | [20 empty] | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
0 | 0.2204 | 0 | 0 | 0 |  | 3.99 | 2.9274 | b1-b3 | -0.21378 | 0.21308 | 1 |  |  | Have Coefficients Remained the Same? | [7 empty] | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | [20 empty] | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
4.5 | 6.693 | 20.25 | 91.125 | 410.0625 | [9 empty] | b0-b3 | -0.022979 | -0.062745 | 0.016788 | 1 |  |  |  | Have Coefficients Remained the Same? | [27 empty] | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
4.5 | 6.723 | 20.25 | 91.125 | 410.0625 |  |  |  | Coefficients are not significantly different. | [5 empty] | b1-b4 | 0.0084949 | -0.012961 | 0.029951 | 1 |  |  |  | b0-b4 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | [23 empty] | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
4.5 | 6.773 | 20.25 | 91.125 | 410.0625 |  |  |  | Consider including the new data as part of original calibration (Step 1). | [5 empty] | b2-b5 | -0.00098256 | -0.0032782 | 0.0013131 | 1 |  |  |  | b1-b5 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
9 | 17.3171 | 81 | 729 | 6561 | [17 empty] | b2-b6 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | [23 empty] | Have Coefficients Remained the Same?
9 | 17.3271 | 81 | 729 | 6561 |  |  |  | Mean Response |  | 95% Uncert Limits |  |  |  | Coefficients are not significantly different. | [7 empty] | b3-b7 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | [23 empty] | b0-b5 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
9 | 17.3371 | 81 | 729 | 6561 |  |  |  | 4 |  | 2.8219 | 5.1781 |  |  | Consider including the new data as part of original calibration (Step 1). | [35 empty] | b1-b6 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
[22 empty] | #VALUE! | [27 empty] | b2-b7 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
[6 empty] | nnn = | 3 | Std Variance of Estimate |  |  | Estimation/Prediction Variance |  |  | Mean Response |  | 95% Uncert Limits | [5 empty] | #VALUE! | [27 empty] | b3-b8 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
[5 empty] | mean = | 4 | 2.9337 | 0.092083 |  |  | 0.39391 |  |  | 4 |  | 3.9633 | 4.0367 | [32 empty] | b4-b9 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
[6 empty] | stdev = | 0.0063102 | [14 empty] | Mean Response |  | 95% Uncert Limits | [7 empty] | p | q | r | a | b | m | q | Solution 1 | Solution 2 | Solution 3 |  |  |  | Best Solution | Indicator | Unique? | A | B
[6 empty] | std err = | 0.1242% | Conc. | 95% Uncert Limits |  |  |  |  | Std Variance of Estimate |  |  | Estimation/Prediction Variance |  |  |  |  | 2.9337 |  | #VALUE! | #VALUE! | [6 empty] | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
[6 empty] | df = | 2 | 2.3473 | 1.7275 | 2.9671 |  |  |  | 0.00014308 |  |  | 0.45661 | [14 empty] | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
[5 empty] | F = | 0.077954 |  |  | -26.41% | 26.41% | [11 empty] | Std Variance of Estimate |  |  | Estimation/Prediction Variance | [6 empty] | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
[5 empty] | F sig? = | 0 | [7 empty] | Conc. | 95% Uncert Limits | [6 empty] | #VALUE! |  |  | #VALUE!
[5 empty] | Pr{>F} = | 0.92526 | [7 empty] | 2.9337 | 2.9105 | 2.9569 | [33 empty] | Mean Response |  | 95% Uncert Limits
[14 empty] | 0.7896% | -0.7908% | 0.7885% | [5 empty] | Conc. | 95% Uncert Limits | [26 empty] | 2.9337 |  | #VALUE! | #VALUE!
[5 empty] | The standard error is okay. | [16 empty] | #VALUE! | #VALUE! | #VALUE! | [37 empty] | quartic equation coefficients (r.h.s.=0) |  |  |  | resolvent cubic equationcoefficients
nn = | 9 | [20 empty] | #VALUE! | #VALUE! | #VALUE! | [25 empty] | Std Variance of Estimate |  |  | Estimation/Prediction Variance | [8 empty] | a | b | c | d | pp | qq | rr | aa | bb | Indicator | mm |  | Conc1 | Conc2 | Conc3 | A | B | yy | R | D | E | Conc1 | Conc2 | Conc3 | Conc4 | In calibration range? |  |  |  | Best
[50 empty] | #VALUE! |  |  | #VALUE! | [8 empty] | -17.8945 | 359.024 | 2397.5035 | -7006.0401 | -359.024 | -14877.9393 | -13565947.2255 | -57844.0252 | -18774428.1662 | 8.0952e+13 | 277.7145 | #NUM! | #NUM! | #NUM! | #NUM! | 263.9272 | 73.0555 | 456.6574 | 13.3299 | #VALUE! | 8.868 | #VALUE! | #VALUE! | 2.2427 | -6.6253 | 0 | 0 | 2.2427 | 0 | 2.2427
Coefficients are not significantly different. | [61 empty] | -17.8945 | 359.024 | 2397.5035 | #VALUE! | -359.024 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0 | 0 | 0 | 0 | 0
Consider including the new data as part of original calibration (Step 1). | [49 empty] | Conc. | 95% Uncert Limits | [10 empty] | -17.8945 | 359.024 | 2397.5035 | #VALUE! | -359.024 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 0 | 0 | 0 | 0 | 0
[50 empty] | 2.2427 | 0 | 0
Estimated Concentration of Candidate Standard | [49 empty] | 0% | -100% | -100%
2.9337
95% Uncertainty |  | Portion of Uncertainty2 Due to Calibration Uncertainty
0.7896% |  | 45.66%
95% Confidence Limits for Candidate Standard Concentration
Lower | Upper
2.9105 | 2.9569
These upper and lower limits are compared with the corresponding
limits estimated on different assay dates to establish that the
candidate standard has not drifted.

## Chart Data
This sheet contains the chart data. The chart data include initial calibration data and the
regression information derived from those data.
Only the initial calibration data (Step 1 in Measurement Data worksheet) are included.
| Recalibration data (Step 6) are not included. | [5 empty] | 0.95 | = Confidence Level for Plotting Confidence Bands
0 | = minimum concentration | [5 empty] | 2.1098 | = t for linear model
9 | = maximum concentration | [5 empty] | 2.1199 | = t for quadratic model
19 | = n, the number of initial calibration points | [5 empty] | 2.1314 | = t for cubic model
[7 empty] | 2.1448 | = t for quartic model
Initial Calibration Data |  |  |  |  | Estimates, Linear Model |  |  |  | Estimates, Quadratic Model |  |  |  | Estimates, Cubic Model |  |  |  | Estimates, Quartic Model
i | Xi | Yi |  |  | Estimate | Lower | Upper | Residual | Estimate | Lower | Upper | Residual | Estimate | Lower | Upper | Residual | Estimate | Lower | Upper | Residual
1 | 0 | 0.2194 | [5 empty] | 1.2972 |  |  |  | 0.022979 |  |  |  | 0.024229 |  |  |  | -0.0011643
2 | 0.5 | 0.7141 | [5 empty] | 0.8416 |  |  |  | -0.0078507 |  |  |  | -0.0074339 |  |  |  | -0.0046125
3 | 1 | 1.2885 | [5 empty] | 0.46574 |  |  |  | -0.0089477 |  |  |  | -0.0091193 |  |  |  | 0.0069798
4 | 1.5 | 1.9132 | [5 empty] | 0.14019 |  |  |  | -0.0097122 |  |  |  | -0.010258 |  |  |  | 0.0085384
5 | 2 | 2.591 | [5 empty] | -0.13226 |  |  |  | -0.0073442 |  |  |  | -0.0080797 |  |  |  | 0.0066086
6 | 2.5 | 3.2866 | [5 empty] | -0.38692 |  |  |  | -0.037144 |  |  |  | -0.037916 |  |  |  | -0.030945
7 | 3 | 4.1078 | [5 empty] | -0.51597 |  |  |  | 0.0086889 |  |  |  | 0.0080025 |  |  |  | 0.0062598
8 | 3.5 | 4.9446 | [5 empty] | -0.62942 |  |  |  | 0.020154 |  |  |  | 0.019645 |  |  |  | 0.010227
9 | 4 | 5.8145 | [5 empty] | -0.70978 |  |  |  | 0.014752 |  |  |  | 0.014482 |  |  |  | -0.00012318
10 | 4.5 | 6.723 | [5 empty] | -0.75153 |  |  |  | -0.0020181 |  |  |  | -0.0020181 |  |  |  | -0.018449
11 | 5 | 7.7284 | [5 empty] | -0.69639 |  |  |  | 0.028144 |  |  |  | 0.028414 |  |  |  | 0.013809
12 | 5.5 | 8.7566 | [5 empty] | -0.61844 |  |  |  | 0.031139 |  |  |  | 0.031648 |  |  |  | 0.022229
13 | 6 | 9.8013 | [5 empty] | -0.52399 |  |  |  | 0.00066681 |  |  |  | 0.0013533 |  |  |  | -0.00038941
14 | 6.5 | 10.8818 | [5 empty] | -0.39375 |  |  |  | -0.043973 |  |  |  | -0.043201 |  |  |  | -0.03623
15 | 7 | 12.0982 | [5 empty] | -0.1276 |  |  |  | -0.0026811 |  |  |  | -0.0019456 |  |  |  | 0.012743
16 | 7.5 | 13.3122 | [5 empty] | 0.13615 |  |  |  | -0.013756 |  |  |  | -0.013211 |  |  |  | 0.0055851
17 | 8 | 14.584 | [5 empty] | 0.45769 |  |  |  | -0.016999 |  |  |  | -0.016828 |  |  |  | -0.00072859
18 | 8.5 | 15.9238 | [5 empty] | 0.84724 |  |  |  | -0.0022097 |  |  |  | -0.0026265 |  |  |  | 0.00019503
19 | 9 | 17.3271 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
20 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
21 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
22 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
23 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
24 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
25 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
26 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
27 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
28 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
29 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
30 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
31 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
32 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
33 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
34 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
35 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
36 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
37 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
38 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
39 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
40 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
41 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
42 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
43 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
44 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
45 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
46 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
47 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
48 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
49 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
50 | -999 | -999 | [5 empty] | 1.3003 |  |  |  | 0.026112 |  |  |  | 0.024862 |  |  |  | -0.00053133
[21 empty] | x'V for Linear |  | x'V for Quadratic |  |  | x'V for Cubic |  |  |  | x'V for Quartic
1 | 0 | 0 | 0 | 0 | -1.0778 | -1.7351 | -0.42036 |  | 0.19642 | 0.16738 | 0.22546 |  | 0.19517 | 0.17064 | 0.21971 |  | 0.22056 | 0.19967 | 0.24145 |  | 0.097088 | -0.015744 | 0.00018765 | -8.1049e-05 | 7.3017e-06 | 0.00028244 | -0.00022757 | 4.9121e-05 | -3.0977e-06 | 0.00020347 | -0.00025914 | 9.7392e-05 | -1.4008e-05 | 6.8146e-07
1 | 0.18 | 0.0324 | 0.005832 | 0.0010498 | -0.73566 | -1.374 | -0.097346 |  | 0.37986 | 0.353 | 0.40671 |  | 0.37894 | 0.35769 | 0.40018 |  | 0.39215 | 0.37538 | 0.40893 |  | 0.094254 | -0.015114 | 0.0001733 | -7.2253e-05 | 6.3925e-06 | 0.00024305 | -0.00018007 | 3.7166e-05 | -2.2795e-06 | 0.0001599 | -0.00016953 | 5.7212e-05 | -7.6659e-06 | 3.5469e-07
1 | 0.36 | 0.1296 | 0.046656 | 0.016796 | -0.39357 | -1.013 | 0.22589 |  | 0.56977 | 0.54494 | 0.59459 |  | 0.56914 | 0.55063 | 0.58765 |  | 0.57251 | 0.55853 | 0.5865 |  | 0.09142 | -0.014484 | 0.00015942 | -6.3791e-05 | 5.5204e-06 | 0.00020674 | -0.00013693 | 2.6396e-05 | -1.5463e-06 | 0.00012216 | -9.4402e-05 | 2.4103e-05 | -2.4979e-06 | 9.0475e-08
1 | 0.54 | 0.2916 | 0.15746 | 0.085031 | -0.051478 | -0.65233 | 0.54937 |  | 0.76615 | 0.74319 | 0.78911 |  | 0.76579 | 0.74945 | 0.78213 |  | 0.76142 | 0.74902 | 0.77382 |  | 0.088586 | -0.013855 | 0.00014602 | -5.5663e-05 | 4.6855e-06 | 0.00017339 | -9.7977e-05 | 1.6763e-05 | -8.946e-07 | 8.9786e-05 | -3.2493e-05 | -2.5896e-06 | 1.6089e-06 | -1.1733e-07
1 | 0.72 | 0.5184 | 0.37325 | 0.26874 | 0.29061 | -0.2919 | 0.87313 |  | 0.96901 | 0.94775 | 0.99028 |  | 0.96888 | 0.95416 | 0.9836 |  | 0.95865 | 0.9469 | 0.9704 |  | 0.085752 | -0.013225 | 0.00013308 | -4.7869e-05 | 3.8876e-06 | 0.0001429 | -6.3043e-05 | 8.2187e-06 | -3.2082e-07 | 6.2331e-05 | 1.743e-05 | -2.3498e-05 | 4.7626e-06 | -2.7461e-07
1 | 0.9 | 0.81 | 0.729 | 0.6561 | 0.63271 | 0.068225 | 1.1972 |  | 1.1784 | 1.1586 | 1.1981 |  | 1.1784 | 1.1648 | 1.192 |  | 1.164 | 1.1523 | 1.1757 |  | 0.082918 | -0.012595 | 0.00012062 | -4.0409e-05 | 3.1269e-06 | 0.00011516 | -3.1962e-05 | 7.1612e-07 | 1.7857e-07 | 3.9366e-05 | 5.6543e-05 | -3.9224e-05 | 7.0665e-06 | -3.87e-07
1 | 1.08 | 1.1664 | 1.2597 | 1.3605 | 0.9748 | 0.42803 | 1.5216 |  | 1.3942 | 1.3757 | 1.4126 |  | 1.3944 | 1.3814 | 1.4074 |  | 1.3773 | 1.3655 | 1.3891 |  | 0.080085 | -0.011965 | 0.00010864 | -3.3282e-05 | 2.4032e-06 | 9.006e-05 | -4.5657e-06 | -5.7929e-06 | 6.0712e-07 | 2.0476e-05 | 8.5976e-05 | -5.0345e-05 | 8.6191e-06 | -4.5988e-07
1 | 1.26 | 1.5876 | 2.0004 | 2.5205 | 1.3169 | 0.78746 | 1.8463 |  | 1.6165 | 1.5992 | 1.6337 |  | 1.6168 | 1.6041 | 1.6295 |  | 1.5983 | 1.5863 | 1.6102 |  | 0.077251 | -0.011336 | 9.7125e-05 | -2.649e-05 | 1.7167e-06 | 6.7493e-05 | 1.9313e-05 | -1.1356e-05 | 9.6836e-07 | 5.2673e-06 | 0.00010681 | -5.741e-05 | 9.5145e-06 | -4.9834e-07
1 | 1.44 | 2.0736 | 2.986 | 4.2998 | 1.659 | 1.1465 | 2.1715 |  | 1.8452 | 1.8289 | 1.8616 |  | 1.8457 | 1.8331 | 1.8584 |  | 1.8268 | 1.8148 | 1.8388 |  | 0.074417 | -0.010706 | 8.6084e-05 | -2.0032e-05 | 1.0673e-06 | 4.735e-05 | 3.9842e-05 | -1.6022e-05 | 1.2658e-06 | -6.6397e-06 | 0.00012006 | -6.0943e-05 | 9.8417e-06 | -5.0722e-07
1 | 1.62 | 2.6244 | 4.2515 | 6.8875 | 2.0011 | 1.5051 | 2.497 |  | 2.0805 | 2.0649 | 2.0961 |  | 2.0811 | 2.0683 | 2.0938 |  | 2.0628 | 2.0508 | 2.0747 |  | 0.071583 | -0.010076 | 7.5517e-05 | -1.3908e-05 | 4.5494e-07 | 2.9522e-05 | 5.7189e-05 | -1.9837e-05 | 1.5031e-06 | -1.5606e-05 | 0.00012671 | -6.1439e-05 | 9.6852e-06 | -4.9112e-07
1 | 1.8 | 3.24 | 5.832 | 10.4976 | 2.3432 | 1.8632 | 2.8231 |  | 2.3222 | 2.3071 | 2.3372 |  | 2.3229 | 2.3099 | 2.3358 |  | 2.3059 | 2.2942 | 2.3176 |  | 0.068749 | -0.0094464 | 6.5423e-05 | -8.1178e-06 | -1.2026e-07 | 1.3902e-05 | 7.1521e-05 | -2.285e-05 | 1.6837e-06 | -2.1976e-05 | 0.00012769 | -5.9368e-05 | 9.1248e-06 | -4.5435e-07
1 | 1.98 | 3.9204 | 7.7624 | 15.3695 | 2.6853 | 2.2208 | 3.1497 |  | 2.5704 | 2.5557 | 2.585 |  | 2.5711 | 2.558 | 2.5842 |  | 2.5562 | 2.5448 | 2.5676 |  | 0.065915 | -0.0088166 | 5.5803e-05 | -2.6616e-06 | -6.5836e-07 | 3.8175e-07 | 8.3005e-05 | -2.5109e-05 | 1.8111e-06 | -2.6077e-05 | 0.00012387 | -5.5172e-05 | 8.2356e-06 | -4.0097e-07
1 | 2.16 | 4.6656 | 10.0777 | 21.7678 | 3.0273 | 2.5778 | 3.4769 |  | 2.825 | 2.8106 | 2.8395 |  | 2.8258 | 2.8126 | 2.839 |  | 2.8133 | 2.8023 | 2.8244 |  | 0.063081 | -0.0081869 | 4.6655e-05 | 2.4606e-06 | -1.1593e-06 | -1.1148e-05 | 9.181e-05 | -2.6661e-05 | 1.889e-06 | -2.8219e-05 | 0.00011607 | -4.9267e-05 | 7.0878e-06 | -3.3477e-07
1 | 2.34 | 5.4756 | 12.8129 | 29.9822 | 3.3694 | 2.9342 | 3.8047 |  | 3.0862 | 3.0718 | 3.1006 |  | 3.087 | 3.0737 | 3.1002 |  | 3.0773 | 3.0666 | 3.0879 |  | 0.060247 | -0.0075571 | 3.7981e-05 | 7.2488e-06 | -1.6232e-06 | -2.0796e-05 | 9.8103e-05 | -2.7554e-05 | 1.9208e-06 | -2.8694e-05 | 0.00010508 | -4.2041e-05 | 5.7472e-06 | -2.5929e-07
1 | 2.52 | 6.3504 | 16.003 | 40.3276 | 3.7115 | 3.2898 | 4.1333 |  | 3.3538 | 3.3394 | 3.3682 |  | 3.3546 | 3.3413 | 3.3678 |  | 3.3479 | 3.3377 | 3.3582 |  | 0.057413 | -0.0069274 | 2.9779e-05 | 1.1703e-05 | -2.05e-06 | -2.867e-05 | 0.00010205 | -2.7836e-05 | 1.9101e-06 | -2.7777e-05 | 9.1608e-05 | -3.3856e-05 | 4.2745e-06 | -1.7781e-07
1 | 2.7 | 7.29 | 19.683 | 53.1441 | 4.0536 | 3.6446 | 4.4626 |  | 3.6279 | 3.6134 | 3.6424 |  | 3.6286 | 3.6155 | 3.6418 |  | 3.6252 | 3.6152 | 3.6352 |  | 0.054579 | -0.0062976 | 2.2051e-05 | 1.5823e-05 | -2.4396e-06 | -3.4878e-05 | 0.00010382 | -2.7556e-05 | 1.8604e-06 | -2.5729e-05 | 7.6339e-05 | -2.5048e-05 | 2.726e-06 | -9.3336e-08
1 | 2.88 | 8.2944 | 23.8879 | 68.7971 | 4.3957 | 3.9986 | 4.7928 |  | 3.9085 | 3.8938 | 3.9232 |  | 3.9092 | 3.8962 | 3.9222 |  | 3.9089 | 3.8991 | 3.9186 |  | 0.051745 | -0.0056678 | 1.4796e-05 | 1.961e-05 | -2.7922e-06 | -3.953e-05 | 0.00010358 | -2.6759e-05 | 1.7754e-06 | -2.2788e-05 | 5.9892e-05 | -1.5926e-05 | 1.1531e-06 | -8.6128e-09
1 | 3.06 | 9.3636 | 28.6526 | 87.677 | 4.7378 | 4.3517 | 5.1239 |  | 4.1955 | 4.1806 | 4.2104 |  | 4.1962 | 4.1834 | 4.209 |  | 4.1989 | 4.1893 | 4.2086 |  | 0.048911 | -0.0050381 | 8.0144e-06 | 2.3062e-05 | -3.1076e-06 | -4.2732e-05 | 0.0001015 | -2.5496e-05 | 1.6584e-06 | -1.918e-05 | 4.2842e-05 | -6.7699e-06 | -3.9731e-07 | 7.3869e-08
1 | 3.24 | 10.4976 | 34.0122 | 110.1996 | 5.0799 | 4.7037 | 5.4561 |  | 4.489 | 4.4739 | 4.5042 |  | 4.4896 | 4.4771 | 4.5022 |  | 4.4953 | 4.4856 | 4.505 |  | 0.046077 | -0.0044083 | 1.7057e-06 | 2.618e-05 | -3.3859e-06 | -4.4594e-05 | 9.7749e-05 | -2.3812e-05 | 1.5131e-06 | -1.5111e-05 | 2.5711e-05 | 2.1646e-06 | -1.8833e-06 | 1.5188e-07
1 | 3.42 | 11.6964 | 40.0017 | 136.8058 | 5.422 | 5.0546 | 5.7894 |  | 4.789 | 4.7737 | 4.8044 |  | 4.7896 | 4.7773 | 4.8018 |  | 4.7979 | 4.7882 | 4.8077 |  | 0.043244 | -0.0037786 | -4.1299e-06 | 2.8964e-05 | -3.6272e-06 | -4.5225e-05 | 9.2487e-05 | -2.1757e-05 | 1.343e-06 | -1.0772e-05 | 8.9701e-06 | 1.065e-05 | -3.2674e-06 | 2.2347e-07
1 | 3.6 | 12.96 | 46.656 | 167.9616 | 5.7641 | 5.4043 | 6.1238 |  | 5.0955 | 5.0799 | 5.1111 |  | 5.096 | 5.084 | 5.1079 |  | 5.1067 | 5.0968 | 5.1166 |  | 0.04041 | -0.0031488 | -9.4922e-06 | 3.1415e-05 | -3.8313e-06 | -4.4732e-05 | 8.5886e-05 | -1.9378e-05 | 1.1516e-06 | -6.3342e-06 | -6.9575e-06 | 1.8484e-05 | -4.5171e-06 | 2.8692e-07
1 | 3.78 | 14.2884 | 54.0102 | 204.1584 | 6.1062 | 5.7528 | 6.4595 |  | 5.4085 | 5.3927 | 5.4242 |  | 5.4088 | 5.3972 | 5.4205 |  | 5.4215 | 5.4115 | 5.4316 |  | 0.037576 | -0.002519 | -1.4381e-05 | 3.3531e-05 | -3.9983e-06 | -4.3223e-05 | 7.8113e-05 | -1.6723e-05 | 9.4255e-07 | -1.9534e-06 | -2.1701e-05 | 2.5494e-05 | -5.6044e-06 | 3.408e-07
1 | 3.96 | 15.6816 | 62.0991 | 245.9126 | 6.4483 | 6.0999 | 6.7966 |  | 5.7279 | 5.712 | 5.7438 |  | 5.7282 | 5.7167 | 5.7396 |  | 5.7425 | 5.7323 | 5.7527 |  | 0.034742 | -0.0018893 | -1.8798e-05 | 3.5313e-05 | -4.1281e-06 | -4.0808e-05 | 6.9336e-05 | -1.3839e-05 | 7.1931e-07 | 2.2322e-06 | -3.4939e-05 | 3.1531e-05 | -6.5062e-06 | 3.8392e-07
1 | 4.14 | 17.1396 | 70.9579 | 293.7659 | 6.7903 | 6.4456 | 7.1351 |  | 6.0538 | 6.0378 | 6.0698 |  | 6.054 | 6.0427 | 6.0652 |  | 6.0695 | 6.0591 | 6.0798 |  | 0.031908 | -0.0012595 | -2.274e-05 | 3.6761e-05 | -4.2209e-06 | -3.7595e-05 | 5.9723e-05 | -1.0774e-05 | 4.8544e-07 | 6.1016e-06 | -4.64e-05 | 3.6475e-05 | -7.204e-06 | 4.1538e-07
1 | 4.32 | 18.6624 | 80.6216 | 348.2852 | 7.1324 | 6.7899 | 7.4749 |  | 6.3862 | 6.3701 | 6.4023 |  | 6.3863 | 6.3751 | 6.3974 |  | 6.4025 | 6.392 | 6.4129 |  | 0.029074 | -0.00062976 | -2.621e-05 | 3.7876e-05 | -4.2766e-06 | -3.3692e-05 | 4.944e-05 | -7.5772e-06 | 2.4449e-07 | 9.5509e-06 | -5.5866e-05 | 4.0232e-05 | -7.6839e-06 | 4.3452e-07
1 | 4.5 | 20.25 | 91.125 | 410.0625 | 7.4745 | 7.1328 | 7.8163 |  | 6.725 | 6.7089 | 6.7411 |  | 6.725 | 6.7139 | 6.7361 |  | 6.7414 | 6.731 | 6.7519 |  | 0.02624 | 0 | -2.9207e-05 | 3.8656e-05 | -4.2951e-06 | -2.9207e-05 | 3.8656e-05 | -4.2951e-06 | 0 | 1.2493e-05 | -6.3165e-05 | 4.2735e-05 | -7.937e-06 | 4.4094e-07
1 | 4.68 | 21.9024 | 102.5032 | 479.7151 | 7.8166 | 7.4741 | 8.1591 |  | 7.0703 | 7.0543 | 7.0864 |  | 7.0702 | 7.0591 | 7.0814 |  | 7.0864 | 7.076 | 7.0969 |  | 0.023406 | 0.00062976 | -3.173e-05 | 3.9103e-05 | -4.2766e-06 | -2.4249e-05 | 2.7538e-05 | -9.7593e-07 | -2.4449e-07 | 1.486e-05 | -6.8179e-05 | 4.3943e-05 | -7.9588e-06 | 4.3452e-07
1 | 4.86 | 23.6196 | 114.7913 | 557.8855 | 8.1587 | 7.814 | 8.5034 |  | 7.4222 | 7.4061 | 7.4382 |  | 7.422 | 7.4107 | 7.4332 |  | 7.4374 | 7.4271 | 7.4478 |  | 0.020572 | 0.0012595 | -3.3781e-05 | 3.9215e-05 | -4.2209e-06 | -1.8926e-05 | 1.6253e-05 | 2.3326e-06 | -4.8544e-07 | 1.6597e-05 | -7.0839e-05 | 4.3844e-05 | -7.7498e-06 | 4.1538e-07
1 | 5.04 | 25.4016 | 128.0241 | 645.2413 | 8.5008 | 8.1525 | 8.8492 |  | 7.7804 | 7.7645 | 7.7963 |  | 7.7801 | 7.7687 | 7.7916 |  | 7.7944 | 7.7842 | 7.8047 |  | 0.017738 | 0.0018893 | -3.5358e-05 | 3.8993e-05 | -4.1281e-06 | -1.3347e-05 | 4.97e-06 | 5.5825e-06 | -7.1931e-07 | 1.7671e-05 | -7.1125e-05 | 4.2449e-05 | -7.315e-06 | 3.8392e-07
1 | 5.22 | 27.2484 | 142.2366 | 742.4753 | 8.8429 | 8.4895 | 9.1963 |  | 8.1452 | 8.1294 | 8.1609 |  | 8.1448 | 8.1331 | 8.1565 |  | 8.1575 | 8.1475 | 8.1676 |  | 0.014904 | 0.002519 | -3.6462e-05 | 3.8438e-05 | -3.9983e-06 | -7.62e-06 | -6.1448e-06 | 8.7261e-06 | -9.4255e-07 | 1.8063e-05 | -6.907e-05 | 3.9801e-05 | -6.6642e-06 | 3.408e-07
1 | 5.4 | 29.16 | 157.464 | 850.3056 | 9.185 | 8.8252 | 9.5447 |  | 8.5164 | 8.5009 | 8.532 |  | 8.516 | 8.504 | 8.5279 |  | 8.5266 | 8.5168 | 8.5365 |  | 0.01207 | 0.0031488 | -3.7093e-05 | 3.7548e-05 | -3.8313e-06 | -1.8533e-06 | -1.6923e-05 | 1.1716e-05 | -1.1516e-06 | 1.7772e-05 | -6.4756e-05 | 3.5965e-05 | -5.812e-06 | 2.8692e-07
1 | 5.58 | 31.1364 | 173.7411 | 969.4754 | 9.5271 | 9.1597 | 9.8945 |  | 8.8941 | 8.8788 | 8.9095 |  | 8.8936 | 8.8813 | 8.9058 |  | 8.9019 | 8.8922 | 8.9117 |  | 0.0092365 | 0.0037786 | -3.725e-05 | 3.6324e-05 | -3.6272e-06 | 3.8445e-06 | -2.7198e-05 | 1.4503e-05 | -1.343e-06 | 1.6814e-05 | -5.8316e-05 | 3.1035e-05 | -4.7774e-06 | 2.2347e-07
1 | 5.76 | 33.1776 | 191.103 | 1100.7531 | 9.8692 | 9.493 | 10.2454 |  | 9.2783 | 9.2632 | 9.2935 |  | 9.2777 | 9.2652 | 9.2902 |  | 9.2834 | 9.2737 | 9.293 |  | 0.0064026 | 0.0044083 | -3.6935e-05 | 3.4767e-05 | -3.3859e-06 | 9.3651e-06 | -3.6802e-05 | 1.7041e-05 | -1.5131e-06 | 1.5224e-05 | -4.9933e-05 | 2.5132e-05 | -3.5845e-06 | 1.5188e-07
1 | 5.94 | 35.2836 | 209.5846 | 1244.9324 | 10.2113 | 9.8251 | 10.5974 |  | 9.669 | 9.6541 | 9.6839 |  | 9.6683 | 9.6555 | 9.6811 |  | 9.6711 | 9.6614 | 9.6807 |  | 0.0035686 | 0.0050381 | -3.6146e-05 | 3.2875e-05 | -3.1076e-06 | 1.46e-05 | -4.5566e-05 | 1.9281e-05 | -1.6584e-06 | 1.3052e-05 | -3.9839e-05 | 1.8403e-05 | -2.262e-06 | 7.3869e-08
1 | 6.12 | 37.4544 | 229.2209 | 1402.8321 | 10.5534 | 10.1563 | 10.9504 |  | 10.0661 | 10.0514 | 10.0808 |  | 10.0654 | 10.0524 | 10.0784 |  | 10.0651 | 10.0553 | 10.0748 |  | 0.00073472 | 0.0056678 | -3.4885e-05 | 3.065e-05 | -2.7922e-06 | 1.9441e-05 | -5.3325e-05 | 2.1175e-05 | -1.7754e-06 | 1.0365e-05 | -2.832e-05 | 1.1023e-05 | -8.4306e-07 | -8.6128e-09
1 | 6.3 | 39.69 | 250.047 | 1575.2961 | 10.8954 | 10.4865 | 11.3044 |  | 10.4697 | 10.4552 | 10.4843 |  | 10.469 | 10.4558 | 10.4821 |  | 10.4655 | 10.4555 | 10.4755 |  | -0.0020992 | 0.0062976 | -3.315e-05 | 2.809e-05 | -2.4396e-06 | 2.378e-05 | -5.9909e-05 | 2.2676e-05 | -1.8604e-06 | 7.2486e-06 | -1.571e-05 | 3.1916e-06 | 6.3412e-07 | -9.3336e-08
1 | 6.48 | 41.9904 | 272.0978 | 1763.1937 | 11.2375 | 10.8158 | 11.6593 |  | 10.8798 | 10.8654 | 10.8942 |  | 10.879 | 10.8658 | 10.8923 |  | 10.8724 | 10.8621 | 10.8827 |  | -0.0049331 | 0.0069274 | -3.0942e-05 | 2.5197e-05 | -2.05e-06 | 2.7508e-05 | -6.5151e-05 | 2.3736e-05 | -1.9101e-06 | 3.8051e-06 | -2.3924e-06 | -4.8623e-06 | 2.1268e-06 | -1.7781e-07
1 | 6.66 | 44.3556 | 295.4083 | 1967.4193 | 11.5796 | 11.1444 | 12.0149 |  | 11.2964 | 11.282 | 11.3108 |  | 11.2956 | 11.2823 | 11.3089 |  | 11.2859 | 11.2753 | 11.2966 |  | -0.007767 | 0.0075571 | -2.8261e-05 | 2.1969e-05 | -1.6232e-06 | 3.0516e-05 | -6.8885e-05 | 2.4308e-05 | -1.9208e-06 | 1.5363e-07 | 1.1197e-05 | -1.2884e-05 | 3.5874e-06 | -2.5929e-07
1 | 6.84 | 46.7856 | 320.0135 | 2188.8924 | 11.9217 | 11.4722 | 12.3712 |  | 11.7194 | 11.7049 | 11.7339 |  | 11.7186 | 11.7054 | 11.7319 |  | 11.7062 | 11.6951 | 11.7172 |  | -0.010601 | 0.0081869 | -2.5106e-05 | 1.8408e-05 | -1.1593e-06 | 3.2697e-05 | -7.0942e-05 | 2.4342e-05 | -1.889e-06 | -3.5692e-06 | 2.4572e-05 | -2.0593e-05 | 4.9639e-06 | -3.3477e-07
1 | 7.02 | 49.2804 | 345.9484 | 2428.5578 | 12.2638 | 11.7994 | 12.7282 |  | 12.1489 | 12.1342 | 12.1636 |  | 12.1482 | 12.1351 | 12.1613 |  | 12.1333 | 12.1218 | 12.1447 |  | -0.013435 | 0.0088166 | -2.1479e-05 | 1.4512e-05 | -6.5836e-07 | 3.3942e-05 | -7.1155e-05 | 2.3792e-05 | -1.8111e-06 | -7.2099e-06 | 3.7198e-05 | -2.768e-05 | 6.1992e-06 | -4.0097e-07
1 | 7.2 | 51.84 | 373.248 | 2687.3856 | 12.6059 | 12.126 | 13.0858 |  | 12.5849 | 12.5699 | 12.6 |  | 12.5842 | 12.5713 | 12.5972 |  | 12.5673 | 12.5556 | 12.579 |  | -0.016269 | 0.0094464 | -1.7378e-05 | 1.0283e-05 | -1.2026e-07 | 3.4143e-05 | -6.9356e-05 | 2.261e-05 | -1.6837e-06 | -1.0597e-05 | 4.8487e-05 | -3.3811e-05 | 7.2317e-06 | -4.5435e-07
1 | 7.38 | 54.4644 | 401.9473 | 2966.3709 | 12.948 | 12.452 | 13.444 |  | 13.0274 | 13.0118 | 13.043 |  | 13.0268 | 13.014 | 13.0395 |  | 13.0085 | 12.9966 | 13.0204 |  | -0.019103 | 0.010076 | -1.2804e-05 | 5.7191e-06 | 4.5494e-07 | 3.3191e-05 | -6.5378e-05 | 2.0747e-05 | -1.5031e-06 | -1.3544e-05 | 5.7803e-05 | -3.8623e-05 | 7.9951e-06 | -4.9112e-07
1 | 7.56 | 57.1536 | 432.0812 | 3266.534 | 13.2901 | 12.7776 | 13.8026 |  | 13.4763 | 13.46 | 13.4927 |  | 13.4758 | 13.4631 | 13.4885 |  | 13.4569 | 13.4449 | 13.4689 |  | -0.021937 | 0.010706 | -7.7573e-06 | 8.2157e-07 | 1.0673e-06 | 3.0977e-05 | -5.9053e-05 | 1.8156e-05 | -1.2658e-06 | -1.5844e-05 | 6.4459e-05 | -4.1729e-05 | 8.4184e-06 | -5.0722e-07
1 | 7.74 | 59.9076 | 463.6848 | 3588.9205 | 13.6322 | 13.1028 | 14.1616 |  | 13.9317 | 13.9145 | 13.949 |  | 13.9314 | 13.9186 | 13.9441 |  | 13.9128 | 13.9008 | 13.9247 |  | -0.024771 | 0.011336 | -2.2372e-06 | -4.4099e-06 | 1.7167e-06 | 2.7395e-05 | -5.0213e-05 | 1.479e-05 | -9.6836e-07 | -1.7276e-05 | 6.7718e-05 | -4.2711e-05 | 8.4256e-06 | -4.9834e-07
1 | 7.92 | 62.7264 | 496.7931 | 3934.6013 | 13.9743 | 13.4275 | 14.521 |  | 14.3936 | 14.3752 | 14.4121 |  | 14.3934 | 14.3804 | 14.4064 |  | 14.3763 | 14.3645 | 14.388 |  | -0.027604 | 0.011965 | 3.756e-06 | -9.9754e-06 | 2.4032e-06 | 2.2334e-05 | -3.8692e-05 | 1.0599e-05 | -6.0712e-07 | -1.7599e-05 | 6.679e-05 | -4.1129e-05 | 7.9365e-06 | -4.5988e-07
1 | 8.1 | 65.61 | 531.441 | 4304.6721 | 14.3164 | 13.7519 | 14.8808 |  | 14.862 | 14.8423 | 14.8818 |  | 14.8619 | 14.8483 | 14.8756 |  | 14.8475 | 14.8359 | 14.8592 |  | -0.030438 | 0.012595 | 1.0222e-05 | -1.5875e-05 | 3.1269e-06 | 1.5687e-05 | -2.4321e-05 | 5.5376e-06 | -1.7857e-07 | -1.6557e-05 | 6.0839e-05 | -3.6513e-05 | 6.8657e-06 | -3.87e-07
1 | 8.28 | 68.5584 | 567.6636 | 4700.2542 | 14.6584 | 14.0759 | 15.241 |  | 15.3368 | 15.3156 | 15.3581 |  | 15.337 | 15.3223 | 15.3517 |  | 15.3267 | 15.315 | 15.3385 |  | -0.033272 | 0.013225 | 1.7162e-05 | -2.2108e-05 | 3.8876e-06 | 7.3449e-06 | -6.9335e-06 | -4.4346e-07 | 3.2082e-07 | -1.3877e-05 | 4.8974e-05 | -2.8367e-05 | 5.1233e-06 | -2.7461e-07
1 | 8.46 | 71.5716 | 605.4957 | 5122.4939 | 15.0005 | 14.3997 | 15.6014 |  | 15.8182 | 15.7952 | 15.8411 |  | 15.8185 | 15.8022 | 15.8349 |  | 15.8142 | 15.8018 | 15.8266 |  | -0.036106 | 0.013855 | 2.4575e-05 | -2.8676e-05 | 4.6855e-06 | -2.8001e-06 | 1.3639e-05 | -7.3916e-06 | 8.946e-07 | -9.2675e-06 | 3.0257e-05 | -1.6169e-05 | 2.6148e-06 | -1.1733e-07
1 | 8.64 | 74.6496 | 644.9725 | 5572.5628 | 15.3426 | 14.7232 | 15.9621 |  | 16.306 | 16.2811 | 16.3308 |  | 16.3066 | 16.2881 | 16.3251 |  | 16.31 | 16.296 | 16.3239 |  | -0.03894 | 0.014484 | 3.2461e-05 | -3.5577e-05 | 5.5204e-06 | -1.4857e-05 | 3.7564e-05 | -1.5355e-05 | 1.5463e-06 | -2.4198e-06 | 3.6984e-06 | 6.3161e-07 | -7.5922e-07 | 9.0475e-08
1 | 8.82 | 77.7924 | 686.129 | 6051.6575 | 15.6847 | 15.0464 | 16.323 |  | 16.8002 | 16.7734 | 16.8271 |  | 16.8012 | 16.7799 | 16.8224 |  | 16.8144 | 16.7976 | 16.8312 |  | -0.041774 | 0.015114 | 4.082e-05 | -4.2812e-05 | 6.3925e-06 | -2.8933e-05 | 6.5008e-05 | -2.4381e-05 | 2.2795e-06 | 6.9912e-06 | -3.1743e-05 | 2.2611e-05 | -5.1028e-06 | 3.5469e-07
1 | 9 | 81 | 729 | 6561 | 16.0268 | 15.3694 | 16.6842 |  | 17.301 | 17.2719 | 17.33 |  | 17.3022 | 17.2777 | 17.3268 |  | 17.3276 | 17.3067 | 17.3485 |  | -0.044608 | 0.015744 | 4.9652e-05 | -5.0382e-05 | 7.3017e-06 | -4.5138e-05 | 9.6139e-05 | -3.4517e-05 | 3.0977e-06 | 1.9308e-05 | -7.7156e-05 | 5.0371e-05 | -1.0525e-05 | 6.8146e-07