## Sheet1
MINOS CABLE TABLE |  |  |  | 2003-03-17
FAR
Description | From | To | Quantity | Length | Cable Type | Connector | Provider | Labels
High Voltage - M16 | LRS1440 | Mux boxes | 1536 | 4-20m | RG-59 | SHV | TAMU
High Voltage - Monitor PMT | LRS1440 | Mon. PMT box | 4 |  | RG-59 | SHV | TAMU
High Voltage - Trigger PMT | LRS1440 | Trig. PMT box | 4? |  | RG-59 | SHV | TAMU
High Voltage - Control | Ethernet adapter | LRS1440 | 4 or 8 | Short ! | RS-232 | DB 9 ? | TAMU
Frontend - Analog cable | VME readout crate | Frontend board | 486 | ~25' | twisted pair |  | Harvard
Frontend - Low Voltage | Readout crate patch panel | Frontend board | 486 | ~25' | twisted pair |  | Harvard
Frontend - Digital | VME readout crate | Frontend board | 486 | ~25' | 50 cond flat |  | Harvard
Frontend - ground strap | Mux boxes | Readout Crate? | 486 | ~25' | insulated gnd. |  | Harvard
PVIC - Copper | VME readout crate | VME readout crate | 4 | 15-20m |  |  | RAL
PVIC - Optical | VME readout crate | Control Room | 4 | <200m | Optical Fiber |  | RAL
Ethernet | Hubs | Everywhere | 3? |  |  |  | FNAL?
GPS antenna feed | Surface GPS antenna | GPS receiver clock | 1 | ~2km | Optical Fiber |  | Oxford
Clock distribution | GPS Clock | VME readout crate | 16 | <200m | Optical Fiber |  | Oxford
Flasher distribution - scintillator | LED boxes | Scintillator manifolds | 19360 |  | Optical Fiber |  | Sussex
Flasher distribution - PIN diode | LED boxes | Frontend board | 640 |  | Optical Fiber |  | Sussex
Flasher distribution - Monitor PMT | LED boxes | Mon. PMT box | 320 |  | Optical Fiber |  | Sussex
Flasher distribution - Trigger PMT | LED boxes | Trig. PMT box | 16 |  | Optical Fiber |  | Sussex
Flasher Control | Ethernet adapter | LED boxes | 8 |  | RS-485 |  | Sussex
Flasher external clock | ? | LED boxes | 8 |  | Coax | BNC
Flasher veto | ? | LED boxes | 8 |  | Coax | BNC
Magnet monitor coils | Steel plate coils | Mag. Monitor PCs | >~500 | <200m |  |  | FNAL
Shunt trip connections | Rack Protection box | Circuit Breaker Panels |  |  |  |  | Installation
Non-powered rack smoke detectors | Frontend racks | Rack Protection box | 60 |  |  |  | DCS group
Environment monitors | Sensors | Field Point or Rack Prot |  |  |  |  | DCS group
Magnet Power supply control/monitor | Power supply | Field Point unit |  |  |  |  | DCS group
AC power cables
NEAR
Description | From | To | Quantity | Length | Cable Type | Connector | Provider | Rating | Comments
High Voltage - M64 | LRS1440 | Alner box | 194 | varies | Belden RG-58/U (pn 82240) | SHV | TAMU | FT6 Plenum rated | Same as used in Far Det
High Voltage - Trigger PMT | LRS1440 | Trig. PMT box | 1 | ? | Belden RG-58/U (pn 82240) | SHV | TAMU | FT6 Plenum rated | Same as used in Far Det
High Voltage - Control | EDAS or RPS Control Box | LRS1440 | 2 | ~5ft | 16-conductor ribbon | ? | UMD | ? | Same as used in Far Det
Alner Box to MINDER Aux - anode | Alner box | MINDER Aux board | 4 types, various quantities | ~ 2ft | Belden-8216 RG-174 | D32 | ANL | UL 1364 | I am unable to find any reference to UL 1364 (or UL 1354, which is the rating given by Belden), but the Belden catalog describes these cables as having polyethelene insulation with PVC jackets. From the standpoint of flammability and smoke, they sound quite undesirable. I recommend finding something else. - DJB
Alner Box (includes Trig PMT) to KEEPER aux - dynode | Alner box | KEEPER Aux board | 195 | ~ 2ft | " " | Lemo | ANL | Same as above? | Same as above - DJB
Optical readout cables | Detector Scint. Modules | Alner box | 573 | varies | optical fiber |  | Sussex | Mitigated by halogen-free fire retardant conduit | Same as used in Far Det
Frontend - Data | MINDER crate | MASTER crate | 584 | 30 ft | multi-pair | 3M 3600B36 | ANL | CL2 | Comment - DJB
MASTER - KEEPER | MASTER Aux | KEEPER board | 44 | 30 ft | multi-pair | 3M 3600B36 | ANL | CL2 | Same comment as above - DJB
PVIC - Copper | MASTER crate | MASTER crate | 4 | ? |  |  | RAL | VW-1 | Same as used in Far Det
PVIC - Optical | MASTER crate | Branch Readout Processor | 4 | <200m | 62.5/125 um OFNR | SC/SC | RAL | NEC 770 | Same as used in Far Det
PVIC - Copper | Branch Readout Processor | Trigger Processor | 8? | 2m |  |  | RAL | VW-1 | Same as used in Far Det
DAQ ROP serial communication | EDAS equivalent device | ROP | 8 | ~ 5ft | RS-232 |  | RAL |  | Same as used in Far Det
Ethernet | Hubs | Everywhere |  |  | Belden 1752A | RJ-45 | FNAL CD Networking | UL 1581 | Same as used in Far Det
Clock distribution - MINDER - 53MHz | Master clock fanout MCF | MTM in MINDER crate | 44 |  | Quabbin 2902 (CAT 5e) |  | PPD EE | NEC CM
Clock distribution - MASTER - 53MHz | Master clock fanout MCF | VTM in MASTER crate | 8 |  | Quabbin 5500 (CAT 5e) |  | PPD EE | NEC CM
Clock distribution - MASTER - GPS | GPS fanout | VTM in MASTER crate | 8 |  | CAT-5 |  | IIT
Clock: GPS to GPS fanout | GPS receiver | GPS fanout | 1 | 1 |  |  | IIT
Clock: GPS to Timing PC | GPS receiver | Timing PC | 1 | 1 | CAT-5 |  | Oxford | CSA CMP | Same as used in Far Det
Fiber in Shaft | Surface | LAN distribution & GPS receiver | 1 bundle | ~400ft | Optical Fiber | none | RBI
LI distribution - scintillator | LED boxes | Scintillator manifolds | 1720 | varies | EHV4001 Optical Fiber |  | Sussex | VW-1 | Same as used in Far Det
LI distribution - PIN diode | LED boxes | PIN Aux | 64 or 66 ? | all the same? | EHV4001 Optical Fiber |  | Sussex | VW-1 | Same as used in Far Det
LI distribution - Trigger PMT | LED boxes | Trig. PMT box | 3 | all the same? | EHV4001 Optical Fiber |  | Sussex | VW-1 | Same as used in Far Det
LI Control | RPS Control Box | LI Crate | 3 | 2 ft | RS-485 |  | Sussex |  | Same as used in Far Det
LI external trigger | Master clock Fanout MCF | LI Crate | 3 | ? | RG-58 Coax | Lemo | FNAL
LI Beam Spill Inhibit | Beam Spill Gate | LI Crate | 3 | ? | RG-58 Coax | Lemo | FNAL
B-dot sensor cables | wrapped on Plates |  | ~400 |  | 50-conductor .025" ribbon cable |  | Uminn | VW-1 | Same as used in Far Det
B-dot coil readout | Connectors on Steel plates | Mag. Monitor PCs | ~400 | <200m | Belden 87761 Twisted Pair |  | UMinn | FT6 Plenum rated | Same as used in Far Det
Environment monitors | Sensors, distributed in Hall | Field Point unit | ? | ? | Belden 8723 2pr |  | DCS group | UL 1581 Vertical Tray | Same as used in Far Det
AC power cables |  | Everywhere |  | 6 ft | Waber Model 602 metal outlet strip |  | FNAL | UL 1363 | Same as used in Far Det
Low Volt Power - in racks | Power supply | Crates | many | varies | various insulated wire |  | FNAL
RPS control - RPS relay box | RPS control | RPS relay box | one per rack | ~ 6ft | Belden 88760 twisted pair |  | DCS group | UL 910 Steiner Tunnel test, FT6 | Same as used in Far Det
RPS in-rack sensors | Smoke detectors, Sensors, fans | RPS control box | many | ~ 6ft | Belden 88760 twisted pair |  | DCS group | UL 910 Steiner Tunnel test, FT6 | Same as used in Far Det
CANbus PS Monitor? | Electronics LVDC PS | Field Point Unit? | 1 set per PS | varies |  |  | ANL?
Note on F5: DJB: The cables actually used in the far detector are RG-58/U, not RG-59. Their flammability rating is FT6 (plenum rating). They are manufactured by Belden, Part #82240.
Note on J55: I'm confused by this line. The designation "multi-pair" suggests that a specific cable has not yet been chosen, yet there is a flammability rating specified. Furthermore, the CL2 rating is for a vertical tray test, but if I'm interpreting the cable diagrams correctly, the Minder and Master crates are on the same level.
Note on J58: Caution! We've had some problems at the far detector with attenuation in this cable. We've had to go to a 6-branch rather than a 4-branch readout, at least temporarily, because the true capability of the cable didn't measure up to the vendor's claim. You should check with the RAL people to see whether you really want to depend on a 200 m length.