MWINILUNGA – KASAMA AND LIVINGSTONE

ZAMBIA 10TH – 20TH SEPTEMBER 2006

Background

The main objectives of this trip was to carryout inspection of Meteorological   Instruments, measure and record Meta data, test and calibrate thermometers and       barometers at these stations. The inspection trip started on 10th and ended on 20th September 2006.  The trip covered three GSN stations namely;- Kasama, Livingstone and Mwinilunga. Our companions during this trip were the Chief Meteorologist, Mr Jacob Nkomoki and the Driver Mr Boyd Ngoma.

- Kasama-

  - Location and Establishment.

Kasama was established as far back as 1958 and lies on Latitude 10º 13( South and Longitude 13º 08( east and its height is 1392 meters above Mean Sea Level.

The station is situated on the eastern side of the airport.  The Meteorological Enclosure measures 20.7m X 14m.  The instruments at this meteorological enclosure are well exposed.  We have, however, observed that most of the instruments at this station are obsolete and unserviceable.

# 3.0	Instruments and met enclosure

3.1	Stevenson Screen

The Steven Screen is very old it cannot be repaired or repainted.

  - Wind System

The wind system obtained from China is old and is not working.  The source of wind information is obtained by observing a wind sock and applying the Beaufort Wind Scale to determine wind speed. We believe this is not a good way of making observations in this new era.

  - Thermometers

The available thermometers were carefully inspected and it was noted that the maximum thermometer could no longer be used as its bulb had a crack hence the mercury could not easily flow through its constriction.  All other thermometers compared favorably with GTH 175/Pt calibration thermometer.  The maximum thermometer needs to be replaced.

Two soil thermometers are available i.e. - 10 cm and 30 cm.  It was also not easy to identify their serial numbers. Thermometers for other depths were not available.

3.4	Barograph and Thermo hygrograph

The available Barograph and Thermo hygrograph are very old and are not working.

3.5	Evaporation Pan

The evaporation pan is old and rusty and it is not in use due to serious shortage of water in Kasama area

3.6	Sunshine Recorder

The sunshine recorder is serviceable

3.7	Barometer

The Digital Barometer (PTB220 from Vaisala) is serviceable and compared favorably with the traveling standard.

3.8	Rain gauges

Automatic rain gauge is very old and unserviceable. The 5” copper rain gauge is new and has recently been installed.

3.9	Communication:-

HF radio communication is used for the collection and disseminates weather observations to the main center in Lusaka.  There is however a land line telephone which is used for administrative matters.  It is used as a back up whenever the radio is out of service.

4.0	Mwinilunga

4.1	Establishment of station

Mwinilunga was established in 1958.  It has an altitude of 1369metres above mean sea level.  It is situated on Latitude 11º 44( South and Longitude 24º 26(  east.

4.0	Instruments and Met Enclosure

The Meteorological Enclosure measures 12.07 m x 7.6m.  The grass in the Met enclosure is cut short and the instruments are adequately exposed.  The following is a description of the status of Instruments.

4...1	Rain -gauges:

A new 5 inch copper rain-gauge was installed in July 2006.  The automatic rain-gauge is very old and is no longer used.

4.2	Thermometers Screen:

A small Marine type screen was installed in February 2006.  The standard Stevenson screen is very old and it cannot be repaired nor repainted.

4.3	Thermometers:

Thermometers are new and have been installed in the new Marine type screen.  Soil thermometers are listed in the record form for this station.  The 5cm and 100cm earth thermometers are not available.

4.4	Barometer

The station uses a Digital Barometer PTB220 which compared favorably with the traveling standard.

The Kew Pattern Barometer is also available but this Barometer is very old and it cannot be used as it has its adjusting knob and sight tube broken.

4.5	Barograph

There is a new barograph but this instrument cannot be used as its specifications do not fall within the station height level.  A suitable Barograph has to be procured as a replacement.

4.6	Sunshine Recorder

The sunshine recorder is available and it is currently not mounted as the mounting screws are not available.  It is to be mounted on a steel base on top of a small ant hill situated some 70 metres away on the eastern side of Meteorological enclosure.

4.7	Evaporimeter

A new class evaporation pan made from aluminum has been procured; this pan is a replacement of the old one.  At the time of our inspection the pan had not yet been installed.  The hook gauge is available and it will be installed as necessary.

4.8	Thermo-hygrograph

A new thermo-hygrograph is available it uses a quartz variable clock however, when the clock is set to run for a day it lags behind by about two hours.  This lagging behind of the clock has now rendered unserviceable.  It was noted from the maintenance manual that this type of clock cannot be serviced or repaired and can only be replaced with a new one hence it is suggested that a new one should be bought.  Details of this thermo-hygrograph are listed on the inspection form which is attached as appendix to this report.

4.9	Cup Counter Anemometer

The Cup Counter Anemometer is currently being used to compute wind speed and a wind vane is used to indicate wind direction.  This is a bit cumbersome for the observer.  It is strongly recommended that a conventional wind system be purchased for this station.

4.10	Communication

An HF Radio telephone is used to transmit weather observations to the main center in Lusaka.  The station has neither phone nor computer and as a result computation of climate reports are done manually.  A suitable communication system is needed at this station.  Replacement of all other instrument may not meet the objective of this project without the provision of a suitable communication system to inject the observation into the GTS.

5.0	Livingstone

5.1	Location

This station is situated at Livingstone Airport on the eastern side of the terminal building. A temporary Met enclosure had been established as a result of construction work going on for the extension of the run way and apron.  Meta data measurements were taken at the temporary met enclosure and half way along the runway.  The readings at the two sites as read from the GPS indicated that the Latitude of the station is 17 º49( Longitude is 25º 49( and the altitude is 991 meters above mean sea level.

6.0	Meteorological Instruments

A few conventional instruments are temporary located at the temporary met enclosure site.  These instruments are the 5” copper rain gauge, an old fiber glass an evaporation pan with a hook gauge and earth thermometers.

6.1	Manual Rain gauge

The rain gauge is new and was installed in June 2006.

6.2	Automatic Rain gauge

This is the typing bucket type.  It is one of the components of the Automatic Weather Station installed half way along the run way.  However this Automatic Weather Station belongs to the National Airports of Zambia.

6.3	The Thermometer Screen

The Stevenson Screen.  It is worn out, a new screen is required.  As indicated above it is installed at the temporary site.

6.4	Thermometers

All thermometers are available.  The Dry, Wet, Maximum and Minimum thermometers compared well with the GTH 175/Pt thermometer when the calibration tests were done.  See details on the attached form.

6.5	Remote Indicating Instruments

An automatic Weather Station comprising of a wind system, pressure, Relative Humidity, Temperature and Radiation sensors has been installed midway along the runway.  The AWS is manufactured by ALMOS in the Netherlands.

6.6	Barometers

Two Barometers are available.  These Barometers are the Aneroid Barometer type M2236/0 manufactured by Mechanism Ltd and the Digital Barometer PPTB manufactured by Vaisala – Finland.  The two Barometers were compared with the traveling standard and were found to be within the acceptable correction factor.  Correction cards are available at the station.

6.7	Barograph

This instrument is not available

  - Sunshine Recoder

The sunshine recorder is available but has not been installed the main reason being that the charts are not available.

6.9	Evaporimeter

This is a class a pan made from fiber glass; it is installed at the temporary site and was not level.

6.10	Thermograph

Not available

6.11	Cloud Search light

The Cloud Search light has been recently installed.  This is a product of ALMOS– in the Netherlands.

6.12	Cup Counter Anemometer

The cup counter anemometer is available. This is type W1200/2 manufactured by Casella.  The instrument is not installed at the temporary site.

It will be installed at the Permanent site once construction work on the extension of the runway and apron is completed.

6.13	Weather Observer Terminal System

The system uses Windows 2000 platform and a custom written software visual base is used as the application software.  The is a product of ALMOS. Model “ALMOS Met 12/24 modem S/N 2061” Rom version 2.04. The system was installed on 5th April 2005.

- Performance of the GSN Stations.

Our daily monitoring of the synops and metars  indicates that all the GSN stations in Zambia are silent. This is mainly due to obsolete instruments and none availability of suitable communications systems at most of these stations. A Performance Chart is attached as an appendix to this report to  illustrate this.

- Conclusion

The GSN Stations we visited in Zambia needs to be attended to promptly. Although we found that some instruments could be used their state is totally below the expected international standard and worse still below the WMO Recommended Standards .The available instruments are obsolete and most of them are not being used. HF communication is used and has proved to be unreliable, especially during rainfall season when static conditions are high. Maintenance Personnel are not adequately trained as a result it takes them too long to restore instruments and communication equipment to normal operation calibration of instruments is not frequently done.

9.0	Recommendations

The inspection of three GSN Stations at Kasama, Mwinilunga and Livingstone revealed that most of the instruments installed at these stations are old and obsolete.  Kasama being the worst as almost all instruments are unserviceable.  In this regard, it is strongly recommended that the instruments listed below are recommended for replacement with immediate effect.

There is a need to improve communication for the transmission of observations to the main center in Lusaka so that it could be injected into the GTS with ease.  Telephone facilities are available at these places.  What is required is to procure PC’s with suitable software and link them through dial up circuits using modems (Wide Area Network).  This is one of the cheapest means of communication that can be explored.

It is strongly recommended that consideration be made for the provision of an up to date means of communication preferably a system using satellite telemeter. It was also noted that calibration of instruments is rarely done. In this regard it is strongly recommended that basic calibration instruments like the electrical thermometer similar to one we used could be made available to the National Met of Zambia and many others like the digital barometer and others which are available in the market these instruments could easily be sent to the CSIR for calibration at reasonable cheap prices. In this way the quality of data could be improved.

9.1      Instruments Recommended for Replacement.

(i) 	3 x Stevenson Screens

(ii)	 2 x Wind Systems

- 3 x  Barographs

(iv)      3 x Thermo-hygrographs

(v)	1 x Max Thermometer

- 2 x Evaporation Pans

- 3 x Automatic rain gauge

- Calibration instruments.