# MEMORANDUM

To:		Distribution

From:		F. Dylla/grn

Subject:	FEL Upgrade Project Weekly Brief  - May 1-5, 2006

Date:		May 5, 2006

Highlights:

We spent the week researching noise in the system to determine possible sources of efficiency degradation.  A large amount of high current/high power lasing was performed and we delivered over 100 coulombs out of two cathode cesiations.  Most of the data taking centered on phase and amplitude noise in the drive laser, buncher, and injector output.  We also attempted to look at the phase noise of the synchrotron light at the wiggler but had insufficient signal to detect.  That data is still under analysis and we will need more tests before we are convinced of its role in the loss of efficiency.  On evening shifts we took a complete set of FEL performance data to compare the lasing at the fundamental, third harmonic, and fifth harmonic.   In all cases the lasing was at 1.05 micron; the wiggler was adjusted so the harmonic was at this wavelength with the beam energy fixed.  We also provided beam time for characterization of the THz beam noise.  The THz FTIR interferometer was commissioned for use with CW beam.

Management:

We presented several talks at the Beam Instrumentation Workshop in Chicago which were well received.  We seem to be ahead of most groups in implementing embedded control in Beam Position Monitors.   This will provide a cheaper, more accurate measurement and greater capabilities in the future.

We are also spending a significant amount of effort planning for our maintenance down time May 15-23.   During this time we will accomplish a host of things: a injector beamline bake out, OCCMs system around the wiggler, installation of the Helium transfer line for the cryomirror demonstration, new 1.6 and 1 micron low-loss outcouplers, halo monitors, extra viewers around the wiggler, and a host of other minor activities.  Depending on the results of our present noise studies there may be additional electronic modifications.

We received the written report from our Operations Review on April 27, 2006.  We'd like to express our appreciation to the review committee Erik Johnson, Jon Cooper, Ben Feinberg, James Tarpinian who worked very hard to produce an exceptional study in the limited time they had available.  This review set the stage for us to become a recognized DOE User facility.

Operations:
   Much of our I&C group was at the Beam Instrumentation Workshop in Chicago so we were a bit shorthanded this week.  The main goal this week was to make measurements on the electron beam to characterize noise that might be reducing performance at high current. 
   On Monday we re-cesiated and tried for a power push at 2.8 microns.   The efficiency was not quite as high as before and we got up to 6.3  kW at 6 mA.  We could get 5.5 kW at 5 mA.  We found a technique that allows the laser operator to maintain the mirrors in an optimum alignment and maintain the efficiency for as long as the laser is on.  We were then able to hold the laser between 5 and 5.5 kW for about one half an hour using this technique.  The power cycles due to  a slow drift in the cavity length due to absorption of scattered  light.  As a precursor to a test plan later in the week we lased briefly at 0.98 microns by changing the resonant wavelength to 4.9  microns in the wiggler.  There is no cavity Q at this wavelength or the third harmonic of this wavelength but there is a good cavity Q at the 5th harmonic so it lased there with a 4 micron detuning curve.   We also measured the spectrum of amplitude noise on the drive laser and found a relatively strong oscillation at 38.7 KHz.  This might be the source of a strong phase modulation seen on the electron beam at high current. 
   On Tuesday we worked on measuring the energy vs. time in the 10-100 kHz frequency range in the injector.  Since we had not done this before and most of our instrumentation group was at the Beam Instrumentation Workshop progress was a bit slow but by day's end we had some measurements of the energy jitter.  Nothing convincing was seen above the noise floor.  We then measured the drive laser phase noise before and after the electro-optic cells and with an internal  oscillator or with the RF control module.  One conclusion of this was what the RF control module reduced noise at low frequencies but increases the noise at frequencies above a few hundred Hz.  It also looks like some of the spectral lines might be intensified by the EO cells but it was not clear by how much.  We then switched over to RF tests in zone 3. 
   Wednesday the line at 38.7 kHz was quite small so we decided to see if the laser performance improved at higher current.  In short, it did not.  We then switched to the 1.06 micron mirrors and took data while lasing at the fundamental, third harmonic, and fifth harmonic.   The performance agreed pretty well with expectations except for slightly lower power than expected for the 5th harmonic.  The power was lower yet more stable at the third harmonic than at the  fundamental. This might be a better way to operate for the LMES experiments.  On Thursday we took a complete set of data looking at energy fluctuations in the injector.  We tried to measure the phase of the spontaneous radiation of the electron beam but were unable to see the signal.  The power was less than when we last measured it so we might have been below the noise floor of the detector. 
   Friday is a down day.  We will be working on trying to find the leak in the gun region and starting a heat clean of the gun.

Beam Physics (WBS 3):

Given progress on phase/energy jitter measurements, we are performing an analysis to understand how - and with what sensitivity - longitudinal rattle in the injector will evolve through the accelerator. Initial estimates based on an extension of considerations in JLab TN 01-048  ( http://www.jlab.org/~douglas/FEL/technote/JLABTN01048.pdf ) indicates that phase jitter out of the cryounit is irrelevant (no surprise - its a parallel to point longitudinal focus from unit to wiggler) but that energy jitter gives longitudinal positional jitter of order 0.01 m/radian of relative energy offset - i.e., 

dl (m) = 0.01 m (dE/E). 

Thus, 0.1% energy jitter will give 10 microns longitudinal motion - or about 30 fsec. This would also correlate with about 0.4 mm of motion at the peak dispersion of 0.4 m in the injection line.

WBS 4 (Injector):

We re-cesiated the cathode on Saturday and again on Wednesday. Since March 20th the cathode has been re-cesiated 7 times and has delivered close to 900 Coulombs and 225 hours of CW and pulse beam for FEL ops. We have been studying the vacuum conditions in the injector and today we did a leak test using the RGA in the gun chamber. We did not find any leaks. The Gun RGA spectrum shows only Hydrogen and traces of He when the gun valve is open. The lightbox RGA shows traces of Methane (mass peak 16) that rises and then lowers when the gun valve is cycled. By tapping in each ion pump in the system, we could only see any evidence of methane increase in the pump next to the cryounit. The thorough vacuum test we performed today indicates there is no leak in the gun, therefore the question of reduced cathode lifetime remains unanswered. 

This afternoon we started a heat clean cycle today to make cathode # 6 on the anodized wafer installed almost two years ago. Since then this wafer has delivered over 3,000 Coulombs.

WBS 6 (RF Systems)

Last Friday we met to discuss the buncher cavity resonance control system.  Several potential modifications to the control algorithm (i.e. RF ON/OFF operating modes, water valve setpoints, etc.)  It was agreed that they would replace the water control valve at the May 15th down and install a better thermal blanket on the cavity during or before the down.  Additionally, it was agreed that we would take parasitic data on the control loop parameters so that we could better understand the transitions between the RF on and off states.  This data was taken early this week.  The matter will again be discussed today.

In addition to attending the Beam Instrument Workshop at FERMI Lab, I assisted Richard Evans in taking and analyzing the time and frequency domain data which indicates the energy jitter in the injector.  Analysis of the data continues.

WBS 8 (Instrumentation):

With the 2006 Beam Instrumentation Workshop taking place at Fermi Lab this past week, we got the chance to exchange some valuable information with other instrumentation and diagnos-tics people from all over the world.  A poster and paper were presented on the Embedded BPM design and it was very well received.  There were several other labs trying to implement the same sort of design with the RTEMS open source operating system loaded directly into the onboard IOC.  With other groups working on the same goal some valuable contacts were made so resources and information can all be shared.  We also picked up some ideas for how to measure other interesting beam parameters that might be useful for further machine development.
    An entire battery of low frequency noise studies of the machine have dominated this weeks’ activities. Our primary goal this week was to measure the energy jitter in the injector region as evidenced by the horizontal beam position jitter at the 0F06 BPM (IPM0F06). The data for this measurement was collected. Several specific frequencies were observed at various beam currents, namely, the 39kHz noise that is the motivation of our studies. The measurement was made possible by the new embedded BPM electrics that have been developed over the past few months. We have received a model 350-160 "wet" EO cell from ConOptics. It is being installed for testing today, the results will be known over the week-end. This is in a further attempt to track down real or imaginary noise on the drive laser.

The Halo Monitor stepper control boxes were installed onto the linear stages this week.  We exercised both the horizontal and vertical forks in and out of the cross and checked for any interferences.  Also, all limit switches were set and tested. The stepper motor driver cables have been installed, channels assignments have been made but the EPICS screens will not be updated until our next scheduled down. The Halo cross was delivered to the Vacuum group for leak testing and bake out. 
    The LPSS installation for Lab 4 is going smoothly with the crash interlock chain being wired in.  An additional card reader was also wired in and tested with the entry software successfully. Time this week was spent on constructing the LPSS master box that will go in the control room. The hardware that we had in house was installed and the necessary parts were ordered so wiring can begin to take place next week. The auto-miniphase was exercised some this week.  The Linac gang phase is still not quite right.  The old algorithm was much more accurate, but slower, and is being put back in. We continue working on the FPGA programming for the General Purpose board. The code was modified for ADC sampling, storage and Coldfire bus bridge. Simulations were done for all functional modules. 
    Fabrication of a new stepper motor driver box for optics began this week; the driver box will be wired to run off of AC instead of the usual battery/dc power.
    In preparation of the absence of the Safety Warden the next two weeks several staff were trained to perform the daily checkoffs.  All safety issues are being forwarded to Ned Walker with EH&S. Wes as a member of the Safety Warden Advisory Team, met with Andrew Hutton to discuss the details of the team's formalized safety warden (SW) program.  After he approved the team held a meeting with all accelerator SWs and their supervisors.  Final revisions to the program are being made.

WBS 9 (Beam Transport):

# IR Machine Re-commissioning and Operations

•	I moved the Camera that is able to view delamination of the GW pole tips to the fourth magnet of the Optical Chicane.  No delamination was seen in the lower pole of the third GW magnet of the optical chicane.

• 	We continue working on using four uninstalled GW dipoles in a quick experiment that uses laser light to look for a pseudoscalar particle that couples to photons.  Budget problems prevented us from measuring our test dipole

•	The cover for the W55 (IR Line) Wiggler remains in manufacture.

• 	I continued working on the design of the permanent cover for the optical transport between the drive laser line and the entrance into the Light Box in the injector.

•  	Dano Oprisko is working on the technician portions the helium system for the lower portion of the waveguides for the Cryomodule in Zone 3.  Our plan is to install one waveguide system during the shut down of May 15.

•	I continued work on the gun for the new injector test stand.  We resolved that the wisest course is to not include a puck exchanger into the new stalk interchange system.  The experience of the CEBAF source group that heat cleaning the puck through a gap in vacuum is may cause more problems than a puck is worth.

WBS 11 (Optics):

#### FEL mirrors

As noted in the Operations report we operated for extended periods of time at 2.8 microns.  The new shields work quite well.  We do observe a relatively rapid change in cavity length as some component heats and expands.  We are monitoring our temperature readbacks to try and determine where the problem is and fix it.  Absorption of power (and thus expansion) of the fixed clips has been proposed and we will look at ways to check that during the upcoming down.  We did more measurements of the absorption of THz radiation by the outcoupler mirrors.  It amounts to about 2W/mA.  Our new 1.6 micron outcouplers arrived and are being QA'ed.  We expect the 1.06 micron OCs early next week.  We plan to braze both into holders.

Other Activities

Optical measurements to resolve the cause of the laser efficiency continue. We used a Stanford Instruments spectrum analyzer and Thorlabs photodiode (~ 200 MHz bandwidth) to look at amplitude noise in the 10-100 kHz region.  We did detect a very weak signal at 38.7 kHz.  We devoted a shift to attempts to measure the phase noise of the wiggler spontaneous radiation but were unsuccessful, due to lower power than measured last week.  Apparently we were closer to the noise floor of the photodiodes than previously thought.  More careful measurements of that threshold will be made with the advanced drive laser's oscillator, and the bias circuit modified as necessary to get the desired performance.  We also measured the phase noise spectrum as a function of rep rate after the EO cells.  While there was essentially no difference in the noise spectrum with rep rate, there was a difference if one looked at the noise of the laser before the EO cells vis-a-vis after the EO cells, especially in the 40 kHz region.  This will be studied further next week.  The streak camera was set up to look at synchrotron light emitted from the first arc, however results are too preliminary at this time.  With many I&C personnel at a conference, we helped Ops with phase noise measurements from beam position monitor (BPM) signals.  Parts for the advanced drive laser integration are beginning to arrive.  Our cryocooled mirror in-vessel loop design progressed this week.  We held further meetings with the Cryo Group on their portion of the hardware, and still plan on having them install this hardware the week of May 15.  After looking at the crowded space where the distribution can was going to be placed, they made the decision to move it.  They should still make the installation schedule, but this incurs some cost.  The Machine Shop continues work on our projects, optical cavity mirror metrology (OCMMS), in-lab optical transport system hardware (Brewster window transition, hutch shutter, & harmonic-blocking filter), and now the User shutter for User Lab 6.

Terahertz:

The THz studies this week were dedicated to further characterization of the new spectrometer and data acquisition software, as well as more noise measurements of the FEL THz emission.  The noise measurements revealed some interesting and surprising periodic noise spikes that were found to be caused by the spectrometer control software and “leaching” through the integral detector electronics where they could be detected with the spectrum analyzer.  It was also found that there was a significant amount of noise passing onto the electronics of the integral detector from the drive electronics of the moveable mirror in the spectrometer.  This could be eliminated by using an external detector, but this detector appears to have a higher noise floor and will need to be compared in detail with the integral detector. 

The spectrometer characterization studies were conducted during dedicated CW beam operations Wednesday night and a few important details were learned.  While varying the scan length of the spectrometer did not affect the data acquisition and curve fitting, the best compromise between resolution and the amount of data to be passed out through the ethernet data socket was found to be at a spectrometer resolution of 1 cm-1.  It was also found that the spectrometer scan velocity had an effect on the spectrum measured by the DAQ software if set too high. This velocity induced error was evident at scan velocities greater than 1.25 cm/sec.  There are still further characterization studies planned, but we have found a range of settings that provide consistent spectrum calculations and the spectrometer itself has been performing reliably all week. 

Another important event that have occurred this week was the final senior thesis presentation made by the William and Mary Physics student Amelia Greer, who was advised and mentored by Gwyn Williams.  Both her written thesis and oral presentation received high praise and excellent reviews from her committee and others in attendance.  Her work helped to bring the THz beamline online as well as to begin characterization of the new spectrometer.  Finally, Wade Brock, a rising senior in Physics at ODU, has proposed to perform beam divergence studies on the THz beam in Lab 3 for his senior thesis.  His work will begin soon and will involve experimental measurements of the beam divergence in Lab 3 to be compared with the calculated divergence.  This will help us to determine how to properly collimate the beam and design an optic to achieve this goal.