SYSTEM TEST PHASE II

PLAN

for the

Radiosonde Replacement System (RRS)

April 2005

U.S. DEPARTMENT OF COMMERCE

National Oceanic and Atmospheric Administration

National Weather Service/Office of Operational Systems

Field Systems Operations Center/Observing Systems Branch

(This Page Left Blank)

Table of Contents	Page

Acronyms	iv

1	Introduction	1

1.1	ST Phases	1

1.2	Test Plan Organization	1

1.3	RRS Subsystems	2

1.4	Purpose of Test	2

1.5	System Test Phase II Strategy	3

1.6	Corollary Documents	4

2	Test Methodology	4

2.1	Prerequisites, Assumptions, and Risks	4

2.2	Test Plan	5

2.2.1	Test Location	5

2.2.2	Pre-Test Activities	6

2.2.3	Documentation	6

2.2.3.1	Required Documentation for Test				          6

2.2.3.2 Forms								          8

2.2.4	Evaluation Areas	 8

2.2.4.1	Workstation Training 						          9

2.2.4.2	Supplemental Operator Familiarization			          9

2.2.4.3	Upper Air Operations						          9

2.2.4.4	Flight Characteristics						        10

2.2.4.5	Security and Software Updates				        11

2.3	Analyses	11

2.3.1	NCEP Analyses	11

2.3.2	NCDC Analyses	12

2.3.3	Flight Performance	12

2.3.4	RRS/MicroART Meteorological Comparisons	14

2.3.5	GPS-IPW Comparisons	15

2.3.6	Reliability, Maintainability, and Availability (RMA) Data Analyses	15

2.4	Communications Tests	15

2.4.1	Back-up Communications	16

2.4.2	Communications Headers	16

2.4.3	NWS Customer Involvement	16

2.5	Post-ST II Recommendations, and Report	17

2.5.1	Evaluation of Objectives	17

2.5.2	Test Result Analysis	18

2.5.3	Test Recommendations and Report	19

3	Test Schedules	20

3.1	Test Objectives by Week	20

3.2	OFS Test Objectives	20

3.2.1	Week 1 Test Objectives	20

3.2.2	Week 2 Objectives	21

3.2.3	Week 3 Objectives	22

3.2.4	Week 4 Test Objectives	22

3.3	Test Milestones	23

4	Test Roles and Responsibilities	26

4.1	Test Review Group (TRG)	26

4.2	Test Membership	26

Tables

Table 1-1.  Major Activities	4

Table 2-1.  RRS Test and Operations Documentation	7

Table 2-2.  RRS Maintenance Documentation	7

Table 2-3.  Forms	8

Table 2-4.  Objectives and Evaluation Methods	17

Table 3-1.  Activities Prior to Start of OFS	23

Table 3-2.  Activities Prior to Start of OFS	23

Table 3-3.  Activities Just Prior to Start of OFS	24

Table 3-4.  Activities During OFS	24

Table 3-5.  Activities After OFS	25

Table 4-1.  Test Personnel	26

Figures

Figure 1.  WFO-KCAR.	5

Figure 2.  Example of RRS comparison with first-guess.	12

Figure 3.  Example of IPW plot for GPS, MicroART, and RRS.	15

Attachments

Appendix A - Legacy System Performance Measures	A-1

Appendix B - Operational Flight Series Test Schedule	B-1

Acronyms

| As | System Availability |
| --- | --- |
| AWIPS | Advanced Weather Interactive Processing System |
| BILS | Balloon Inflation and Launch Shelter |
| BIT | Built-In-Test |
| BWWFO | Baltimore, MD/Washington, DC Weather Forecast Office |
| CCB | Configuration Control Board |
| CLIMAT | NWS Monthly Site Upper Air Summary Product |
| COTR | Contracting Officer(s Technical Representative |
| COTS | Commercial Off-The-Shelf |
| CDU | Control Display Unit |
| ERH | NWS Eastern Region Headquarters |
| EMRS | Engineering Management Reporting System |
| ESA | Electronics System Administrator |
| FTP | File Transfer Protocol |
| GPS | Global Positioning System |
| GUI | Graphical User Interface |
| hPa | Hecto-Pascal |
| km | Kilometer |
| kts | Knots |
| LDAD | Local Data Acquisition Device |
| LRIP | Low-Rate Initial Production |
| LRU | Line Replaceable Unit |
| LWX | National Weather Service identifier for the BWWFO. |
| MIC | Meteorologist-In-Charge |
| NCDC | National Climatic Data Center |
| NCEP | National Centers for Environmental Prediction |
| NCF | (AWIPS) National Control Facility |
| NHDA | NWS Headquarters Development AWIPS |
| NLSC | National Logistics Support Center |
| NRC | National Reconditioning Center |
| NMTW | NWS Headquarters Modernization Test and Integration AWIPS, WFO System |
| NWS | National Weather Service |
| NWSO | National Weather Service Observing |
| NWSTC | NWS Training Center |
| NWSTG | NWS Telecommunication Gateway |
| OAT | Operational Acceptance Test |
| OBIT | Offline Built In Test utility suite |
| OMS | Offline Maintenance Suite |
| OPS11 | NWS Engineering and Acquisition Branch |
| OPS12 | NWS Maintenance Branch |
| OPS13 | NWS Configuration Branch |
| OPS22 | NWS Observing Systems Branch |
| OPS23 | NWS Software Branch |
| OPS24 | NWS Test and Evaluation Branch |
| OS7 | NWS Observing Services Division |
| OST11 | NWS Program Management Branch |
| OST31 | NWS Analysis Branch |
| PAMS | Product Availability Monitoring System |
| PCA | Physical Configuration Audit |
| PDB | Precision Digital Barometer |
| RH | Relative Humidity |
| RMA | Reliability, Maintainability, and Availability |
| RRS | Radiosonde Replacement System |
| RSIM | RRS System Integration Manager |
| RSOIS | Radiosonde Surface Observation Instrumentation System |
| RWS | RRS Workstation Subsystem |
| SIR | System Issue Report |
| SIT | System Integration Test |
| SRS | System Requirement Specification |
| SPS | Signal Processing System |
| SR&DC | Sterling Research and Development Center |
| ST | System Test |
| S/W | Software |
| TRG | Test Review Group |
| TRS | Telemetry Receiving System |
| UPS | Uninterruptible Power Supply |
| UTC | Universal Time Coordinated |
| WFO | Weather Forecast Office |
| WMO | World Meteorological Organization |
| WSH | National Weather Service Headquarters |
| WSOM | Weather Service Operations Manual |
| XDP | External Data Pump |

# Introduction

The National Weather Service (NWS) will transition from the current radiosonde system into the new Radiosonde Replacement System (RRS) to collect and process upper air data using modern technology.  The RRS will be nationally deployed after the government successfully completes a series of tests, the last two of which are the System Test (ST) and the Operational Acceptance Test (OAT).

This plan describes the government tests performed during ST Phase II.  The ST Phase II will verify the RRS:

Meets NWS upper air operational usability for use at field sites.  Field staff will conduct all functions associated with the operations phase.

Meets field performance metrics and Concept of Operation performance standards.

Can transmit upper air products through various telecommunications networks to NWS internal users and external customers.

The ST II will be performed at the National Weather Forecast Office  located in Caribou, Maine. Assistance with data analysis will be provided as requested by the National Centers for Environmental Prediction (NCEP) located in Camp Springs, Maryland; and the National Climatic Data Center (NCDC) located in Asheville, North Carolina.

## ST Phases

The ST will be conducted in two phases. The following sections discuss the requirements for each phase of testing. ST Phase I will verify the functional software requirements from the RWS Software Requirements Specification have been met to the satisfaction of NWS Management.  ST Phase II will simulate operational use of the RRS and will commence after successful completion of Phase I and prior to OAT. Thus, ST Phase II is the capstone test ensuring the RRS is ready for deployment to field sites and can perform in a similar manner to the legacy system it is replacing.

## Test Plan Organization

This plan is a subsidiary of the System Test Plan for the RRS, dated July 2003.  It is composed of four sections.  Section 1 contains introductory material on the RRS major components and presents the test strategy, test objectives, and test prerequisites for this phase of testing.  Section 2 describes the methodology for conducting the ST Phase II, the facilities employed, and pre- and post-test activities.  Section 3 describes the critical milestones during the test and Section 4 contains information on test personnel and their responsibilities.  The following appendices are included:

(	Appendix A - Legacy System Performance Measures

(	Appendix B - Operational Flight Series Test Schedule

## RRS Subsystems

The RRS consists of the following major subsystems:

Radiosonde Surface Observation Instrumentation System (RSOIS) and Precision Digital Barometer (PDB) ( collects pre-release surface data. RSOIS will be tested with both the radio and the optical cable connected.  There will also be a portion of the test where the RSOIS will not be interfaced to the RWS, thus, requiring manual entry of the surface data.

Telemetry Receiving System (TRS) and associated Uninterruptible Power Supply (UPS) ( tracks the Global Positioning System (GPS) radiosonde and receives radiosonde telemetry signal; the UPS allows radiosonde tracking and data collection to continue in the event of local power failure.

Signal Processing System (SPS) ( translates the telemetry signal into displayable upper air parameters, both received and processed.  The SPS maintenance port is activated for use with the OBIT software for advanced diagnostics.

RWS ( allows RRS operators to receive and process data from the TRS, SPS, RSOIS, and PDB.  The RWS also controls the TRS, displays received and processed data, and disseminates formatted products to the Advanced Weather Interactive Processing System (AWIPS).  The primary software under test will be RWS Build 1.0.4.0.  After the ST, the approved build delivered to the OAT sites will be designated as 1.1.

GPS radiosondes - allow upper air wind direction and speed to be determined from differential triangulation data transmitted by the 24-satellite GPS constellation.  (Note, the initial ST Phase II will only use Sippican GPS radiosondes.)

## Purpose of Test

The purpose of the ST Phase II is to verify the RRS is fully functional in an operational sense, reliable when compared to the legacy system, products can be transmitted reliably to NWS internal users and external customers.  This phase of  ST must meet the following objectives under operational field conditions:

Validate RRS operational usability during a quasi-operational environment.

Exercise all RRS subsystem interfaces to other NWS systems [e.g., the Advanced Weather Interactive Processing System (AWIPS)/Local Data Acquisition Device (LDAD)] and back-up telecommunications.

Collect performance data for evaluating meteorological algorithms including the solar radiation correction and generation of coded messages.

Collect data for use during the Reliability, Maintainability, and Availability (RMA) analysis using the Engineering Management Reporting System (EMRS).

Validate simulated end-to-end functionality to NCEP and NCDC and other NWS customers for their product verification of decoders and data bases.

Assess the operational usability of RRS-generated data and products by a limited set of operational forecast applications and through comparisons with MicroART using RAOB, a commercial upper air analysis product.

Evaluate the installation of security and software updates without impacting operations.

Assess operator/maintenance training and supporting documentation.

Limited assessment of the logistic process.

Refer to Section 2.2.4 for a discussion of the areas to be evaluated:

## System Test Phase II Strategy

The basic strategy for conducting ST II is to operate the RRS in a (quasi-operational mode and evaluate its performance in various areas as shown in Table 1-1. In the first column, pre-test activities are listed, which will be completed before the actual test begins.

This will be followed by actual use of RRS in operations as conducted at a real upper air site using field staff to conduct the operations (refer to the second column in Table 1-1).  Core aspects such as flying radiosondes coincident with MicroART flights at synoptic and asynoptic times will be performed and the results tabulated exactly like a field site.  To not cause problems for the field site, a unique test header (KCRT, station identifier 69992) has been established on all messages transmitted over AWIPS.  In this way, there shouldn(t be any confusion with official products issued.

From this data, performance metrics will be generated to determine if the RRS is functioning at sufficient levels to warrant fielding the system.  Comparisons will be made with MicroART, the Forecast Systems Laboratory(s GPS-Integrated Precipitable Water (IPW) sensor, and assessed by both NCEP for their numerical weather models, and NCDC for its quality control procedures.

In addition to the above, a detailed end-to-end telecommunications test will be conducted, coordinated with both NWS and non-NWS users.  The purpose of this portion of the test will be to validate product throughput from KCAR via AWIPS/NOAAPORT to the NWS Telecommunications Gateway (NWSTG) for distribution to users.  There are a small number of product differences between the legacy products and the RRS products which need to be tested with customer decoders and data bases.

Two Key Decision Points (KDP-1 and KDP-2) will occur during the test to determine the state of the testing.  The KDPs are specified times during the test where the Test Review Group (TRG) (see Section 4.1 of this plan and the ST Plan for further details about the TRG) will review the test results and make any necessary decisions about the state of the test. At the completion of the test, preliminary results and a final report will be developed to ascertain the state of the system and its suitability for use in an operational environment.

## Corollary Documents

The reader of this plan is expected to be familiar with the System Test Plan for the RRS, dated July 2003.  In addition, this document has been prepared for those with an understanding of upper air systems and the Radiosonde Replacement Program.  A review of the Radiosonde Replacement System (RRS) Deployment Plan, June 2003, V9.4 is considered critical for understanding the major steps in the deployment of this system.  Documents required in support of the ST Phase II are found in Section 2.2.3.

| Table 1-1. Major Activities |
| --- |
| Pre-Test Activities | Operational Testing | Post-Test Activities |
| Ground system certified | Conduct synoptic/asynoptic flights | Analysis of test results |
| Radiosondes in stock at NLSC | Products sent over NCEP and NWS users | Final Test Review Group meeting conducted |
| AWIPS configured | NCDC receives archive files | Complete test reports |
| Documentation available | Security updates completed |  |
| Training completed | New software version updated |  |
| Software installed & switched to operations mode | Data analyses performed by NCEP & NCDC |  |
| Supplemental Observer Familiarization completed | Flight performance statistics generated |  |
|  | Meteorological comparisons with legacy system conducted Reliability, Maintainability, Availability data analysis performed |  |
|  | Training survey completed |  |

# Test Methodology

The following sections describe the test methodology, which will be employed during ST Phase II.  It is intended to provide a logical progression of the major activities and responsibilities for each aspect of the test.

## Prerequisites, Assumptions, and Risks

The prerequisites, assumptions, and risks associated with this phase of testing can be found in the System Test Plan for the RRS, dated July 2003.

## Test Plan

The sections below delineate the general plan for the conduct of the ST Phase II test.

### Test Location

The ST Phase II will be conducted in Caribou, Maine (referred to as KCAR, herein).  This site makes for an ideal test facility, because the RRS is configured exactly the way other field upper air sites will be done and it affords a high degree of operational assessment without interfering with the commissioned system, MicroART.  Figure 1shows a picture of the WFO environment.

Factors which were considered  in selecting Caribou as the ST II test site include:

- Office has a separate inflation building housing a TRS under its radome and installed just like field sites will be during the deployment phase.

- Office has historically been the NWS upper air winter weather test site.

- Comparisons with the legacy system, MicroART, are possible.

- As a predetermined winter test site, many of the training, installation and checkout aspects along with field operations have been evaluated in preparation for ST Phase II.

- Site has been pre-configured and tested to be ready for RRS comminations tests.

### Pre-Test Activities

The following pre-test activities have been or must be completed before the test begins.

| ( Ground Equipment | - TRS already installed at KCAR |
| --- | --- |
|  | - SPS/Base GPS installed at KCAR |
|  | - RSOIS with optical cable & base station already installed at KCAR |
|  | - PDB already installed at KCAR |
|  | - GPS Repeater already installed |
|  | - RWS XP workstation configured and connected to the office LAN |
|  | - Ghosting CD, RWS and OMS s/w should be delivered for the test with RWS Version Description Document (VDD) and OMS VDD |
|  | - Tool kit |
| ( System Certifications | - Engineering and software staff complete certifications at KCAR, like was done for ST Phase 1A ( checklists |
| ( Radiosondes | - 3 boxes delivered from NLSC from LRIP |
|  | - Power supply (part of tool kit), if available |
| ( AWIPS | - LDAD port configured for RRS per EHB-13, AWIPS RWS-LDAD System Administration Note 15 |
|  | - New comms line interfaced to LDAD port |
|  | - ISP interfaced for security updates |
|  | - AWIPS set up to ingest and display RRS products with test headers |
| ( Documentation | - Operator training materials shipped to field staff |
|  | - Maintenance materials available |
|  | - System administration-related procedures available |
|  | - Surveys for training completed |
| ( Training | - 1 ET/RMS Trained via NWSTC |
|  | - 3-6 HMTs complete training via distance learning and proficiency test |

### Documentation

Two sets of documents are required for ST Phase II: required documentation for conducting the test and a series of forms as part of the test paperwork.  These can be found in the subsections below.

      - Required Documentation for Test

RRS support documentation such as hardware/software installation and operational instructions, will be used in the ST Phase II.  Reference to these documents will be made as required throughout the test.  The list of test and operational documentation and procedures includes, but is not limited to, those items contained in Table 2-1.  Table 2-2 contains the maintenance-related and system administration documentation.

Table 2-1.  RRS Test and Operations Documentation

System Test Phase II Procedures for the RRS (final)

RWS Build 1.1 VDD and OMS VDD

The RRS home page: http://www.ua.nws.noaa.gov/RRS.htm

RRS Operator Training Guide (final draft)

Weather Service Operations Manual (WSOM) 30-2109 (EHB-9) Update (final draft)

WSOM 30-2101 (EHB-1) Update (final draft)

WSOM 30-2104 (EHB-4) Update (final draft)

National Weather Service Observing (NWSO) Handbook 10 Update (final draft)

Table 2-2.  RRS Maintenance Documentation

| Description | Manual Number | Current Draft Date | Status |
| --- | --- | --- | --- |
| RSOIS O & M Manual, Rev B | NWS EHB 9-201 | 13 April 2004 | Revision B submitted on 3/22/04 for Final Publication Process. |
| TRS O & M Manual, w/Attachments A thru E | NWS EHB 9-753 | 06 May 2004 | Submitted on 3/30/04 for Final Publication Process |
| SIPPICAN SPS & Radiosonde O & M Manual & SPS Software Operations | NWS EHB 9-801 | 18 March 2004 | Submitted on 3/30/04 for Final Publication Process. |
| Radiosonde Workstation (RWS) O & M Manual, w/Attachments A & B | NWS EHB 9-901 | 05 May 2004 | Document distributed on 13 April 2004 for technical review. |
| SIPPICAN SPS Users Manual for SPS Software used on the RWS | NWS EHB 9-902 | 18 March 2004 | Submitted on 3/30/04 for Final Publication Process. |
| RRS Fault Analysis Manual | NWS EHB 9-903 | 05 May 2004 | Fault Analysis is available in individual O&M manuals. (This manual is under development). |
| RRS Systems Administration Manual | NWS EHB 9-904 | 05 May 2004 | Document distributed on 13 April 2004 for technical review. |

      - Forms

During the test, a series of forms including surveys, operations-related, EMRS, and checklists will be completed during and after the test.  The complete list can be found in Table 2-3.

Table 2-3.  Forms

| Form Type | Form Name | Purpose | Responsible organization |
| --- | --- | --- | --- |
| Operational Form | B-85 | Enter logistics information | OPS22/OS7 |
| Operational Form | B-29 | Flight Performance | OPS22/OS7 |
| Operational Form | B-47 | NCDC Report | OPS22/OS7 |
| Operational Form | H-6 | Reject radiosondes | OPS22/OS7 |
| Survey | Proficiency Test | Evaluate the RWS training and proficiency | OS7/OPS24 |
| Maintenance Form | EMRS | Entering maintenance-related information | OPS21 |

### Evaluation Areas

The following major areas will be evaluated during the ST Phase II:

There will be an assessment of the training process and materials. This will include an assessment of the proficiency exam and the activities associated with familiarizing the operations staff with the RRS.  Note: field staff from non-KCAR locations will be involved in this test.

Verifying that RRS can operate in a 24/7 mode with system/flight performance similar to or better than MicroART.

Demonstrate, using available system administration documentation, to perform security updates, and the installation of a new software load without having an adverse impact on operations.

Although there will not be any specific evaluation of the maintenance aspects of the system, there may be a need for NLSC to deliver a replacement part to the site and the site ET/RMS staff replace the failed component.  If problems or issues exist with either the logistics support, maintenance documentation, or level of training during the test these will be relayed to the appropriate engineering staff.

The sections below discuss each of these in detail.

Workstation Training

Since the workstation, application software, training materials, and the proficiency test, will be shipped to the sites where the operators will be trained in advance, this will be the first area to be evaluated.  The following is a list of activities that will be completed by the operators participating in the test.

(1) Reviewing and commenting on the new RRS Handbook WSOH#10 and RRS Operator Training Guide

(2) Loading the RRS software (and sample flights on a CD for reworking) on a PC in your office for training purposes

(3) Reviewing a RRS training DVD

(4) Taking an open book examination on observations with RRS and pass the exam.

Supplemental Operator Familiarization

One of the key deployment elements is familiarizing the office staff with the various components of the RRS.  The approach is for the deployment team to provide hands on familiarization training to  site personnel using the RRS equipment after completion of the RRS INCO. This process is called the Supplemental Operator Familiarization (SOF).  For this to take place, it is assumed that site personnel will have completed the required training material and have passed the RRS certification exam prior to system deployment. It is also assumed that since the RRS workstation and software will be delivered to the sites well in advance of deployment, the field staff will already be familiar with the basic software functions.  The concept of the SOF is to value add the training material with actual hands on experience.  The SOF will consist of , but not limited to, operating the TRS, activating the radiosonde and performing the baseline, acquiring GPS signals, and launching the radiosondes.  Since the deployment team will only have a few days to conduct this activity, it is imperative to maximize the limited time in terms of providing this familiarization to the office staff in an efficient manner.  For the purposes of this test, the SOF will be provided to both the KCAR and non-KCAR staff to ensure a good assessment of this activity.  A survey will be provided to the operations staff to identify strengths and weaknesses with this approach.

Upper Air Operations

The next phase of testing will be a series of synoptic and non-synoptic (Special) flights conducted over a contiguous 30-day period. This is referred to as the Operational Flight Series (OFS).  The OFS  will consist of  approximately 85 upper air observations being taken for the purpose of developing performance characteristics for RRS. One other important feature of the OFS is that it will facilitate the RWS comparing current flight data with flights conducted 12-hrs and 24-hrs ago.  Unlike previous tests where a system/radiosonde failure caused a System Issue Report (SIR) to be written; in this test, this type of failure will result in a second or third release.

Since this is a performance-based test, the number of second and third releases for RRS will be compared with the average number of similar-type failures encountered with MicroART.  These are delineated in Appendix A as the Legacy System Performance Measures (LSPM) Tables.

Flight Characteristics

Flights will be conducted exactly as if it were the actual KCAR field site, with one-person releases, messages transmitted via AWIPS (including back-ups), and the archive data sent to NCDC after each flight.  Additional tests such as transmitting the archive data, say after 2 flights, or multiple releases, will also be performed. NCEP will receive all these products for assessing any issues with the data and will also generate their standard performance measures.  The Test Director will coordinate with NCEP and NCDC on the particulars of each test.

The ascension number assigned to the MicroART flight and the associated first (official( synoptic RRS flight will be documented as part of the record.  Note, the ascension numbers do not have to match throughout the test, just have a record of both. For the first three weeks of the test, synoptic flights will be conducted without inducing any special situations during the synoptic flights.  In this way, actual performance statistics can be acquired.  However, different configurations such as changing the RSOIS connection from optical cable to radio will be conducted during this period.

During asynoptic times a host of special situations can be performed without interfering with the synoptic flight performance.  On several occasions during asynoptic flights, dummy failures will be induced to verify that, for example, second and third release conditions can be met without problems. When the first flight fails, then a second release will be taken.  If that one also fails, a third one will also be taken.  Two flights might have the first flight fail due to an induced radiosonde failure, e.g., cut temperature sensor, and then a second flight will be attempted as normal.  Note, any induced failures will not be counted against the RRS performance statistics (Appendix A) and additional flights will have the 1st & 2nd flights fail and the third flight will be normal.  If second and third flights are required during synoptic flight conditions, then these could replace the (dummy( situations.

Additional special flights will be conducted if the Test Director believes a special condition is warranted, say, a night-time flight. All special flights will be transmitted via AWIPS.  BUFR files will be sent to NCDC after every flight and operational forms as specified in Section 2.2.3.2 will be completed.

If any abnormalities (other than those induced for test purposes) or indications of noncompliant operations are observed during the ST Phase II, SIRs will be written to document the problem and called to the immediate attention of the Test Director.  The Test Director will present the discrepancies to the TRG for classification and recommendation.

Security and RWS Software Updates

At least once during the OFS, a download of one or more security updates must be attempted, using EHB9-904. The field ESA or ITO staff will perform this update using the System Administrator(s manual for guidance.  The update will be performed by downloading the update via an Internet Service Provider connection directly into the RRS workstation.  After the RWS has been updated and checked out, flight operations will continue with no adverse effects. A report on the effectiveness of this activity to support field operations will be included in the final report. At about half way through the test the RWS software will be updated using the appropriate manuals.

## Analyses

The following sections delineate the types of analyses required during the test.

### NCEP Analyses

NCEP personnel will import the RRS test products into their meteorological models and perform quality control of the data. They will also compare the RRS products with official products from WFO-KCAR.  This comparison will occur during the ST Phase II using the WFO-KCAR legacy upper air system and RRS launched balloons.  This will be accomplished by launching two balloons  within 15 minutes of one another. NCEP will add the ST II system WMO identifier (69992) to their (THKS( web page (http://www.ncep.noaa.gov/NCO/DMQAB/QAP/thanks/) and display the status of the data along with the data from the WFO-KCAR.

In addition, NCEP will determine if upper air data from the 69992 site is:

(	Being correctly ingested into models, analyses, and can determine estimated height/temperature biases

(	Identifying the freezing levels correctly

(	Not rejecting significant amounts of data from QC algorithms

An example of one of these charts is shown in Figure 2.

### NCDC Analyses

The RWS  archive flight data will be  submitted to the NCDC over the Internet by File Transfer Protocol (FTP) after every flight. OPS22 and NCDC staff will evaluate the files for format and accuracy of content.  SIRs will be generated if problems are identified. This analysis will be completed for products generated from live flights and products generated after performing a rework.

NCDC will ingest these data into their archival system and generate their normal quality-control reports.  OPS22 and other staff will compare the RRS results with the MicroART from the KCAR site to determine if any performance issues exist.

### Flight Performance

The objective of this portion of the test is to verify RRS flight characteristics for approximately 70 consecutive flights and its ability to produce flight performance similar to the legacy system. These must be in good agreement with the legacy system with none categorized as poor by the end of the assessment.

A baseline set of performance measures have been derived from the legacy systems and are referred to as the Legacy System Performance Measures (LSPM). Each is supported from data collected and processed from current parameters over the years.  These will then serve as a comparison set and form a part of the report.  The specific criteria are listed in Appendix A. Statistics for 2nd/3rd releases, unsuccessful releases, missed fights and rejected radiosondes, etc., will also be computed comparing MicroART with RRS.  To collect this information, upper air forms (Section 2.2.3.2, Table 2-3) will be filled out as is done at the field site using the Management Information Retrieval System (MIRS).  Statistics and graphics will also be generated by OPS22 for the RRS to compare with MicroART.

Possible outcomes: The following outcomes are possible for each LSPM when compared to the RRS for the duration of the test:

Excellent Performance  - the RRS performance measure far exceeds the minimum criteria for acceptable field use when compared to the LSPM

Good Performance - the RRS performance measure exceeds or meets the minimum criteria for acceptable field use when compared to the LSPM

Marginal Performance - the RRS performance measure just meets the minimum criteria for acceptable field use when compared to the LSPM

Poor Performance - the RRS performance does not meet the minimum criteria for acceptable field use when compared to the LSPM

By the end of the test period, all performance measures must be in the Marginal Performance or better categories, i.e., for this aspect of the test to be deemed successful, none can be in the poor category.

Test Conditions: After each flight, data related to each performance measure will be collected and documented.   The TRG will assess the RRS results against the LSPMs as shown in the Tables in Appendix A.  When problems occur, a determination will be made as to whether they will impact the operation of the system or a work-around can be established.  If a performance measure status is moving from marginal to poor, then various test conditions are possible:

Condition  1.  Continue testing and monitor.

Condition  2.  Pause and fix problem, then continue.

Condition  3.  Halt test and restart from beginning.

A review of the test status will be conducted each day.  Condition 3 will only be recommended  if any of the RRS performance measures (RRS PM) are categorized as poor (see below) and the prognosis also appears to be poor.  For example, if the software/hardware were very unstable and not able to process flight data in order to meet the LSPM, then it makes sense to assign Condition 3.  If there are bugs, glitches, etc. needing corrective action, then Condition 2 would probably be valid.  A condition considered more of a nuisance would probably call for a Condition 1 state.  The full operational assessment should not be conducted until there is confidence Condition 3 situations will be of low probability, as a result of ST Phase I testing.

### RRS/MicroART Meteorological Comparisons

Because of the impact RRS will have on field operations, it is necessary to produce comparison charts of upper air meteorological profiles with MicroART.  Perfect agreement is not critical here, rather the temperature, pressure/height, relative humidity, and wind plots must provide a certain degree of agreement, since both are sensing the same atmosphere. The solar radiation correction will be activated for the RRS, and therefore, the temperature profiles are expected to be cooler than the MicroART ones. Plots as produced by the RAOB program will be collected for surface-to-700 hPa,  surface-to-400 hPa, surface-to-100 hPa, and surface-to-flight termination.  Pressure/height curves for both systems will be acquired to compare profiles and calculated parameters from the compared flights will also be evaluated to determine if any problems exist.  Areas, which will be specifically evaluated using the RAOB program, are as follows:

(	Comparison of surface values between RRS and MicroART for the early part of the sounding.

(	Ekman spiral in the first two kilometers of wind data.

(	Low-level and mid-to-high tropospheric jets along with max winds and shears.

(	Cloud formation including areas of high and low RH, and integrated precipitable water.

(	Identification of freezing level crossings.

(	Thickness values.

(	Selection of Tropopause(s).

(	Stratospheric temperature profiles with RRS solar radiation corrections applied.

OPS22 with support from the KCAR forecast office and OS7 staff will review the comparison plots.  If they find any inherent problems, they will inform the Test Director for resolution of the problem. To simplify the meteorological analysis and to save on resources, some of the asynoptic flights may be conducted as a dual with one RRS and one MicroART radiosonde on the same balloon.

### GPS-IPW Comparisons

Comparisons of GPS-IPW values produced from FSL(s differential GPS system will be used to determine if there are biases with the total precipitable water values from the RRS observation as obtained from the RAOB program.

### Reliability, Maintainability, and Availability (RMA) Data Analyses

RMA information will be collected during the ST Phase II using the EMRS (see Attachment C of the RRS ST plan).  The collected data will be used to assess the potential reliability, maintainability, and availability only during this phase. The EMRS data will be provided to the appropriate staff for assessing the system reliability.  This reliability value will not be a prediction of future RRS reliability, but only an initial estimation.

## Communications Tests

The WFO-KCAR AWIPS will be used to receive coded messages from the RWS/RRS and to send upper air products to the NCEP and other NWS customers via the NWSTG using the current communications path. In general, products are issued from the local site through AWIPS to the Network Control Facility for satellite broadcast over NOAAPORT and to the NWSTG  for world-wide distribution.

### Back-up Communications

At least 2 synoptic flights will transmit all of its products over the 1st backup, which is the ERH LDAD located in Long Island, NY, instead of the primary.  At least two additional flights will use the 2nd backup at SRH in Dallas/Fort Worth in a similar manner.  The Product Availability Monitoring System (PAMS) will be used to collect the frequency/timeliness of upper air products transmitted from the WFO-KCAR AWIPS to the NCF and then back to the WFO.  NCEP may also provide information on the receipt times to their NCEP computer, especially if long delays exist.  Performance measurement requirements for timeliness and frequency can be found in Appendix A.

### Communications Headers

Products issued over the telecommunications networks will require message headers to be appended to the products.  This facilitates the transmission of these products over AWIPS and through the back-up communications line using the following headers:

| AWIPS PIL | Individual Test WMO Header | Individual WMO Header | Collective |
| --- | --- | --- | --- |
| CRTMANCRT | USUS97 KCRT | USUSii KCRT | USUS05 KWBC |
| CRTSGLCRT | UMUS97 KCRT | UMUSii KCRT | UKUS50 KWBC and UGUS50 KWBC |
| CRTABVCRT | UFUS97 KCRT | UFUSii KCRT | ULUS05 KWBC, UEUS05 KWBC, ULUS50 KWBC and UEUS50 KWBC |
| CRTFZLCRT | UXUS97 KCRT | UXUSii KCRT | UXUS70 KWBC |
| CRTULGCRT | NXUS97 KCRT | NXUSii KCRT | (Note, ULG product not required for this test) |

BUFR Collectives: IUST41 KWBC, IUST42 KWBC and IUST46 KWBC

The NWSTG will verify these products are being routed correctly to NWS customers and also verify they are being sent to NCEP.  Having the ii-portion of the header denoted as 97 prevents the products from being official, even though the contents will be actual data collected from the RRS.

### NWS Customer Involvement

Although there are only a few changes with respect to the actual coded messages, they can become a problem if these changes are not handled correctly by decoders and data bases.  NWS will be issuing a series of Public Notification Statements (PNS) to NWS customers indicating the availability of test messages and where to find information about the coded message changes.   The necessary product header information as described in Section 2.4.2 will be included.  NWS customers will be invited to participate in the test through the PNS as either passive or active participants.  Passive participants can use the information provided to check out their decoders and other processing functions and do not need to inform the NWS of their results unless they need some assistance.  Active participants are those which the NWS would like to hear from,  to ensure at least some of the participants exercised their processing systems during the test and verified no adverse effects.  Some potential NWS customers who could become active participants include:

(	FAA, AFWA, and FNMOC ( government agencies

(	Unidata ( universities

(	WSI ( commercial

The actual active participants will be selected after coordination within NWS is completed on potential candidates and a request is made to them.

## Post-ST II Recommendations, and Report

The following sections describe how the data collected during the ST II will be analyzed and how a  recommendation to proceed is determined.

### Evaluation of Objectives

The ST Phase II Objectives identified in Section 1.4 will be evaluated by the methods specified in Table 2-4.

Table 2-4.  Objectives and Evaluation Methods

| Objective | Evaluation Method |
| --- | --- |
| Validate RRS operational usability and maintainability. | Field staff will exercise RRS in (quasi-operational mode and verify system is ready for field use. Maintenance staff can maintain the system in accordance with established documentation. Problems needing corrective action before field deployment will be captured in SIRs. |
| Exercise all RRS subsystem interfaces to other NWS systems [e.g., the AWIPS/Local Data Acquisition Device (LDAD)]. | Successful completion of flight and transmission of RRS observations, followed by transmission of archive data. |
| Collect performance data for evaluating meteorological algorithms including solar radiation correction and generation of coded messages. | Analysis by OPS22 and OS7 subject matter experts verifies RRS is operating nominally. |
| Collect data for use during the Reliability, Maintainability, and Availability (RMA) analysis using the Engineering Management Reporting System. | OPS11/OPS12/OPS13 to analyze data collected using a modified Engineering Management Reporting System A-26 form. |
| Validate end-to-end functionality at an operational site. This will include RRS product format and data evaluation. | NWSTG will evaluate the products by generating their monthly CLIMAT report and checking throughput to NWS customers. NCEP will import RRS products into their meteorological models and perform quality control of the data. NCDC will evaluate archive generation and data quality for their products. NWS Customers can successfully decode and integrate the products. |
| Assess the operational usability of RRS-generated data and products by a limited set of operational forecast applications. | OPS24 and OPS22 will examine RRS products distributed by the AWIPS NCF and their use in AWIPS UA applications. |
| Assess Maintenance Actions/Fault Isolation using the OMS and the Offline Built-In-Test (OBIT) applications. | Evaluated by the field electronics staff with support from OPS11/OPS12 as faults arise. |
| Verify the Radiosonde Replacement System Project Agreement document minimum RRS performance requirements. | Field staff will collect data for analysis as follows: 1. Percentage of all naturally terminating test flights processing data above the 400 hecto-Pascal (hPa) level is equal to or greater than 95%. 2. Percentage of all naturally terminating test flights processing data above the 10 hPa level is equal to or greater than 60%. 3. Percentage of all test flights with 90% of wind and temperature values passing NCEP quality control tests will be 97% or greater. 4. Percentage of all test flights transmitting messages within three hours of balloon release is 95% or greater. |
| Verify the Legacy System Performance Measures (LSPM) are being met. | Data collected by the field staff and OPS 22 and analyzed by OPS22/OS7. |

### Test Result Analysis

During the real time analysis of the test results, the criteria presented below  will be monitored and presented by the test team during the weekly TRG. If at any time during the test this criteria is not met, the test director will notify the TRG for further analysis of the problem and to determine if the problem is serious enough to suspend testing.  The TRG may suspend the test, if at any time a problem prohibits the following criteria from being met.

Successful completion of the test procedure for each test case will be used as one of the pass/fail criteria.  If any abnormalities are observed during the test procedures, SIRs will be written to document the problems and will be provided to the TRG for adjudication.  The TRG will prioritize all SIRs and determine if they warrant being submitted to the RRS CCB.  The TRG will not submit a recommendation to the RRS Program Manager to proceed if there are any unresolved Impact 1 or 2 SIRs.

The second result analysis will be an examination of RRS upper air observation products output for content correctness as well as initial operational performance characteristics.  This analysis will be performed by OPS22 and OS7 subject matter experts coordinating data analysis with NCEP and NCDC.

The Legacy System Performance Measures are deemed minimally acceptable. By the end of the test period, all performance measures must be in the Marginal Performance or better categories. (i.e., for this aspect of the test to be deemed successful, none should be in the poor category.)

The Radiosonde Replacement System Project Agreement document states the minimum RRS performance requirements in terms of operational use. The following minimum performance requirements are extrapolated from this document and tailored to the ST:

Percentage of all naturally terminating test flights processing data above the 400 hecto-Pascal (hPa) level is equal to or greater than 95%.

Percentage of all naturally terminating test flights processing data above the 10 hPa level is equal to or greater than 60%.

Percentage of all test flights with 90% of wind and temperature values passing NCEP quality control tests will be 97% or greater.

Percentage of all test flights transmitting messages within three hours of balloon release is 95% or greater.

The statistics for the above requirements will be calculated for each system independently and as a group for all systems used during the ST-II.  The statistics will be adjusted to remove, where and when necessary, inappropriate data such as defective balloons or an operator induced failure to test a specific requirement.

### Test Recommendations and Report

Following completion of the ST Phase II, the Test Director will conduct a (Wrap-up( meeting with the TRG describing what was tested, a summary of any discrepancies found, major findings, and recommendations.  The TRG will review the materials presented by the Test Director and make a recommendation to the RRS Program Manager whether to proceed with OAT.  After the (Wrap-up( meeting, the Test Director will prepare a formal report of all test activities, including details of any deficiencies.  The report will also include findings and recommendations.

# Test Schedules

The following sections contain the schedules for various aspects of the test.

## Test Objectives by Week

Specific test objectives will be met each week of the test.  Even though testing during ST Phase II will be a 24/7 operation to mimic actual field operations, a test week will be defined as starting with the  Friday 00 UTC sounding, and concluding on Thursday of the following week at 12 UTC.  In this way, a number of test-related activities can be concluded throughout the week leading up to the start of the new week.

## OFS Test Objectives

In order to meet the evaluation areas of Section 2.2.4 above, the types of  following test objectives will be adhered to during each week of the test.  However, the Test Director has the flexibility to modify this plan as conditions warrant and terminate the test at any time if there is a determination that continuance of the test would be counter-productive.  Documentation found in Section 2.2.3 will be used to support the operational use, system administration, and maintenance of the system.

### Week 1 Test Objectives

The following test objectives will be conducted, and the results evaluated, during this week of testing:

Synoptic Flight Objectives

(	Perform 00 &12 UTC flights exactly like would occur at a field site:

- 1-person releases and different configurations, e.g., each flight will have a different frequency.

- All 00 UTC flights use primary AWIPS & all 12 UTC flights will use 1st back-up dial to ERH AWIPS.  A select set will also be transmitted to the second back-up at SRH.

- If sonde/system failure requires 2nd/3rd release, then appropriate action will take place.

- Some flights will have launches move towards the end of the synoptic window.

- Comparisons with MicroART using RAOB software will be conducted and results transmitted to OPS22/OS7 for review.

(	Updated security profile from ISP download, per regional procedures.

(	Primary objective is to collect flight performance data to compare against Legacy System Performance Measures tables (Appendix A) & ensure number of rejected radiosondes is within specifications, based on H-6 Form.

- Evaluate RRS product throughput using PAMS.

- Determine if RRS is stable in terms of day-in/day-out operational/system availability.

- Flight/System performance must be at least marginal, while goal is good-to-excellent range.

- NCEP /NCDC reports/graphics are deemed positive and NCDC verifies BUFR messages correspond with RWS files produced.

- All flights will reside in RWS to verify system performance.

- EMRS updated as required.

(	Forms will be filled out just like regular site and transmitted to OPS22 for assessment, when requested.

(	Local applications will be checked out, if possible, e.g., AWIPS Skew-T, Log P diagram.

(	RSOIS optical cable used to receive surface data.

Asynoptic Flight Objectives

(	Test special cases not found during (normal( flights.

- TRS off-target launches, i.e., 180 degrees off target at launch.

- Special meteorological conditions cases. (e.g.,  launch near thunderstorm)

- multiple releases situations. (i.e., 2nd/3rd release situations)

### Week 2 Objectives

The following test objectives will be conducted, and the results evaluated during this week of testing:

Synoptic Flight Objectives

(	Same as week 1, continue compiling performance results.

- Skip transmission of archive data to NCDC for one flight and then transmit both synoptic BUFR messages after next synoptic flight. (Note, NCDC must inform site that BUFR message was not received before next synoptic flight.)

- RSOIS spread spectrum radio used instead of optical cable.

- KDP-1 on results to-date.

- All flights will reside in RWS to verify system performance.

Asynoptic Flight Objectives

(	Test Special Observation cases similar to week 1.

### Week 3 Objectives

The following test objectives will be conducted, and the results evaluated, during this week of testing:

Synoptic Flight Objectives

(	Same as week 1, continue compiling performance results.

- Skip entire flight to demonstrate ascension numbers are correct after no release case and coded messages transmitted (without data( from RWS to NCEP.

- RSOIS not used, instead surface values entered via ASOS.

- KDP-2 on results to-date.

- All flights will reside in RWS to verify system performance.

Asynoptic Flight Objectives

(	Same as week 1, except:

- Replace RWS with a replacement computer, then verify no impact on operations.

### Week 4 Test Objectives

The following test objectives will be conducted, and the results evaluated, during this week of testing:

Synoptic Flight Objectives

(	Same as week-1, continue compiling performance results.

Asynoptic Flight Objectives

(	Any remaining situations not tested.

- Install new software & use backup drive to restore data using admin.

- Download security from Internet.

- Retransmit messages from rework.

- Test 2nd/3rd releases if none leave occurred to-date.

- All flights will reside in RWS to verify system performance.

- Test Power failure cases.

## Test Milestones

The following tables delineate the major time steps leading to commencement of the OFS. For the purposes of this test, T=0 is when the entire RRS is connected within the office environment and ready for certification.  This means T=0 may require several days to accomplish, and so it is not considered a single day, but rather a period of time leading to the certification of the RRS.

Table 3-1 delineates those activities required 45 days or more prior to beginning the test (T > - 45).  Table 3-2 lists those actions required for completion within 45 days of the start of the test and Table 3-3 defines those activities that must be completed successfully within a week around T=0.  At this point, Table 3-4 activities support the evaluations as delineated under Section 2.2.4 and subsections.  Test reports and final coordination activities are listed in Table 3-5.

Table 3-1.  Activities Prior to Start of OFS

| Timeframe | Areas Tested | Responsible Org. |
| --- | --- | --- |
| T > - 45 days | Radiosondes shipped from NLSC to KCAR (3 boxes) | KCAR via CLS |
|  | Operations documentation/ software/1-radiosonde shipped to KCAR & volunteers | OS7 |
|  | At least 1-ET/RMS trained @ NWSTC | ERH |
|  | Master station file for KCAR established | OPS22/Region |

Table 3-2.  Activities Prior to Start of OFS

| Timeframe | Areas Tested | Responsible Org. |
| --- | --- | --- |
| 0 < T < (-45days) | Field/Test operations staff train on RWS | Test staff |
|  | GPS repeater calibrated @ KCAR | RMS @ KCAR |
|  | System certifications completed (checklists) | OPS11/OPS12/OPS23 |
|  | Remaining documentation available to KCAR via web | OPS12 |
|  | FTP to NCDC link tested | OPS22/KCAR staff |
|  | Field Test staff complete proficiency testing and survey | KCAR staff & field Volunteers/ OS7 |

Table 3-3.  Activities Just Prior to Start of OFS

| Timeframe | Areas Tested | Responsible Org. |
| --- | --- | --- |
| T = 0 | TRS connected to RWS & RSOIS optical cable, then radio, then neither, i.e., use ASOS | OPS22/OS&T DM |
|  | 2-5 test flights conducted | OPS22/KCAR staff |
|  | AWIPS/Internet/Backup comms lines connected and security patches downloaded | KCAR/ESA/ITO |
|  | AWIPS/LDAD setup for ingest and display RRS test products on D2D. | KCAR staff/OPS22/OPS24 |
|  | Test mode switch to Operations mode via re-install of s/w, i.e., data base wiped-out & ASC No. set to correct one. | OPS22/ESA/ITO |
|  | System certifications completed (checklists) | OPS11/OPS12/OPS23 |
|  | Test obs sent to NCEP and BUFR files to NCDC via FTP | KCAR staff/OPS22 |
|  | Begin Familiarization on RRS conducted with staff | OPS22/KCAR staff |

Table 3-4.  Activities During OFS

| Timeframe | Areas Tested | Responsible Org. |
| --- | --- | --- |
| 0 < T < 30 days | Begin operational flights: (2 synoptic/day and 1-2 Asynoptic/day | KCAR staff/OPS22 |
|  | Obs sent to NCEP & NCDC | KCAR staff |
|  | Complete familiarization on RRS conducted with staff |  |
|  | Collect performance statistics | OPS22 |
|  | Test special situations | KCAR STAFF/OPS22 |
|  | Key Decisions PT-1 for @ T = 15 days KDP-2 @T = 30 days | PM/TRG |
|  | Test local AWIPS applications | OPS22/KCAR staff |
|  | Checks AWIPS Applications | KCAR staff |
|  | Security patch Update | KCAR ESA/ITO |
|  | Install new s/w load | KCAR ESA/ITO |

Table 3-5.  Activities After OFS

| Timeframe | Areas Tested | Responsible Org. |
| --- | --- | --- |
| 45 < T < 60 days | Preliminary Report | OPS22 |
|  | NCEP Report sent to NWSH | NCEP |
|  | NCDC Report sent to NWSH | NCDC |
|  | Final Report Completed & issued | OPS22 |

# Test Roles and Responsibilities

The following sections delineate the test roles and responsibilities.

## Test Review Group (TRG)

The TRG is an assembly of subject-matter experts and is chaired by the Chief, Test and Evaluation Branch (OPS24).  The role of the TRG is to evaluate each observed deficiency (both repeatable and intermittent problems) as documented by a SIR, on its impact on daily field service operations and to make recommendations to the RRS Program Manager on SIR criticality.  Membership in the TRG can be found in the System Test Plan for the RRS, dated July 2003.

## Test Membership

There will be two test teams on site. Each team will consist of one or two volunteer field personnel and one test monitor from the OPS 22 staff in Sterling. Volunteer field personnel can be from any site, but in the selection,  priority will be given to personnel from the OAT sites. Volunteers  will be on site for approximately two weeks .  At the end of the first two week period, field personnel on site will depart and new field personnel will arrive. Personnel stationed in Caribou will be utilized as appropriate if conditions permit.

One team  works a day shift starting with the 12Z sounding, the other team will work evening/night covering the 00Z sounding. Each shift in most cases will also conduct one special or asynoptic sounding.  OPS22 will provide the specific scenario for each flight conducted by field personnel.

The test director will coordinate all activities with the OPS22 test monitor.

Table 4-1.  Test Personnel

| Name/Organization | Function | Phone | E-mail |
| --- | --- | --- | --- |
| James Fitzgibbon (OPS22) | RRS ST II Test Director | 703-661-1243 | james.fitzgibbon@noaa.gov |
| Jae Lee/ (OPS24) | RRS Test Advisor | 301-713-0326 x158 | jae.lee@noaa.gov |
| Bill Blackmore (OPS22) | OPS22 Analytical Support | 301-713-2091 x107 | william.blackmore@noaa.gov |
| Fred Branski (CIO11) | NWSTG Test Support | 301-713-0864 x146 | fred.branski@noaa.gov |
| William Ryman (OS612) | NWSTC Test Support | 816-880-9368 x242 | william.ryman@noaa.gov |
| Carl Bower (OS7) | Analytical Support (Data Processing and Algorithm Analysis) | 301-713-0722 x145 | carl.bower@noaa.gov |
| Bob Thomas (OS7) | Analytical Support | 301-713-0722 x127 | robert.thomas@noaa.gov |
| Ken Bashford | Survey Support | 301-713-0326 x113 | kenneth.bashford@noaa.gov |
| Field Personnel (TBD) | RRS Test Support @ KCAR | N/A | N/A |
| Jeff Ator (NP12) | NCEP Analytical Support | 301-763-8000 x7104 | jeff.ator@noaa.gov |
| Richard Robinson (CIO) | NWSTG Support | 301-713-0864 x179 | richard.robinson@noaa.gov |
| Dr. Bradley Ballish (NCEP) | NCEP Analytical Support | 301-763-8000 x7159 | bradley.ballish@noaa.gov |
| Wayne Martin (OPS21) | AWIPS Support | 301-713-1724 x166 | wayne.martin@noaa.gov |
| Eddie Roberts (OPS23) | RWS Support | 301-713-0191 x154 | edward.roberts@noaa.gov |
| Ivan Navarro (OPS11) | Radiosonde Support | 301-713-0844 x134 | ivan.navarro@noaa.gov |
| Darryl Modracek (OPS11) | TRS Support | 301-713-1842 x111 | darryl.modracek@noaa.gov |
| Kenneth Clark (OPS12) | Maintenance Support | 301-713-1833 x190 | kenneth.clark@noaa.gov |
| John Monte (OPS12) | RRS Support | 301-713-1845 x117 | john.monte@noaa.gov |
| Eric Parr (OPS14) | Logistics Support | 301-713-1830 x111 | eric.parr@noaa.gov |
| Stuart Hinson/Larry Griffin (CC11) | NCDC Support | 828-271-4437 | Stuart.Hinson@noaa.gov / larry.j.griffin@noaa.gov |

The following describes the major roles and responsibilities of the test personnel.

Test Director - Ensures all tests defined for the ST Phase II are completed and the results properly documented in the ST II report.  Responsible for collecting and presenting all test trouble reports to the TRG for classification and convening emergency TRG meeting, if necessary.  Following completion of the ST Phase II, the Test Director will conduct a (wrap-up( meeting of the TRG, provide details to the RRS Program Manager what was tested, report the ST II conclusions, and recommend whether to proceed with the next phase.  This person ensures all test trouble reports documented and classified during the ST II are forwarded to the proper WSH organization or board for adjudication; writes/issues the ST II preliminary and final reports to document the test results and recommendations.  The Test Director also serves as a member of the test team.

Test Team - Responsible for performing individual test procedures as assigned; documents the results of each test in test logs; electronically completes trouble report forms when problems/discrepancies are observed; provides the Test Director with all completed forms (if not entered electronically).  Informs the Test Director with comprehensive technical information on how the tests were conducted and any problems encountered.

Test Support - Responsible for providing technical support and information as required when RRS questions arise; forwards RRS hardware and software discrepancies to the respective contractor for resolution.

NWSTG Test Support -  Responsible for evaluating RRS product throughput and the ability of the NWSTG to generate monthly upper air CLIMAT products using RRS products.

NWSTC Test Support - Responsible for evaluating RRS maintainability; conducting RRS maintenance training; performing hardware evaluations as time permits.

NCEP Test Support - Responsible for evaluating RRS products for format and usability in meteorological models and performing quality control checks.

NCDC Test Support - Responsible for determining if archived RRS atmospheric data from radiosonde flights is in the proper format for use in producing NCDC climatic products.

Field Personnel Test Support - Responsible for evaluating system functionality and usability; participates in documentation reviews; prepares radiosonde during pre-flight and assists in both synoptic and special balloon launches.  Responsible for assisting with product quality control and collecting performance data.

Analytical Support - NCEP and NCDC personnel are responsible for evaluating synoptic products for accuracy, usability, timeliness, and the ability to be archived and retrieved.  OPS22 and OS7 personnel are responsible for evaluating all RRS algorithms and products for accuracy and maintaining ST flight logs.

Appendix A - Legacy System Performance Measures

General System

| Performance Measure | Excellent Performance | Good Performance | Marginal Performance | Poor Performance | RRS Performance |
| --- | --- | --- | --- | --- | --- |
| Hardware:* | No missed flights** | No more than 1 missed flight | No more than 3 missed flights | More than 3 missed flights |  |
| System Failures*** | No missed flights | No more than 1 missed flight | No more than 3 missed flights | More than 3 missed flights |  |
| Start-up | System is functional within 10 minutes of activation | System is functional within 15 minutes of activation | System is functional within 20 minutes of activation | System isn(t functional within 20 minutes of activation |  |
| In-flight | No loss of PTU or winds due to system malfunctions | No more than 5 min loss of PTU/winds due to system malfunctions | No more than 10 min loss of PTU/winds due to system malfunctions | More than 10 min loss of PTU/winds due to system malfunctions |  |
| Ranging | No loss of winds at less than 250 km from launch point | No more than 1 flight having loss of winds less than 250km (see in-flight) | No more than 3 flight have loss of winds less than 250km (see in-flight) | More than 3 flights have loss of winds less than 250km (see in-flight) |  |
| Shut Down | No system malfunction | No loss of PTU or winds and no more than 1 system malfunction | No more than 1 flight having loss of PTU or winds and no more than 3 system malfunctions | More than 1 flight having loss of PTU or winds and more than 3 system malfunctions |  |

*Hardware includes: TRS, SPS, RWS, Multiplexer, DGPS Antenna, interfaces, Cables, RSOIS

** A Missed flight is defined as a failure causing the sounding not to be taken during its observational window.

*** System failures include hardware and software supporting the hardware, but not the radiosonde.

COMMUNICATIONS

| Performance Measure | Excellent Performance | Good Performance | Marginal Performance | Poor Performance | RRS Performance |
| --- | --- | --- | --- | --- | --- |
| From RRS to NCF/ NOAAPORT* |  |  |  |  |  |
| Frequency | 99.0-100% | 98.9 - 98.4% | 98.3-98.0% | Less than 98% |  |
| Timeliness | 0.5 min or less for all U/A products | 0.5 - 1.0 mins | 1.0 - 1.25 min. | greater than 1.25 min |  |
| From RRS to WFO* |  |  |  |  |  |
| Frequency | 98.5%-100% | 98.4 - 98.1% | 98.0 - 97.0% | Less than 97% |  |
| Timeliness | 3 min for all U/A products | 3.1 to 5 min. | 5.1 to 7 min. | Greater than 7 min. |  |
| From RRS to NCEP |  |  |  |  |  |
| Frequency | 98.0-100% | 97.9 - 97.4% | 97.3-97.0% | Less than 97% |  |
| Timeliness | 5 min for all U/A products | 5.1 to 7 min. | 7.1 to 15 min. | greater than 15 min. |  |
| From site to NCDC (FTP) | No BUFR files missing from NCDC(s HDSS | 1 missed transmission & NCDC notifies NWS to retransmit, which is successful | 2-3 transmission missing & NCDC notifies to retransmit, which are successful. | 1 or more missed transmissions at NCDC after NCDC notifies NWS to retransmit |  |
| From RRS to TG (CLIMAT-TEMP) | N/A | 100% | N/A | 0% |  |

*Includes transmissions over back up communications paths as well.

Software Performance

| Performance Measure | Excellent Performance | Good Performance | Marginal Performance | Poor Performance | RRS Performance |
| --- | --- | --- | --- | --- | --- |
| Software Halts | None | No more than 1 missed flight | No more than 3 missed flight | Nore than 3 missed flights |  |
| In-flight s/w functions | All function nominally | No more than 1 s/w function does not operate correctly, but does not impair flight. | No more than 3 s/w functions do not operate correctly, but none impair flight. | One or more functions do not operate correctly and impair the collection or processing of flight data or coded messages |  |
| Performance in response to operator commands | No System Slow-downs | Minor System Slowdowns | Some slowdown, but but none impair flight processing of coded messages | Major System Slowdowns affecting overall performance, e.g., 15 minutes to display coded messages |  |
| Non-flight s/w functions, e.g., off-line utilities | All function nominally | No more than 1 s/w function does not operate correctly, but does not impair the flight and OBIT is functioning. | No more than 3 s/w function do not operate correctly, but does not impair the flight and OBIT is functioning. | OBIT is not functioning or another utility is not functioning and impairs the flight |  |

Radiosonde Performance

| Performance Measure | Excellent Performance | Good Performance | Marginal Performance | Poor Performance | RRS Performance |
| --- | --- | --- | --- | --- | --- |
| Data Quality: |  |  |  |  |  |
| Percent of RAOBS with 10% or more temperature levels rejected by NCEP | <1.0% | 1 to 2.9% | 3 to 3.9% | >4.0% |  |
| Percent of total temperature errors detected by NCDC | <0.3% | 0.3 to 1.9% | 2.0 to 3.9% | >4.0% |  |
| Percent of total height errors detected by NCDC | 0% | 0% | 0.05% to 0.1% | >0.1% |  |
| Percent of total pressure errors detected by NCDC | 0.5% | 0.5 to 1.2% | 1.3% to 1.9% | >2.0% |  |
| Percent of relative Humidity errors detected by NCDC | 0% | 0% | 0.05% | >0.1 |  |
| Percent of wind errors detected by NCDC | <1% | 1-3% | 3.1-5% | > 5% |  |
| Flight Performance: |  |  |  |  |  |
| No. Flights not reaching 400 hPa | None | Not more than 1 | Not more than 2 | More than 2 |  |
| Percent higher than 100 HPA | 100% | 90-99.9% | 85-89.9% | Less than 85% |  |
| Percent higher than 20 HPA | 95-100% | 90-94.9% | 85-89.9% | Less than 85% |  |
| Percent higher than 10 HPA | 80-100% | 70-79.9% | 60-69.9% | Less than 60% |  |
| Second releases* | None | No more than 2 | No more than 3 | More than 3 |  |
| Third releases* | None | No more than 1 | No more than 1 | More than 2 |  |
| Average termination pressure** | Less than 10 hPa | 10.1-15.9 hPa | 16-20 hPa | Greater than 20 hPa |  |
| Rejects during baseline*** | None | No more than 2 | No more than 5 | More than 5 |  |
| Missed Flights | None | No more than 1 | No more than 2 | More than 2 |  |

* These only include non-induced failures of any type causing a second or third release.

** It is recognized these are also a function of balloon, therefore early balloon burst will not be counted against RRS POPM.

*** Note, radiosondes rejected during baseline will be retested in the next scheduled observation window as is currently done.

Aggregate Table

| Performance Measure | Excellent Performance | Good Performance | Marginal Performance | Poor Performance | RRS Performance |
| --- | --- | --- | --- | --- | --- |
| System and software failures: | 0 missed flights | no more than 1 missed flight | 2-3 missed flights | Greater than 3 missed flights |  |
| Radiosonde failures* | 0 R/S in-flight failures | 1-2 r/s in-flight failures causing 2nd/3rd release | 3 R/S in-flight failures causing 2nd/3rd release | Greater than 3 R/S in-flight failures causing 2nd/3rd release |  |
| Comms failure (And backups) | All flights arrive @ destination: a.NCF/NOAAPORT b. NCEP c. NCDC | 1. FLT doesnt make it to either: a.NCF/NOAAPORT b. NCEP c. NCDC | 2-3 flts don(t make it to either: a.NCF/NOAAPORT b. NCEP c. NCDC | More than 3 flts don(t make it to either: a.NCF/NOAAPORT b. NCEP c. NCDC |  |

Note, for comms failures, a running total for NCF/NOAAPORT, NCEP, and NCDC will be used to determine which category should be assigned.

* Radiosonde failure reason is determined from Flight Summary data and the failure occurs at pressures greater than 400 hPa.

Applications/Model/Analyses

| Performance Measure | Excellent Performance | Good Performance | Marginal Performance | Poor Performance | RRS Performance |
| --- | --- | --- | --- | --- | --- |
| AWIPS: |  |  |  |  |  |
| U/A Profile S/W | All work currently | no more than 1 | No more than 2 | more than 2 |  |
| U/A Indice Calc. | All work currently | No more than 1 | No more than 2 | more than 2 |  |
| Decoder/database coded MSGS | Messages decoded/db without problem | No more than 1 fails to reach data base | No more than 2 fail to reach data base | more than 2 fail to reach data base |  |
| Other Applications | All work correctly | No more than 1 | No more than 2 | more than 2 |  |
| NWSTG CLIMAT |  |  |  |  |  |
| NCEP: Analyses Models Freezing Level QC routines | All work correctly | No more than 1 fails to | No more than 2 | more than 2 |  |

Appendix B - Operational Flight Series Test Schedule

| Date | Time | Ascension No. | Special Flight Charactaristics |
| --- | --- | --- | --- |
|  | Any time | None | Test mode with in-flight simulator |
|  | Any time | None | Test mode with in-flight simulator |
|  | Any time | None | Test mode with in-flight simulator |
|  | Any time | None | Test mode with in-flight simulator, Reload software and set to ops mode Asc No. = current ascension number for MicroART + 1 |
| Day 1 | 12 | 702 | Nominal Sippican Only |
| Day 1 | 0 | 703 | Nominal Sippican Only; Use Rework to retransmit corrected TTBB message |
| Day |  |  |  |
| Day |  |  |  |
| Day |  |  |  |
| Day | 0 | 707 | Nominal Sippican Only |
| Day | 12 | 708 | Nominal Sippican Only; Leave RH cap on instrument |
| Day | 0 | 1 | Nominal Sippican Only, |
| Day | 12 | 2 | Nominal Sippican Only; Reject 1 sonde during baseline and document |
| Day | 0 | 3 | Nominal Sippican Only, |
| Day | 12 | 4 | 2nd Back up comms path; Primary and 1st backup fail |
| Day | 0 | 5 | Nominal Sippican Only, |
| Day | 0 | 6 | Nominal Sippican Only, |
| Day | 12 | 7 | 2nd release flight** |
| Day | 0 | 8 | Nominal Sippican Only, |
| Day | 12 | 9 | Nominal Sippican Only, |
| Day | 0 | 10 | Nominal Sippican Only, |
| Day | 12 | 11 | 1st Back up comms path; Use Rework to transmit |
| Day | 0 | 12 | Nominal Sippican Only, |
| Day | 12 | 13 | Nominal Sippican Only, |
| Day | 0 | 14 | Nominal Sippican Only, |
| Day | 12 | 15 | Nominal Sippican Only, |
| Day | 0 | 16 | 1st & 2nd release fail, 3rd release flight is nominal** |
| Day | 12 | 17 | Nominal Sippican Only, |
| Day | 0 | 18 | Nominal Sippican Only, |
| Day | 12 | 19 | Power Failure near end of flight; Rework to transmit ABV message |
| Day | 0 | 20 | Nominal Sippican Only, |
| Day | 12 | 21 | Nominal Sippican Only; Reject 2 radiosondes during baseline** |
| Day | 0 | 22 | Nominal Sippican Only, |
| Day | 12 | 23 | 2nd release flight: |
| Day | 0 | 24 | Nominal Sippican Only, |
| Day | 12 | 25 | Nominal Sippican Only, |
| Day | 0 | 26 | Nominal Sippican Only, |
| Day | 12 | 27 | 1st Back up comms path |
| Day | 0 | 28 | Nominal Sippican Only, |
| Day | 6 | 29 | Special; Change pressure after flight using rework and then retransmit messages |
| Day | 12 | 30 | Nominal Sippican Only, |
| Day | 18 | 31 | Special, Use CD backup to retransmit all messages |
| Day | 0 | 32 | 2nd Back up comms path |
| Day | 6 | 33 | Special; Use rework to retransmit all messages |
| Day | 12 | 34 | Nominal Sippican Only, |
| Day | 18 | 35 | Special |
| Day | 0 | 36 | Nominal Sippican Only, |
| Day | 6 | 37 | Special |
| Day | 12 | 38 | Nominal Sippican Only, |
| Day | 18 | 39 | Special, 3 releases, all flights fail, reprocess best flight and transmit** |
| Day | 0 | 40 | Nominal Sippican Only, |
| Day | 12 | 41 | Nominal Sippican Only, |
| Day | 0 | 42 | Nominal Sippican Only, |
| Day | 12 | 43 | Nominal Sippican Only, |
| Day | 0 | ---- | No obs taken; Issue (no data( message |
| Day | 12 | 44 | Nominal Sippican Only, |
| Day | 0 | 45 | Nominal Sippican Only, |
| Day | 12 | 46 | Nominal Sippican Only, |
| Day | 0 | 47 | Nominal Sippican Only, |
| Day | 12 | 48 | Nominal Sippican Only, |
| Day | 0 | 49 | Nominal Sippican Only, |
| Day | 12 | 50 | Nominal Sippican Only, |
| Day | 0 | 51 | Nominal Sippican Only, |
| Day | 12 | 52 | Nominal Sippican Only, |
| Day | 0 | 53 | Nominal Sippican Only, |
| Day | 12 | 54 | Nominal Sippican Only, |
| Day | 0 | 55 | Nominal Sippican Only, |
| Day | 12 | 56 | Nominal Sippican Only, |
| Day | 0 | 57 | Nominal Sippican Only, |
| Day | 12 | 58 | Nominal Sippican Only, |
| Day | 0 | 59 | Nominal Sippican Only, |
| Day | 12 | 60 | Nominal Sippican Only, |
| Day | 0 | 61 | Nominal Sippican Only, |
| Day | 12 | 62 | Nominal Sippican Only, |
| Day | 0 | 63 | Nominal Sippican Only, |
| Day | 12 | 64 | Nominal Sippican Only, |
| Day | 0 | 65 | Nominal Sippican Only, |
| Day | 12 | 66 | Nominal Sippican Only, |

* If a second or third release should or before the scheduled fight, then the actual occurrence can replace one-for-one the scheduled condition.

** Any situations induced to create a special situation will not be used for performance measure calculations.

# Footnotes

RAOB, the complete rawinsonde observation program, is produced by Environmental Research Services, (, 1994-2002.