### LARP: Magnet-Connection Interface Documentation Guidelines

Author: A.F. Lietzke (LBNL)

(Adapted from R. Carcagno, VMTF-01-001, FNAL)

“VMTF Test Stand: Magnet/DAQ System Interface”

Review History:

| Revision A: | Date: 06/09/2005 |
| --- | --- |
| Revision B: | Date: 06/30/2005 |
| Revision C: | Date: 07/06/2005 |
| Revision D: | Date: 07/28/2005 |
| Revision E: | Date: 07/28/2005 |
| Last Reviewed: P. Wanderer, S. Feher, M. Nyman, A. McInturff | Date: 07/28/2005 |
| Approved for Trial: P. Wanderer, S. Feher, M. Nyman, A. McInturff | Date: 07/28/2005 |
| Final Approval: | Date: |

Please route feedback and revision requests to AFLietzke@lbl.gov.

------------------------------ TABLE OF CONTENTS ----------------------------

LARP: Magnet-Connection Interface Documentation Guidelines	1

PURPOSE OF Document:	3

SCOPE OF DOCUMENT	3

DOCUMENTATION EXPECTATIONS:	3

DIVISION OF RESPONSIBILITIES:	3

Summary of MAGNET & Magnet-test Sub-document	4

MECHANICAL INTERFACE Sub-document	5

POWER INTERFACE Sub-document	6

PROTECTION INTERFACE Sub-document	6

Diagnostics INTERFACE Sub-document	6

DIAGNOSTIC DEVICE CALIBRATIONs Sub-document	7

PRE-SHIPMENT Q-A MEASUREMENTS Sub-document	7

Magnet-Diagnostic Interface Conventions	7

RECOMMENDED DIAGNOSTIC WIRING PRACTICE	8

Heater ConnectorS	9

Voltage-Tap Connectors	10

Low Voltage (Strain-Gauge & RTD) Connectors	12

Schematics AND MAPS	14

# PURPOSE OF Document:

##### The purpose of this document is to present explicit guidelines and templates for creating the “Magnet-Connection Interface Documents”.  The magnet test “Connector-DAQ” interface will be addressed in a separate document.

# SCOPE OF DOCUMENT

This document specifies the following interface-connection documentation aspects and sub-sections:

- Documentation expectations

- Division of responsibilities.

- Sub-document description and guidelines.

- Documentation examples

# DOCUMENTATION EXPECTATIONS:

The “Magnet-Connection Interface” documentation is expected to include the following sub-documents:

- Introduction to magnet: magnet description, feature emphasis, primary testing purpose(s), principle test goal(s), performance expectations and operational limitations.

- Mechanical interface: requirements and features, recommended procedures for safe transportation, lifting, tilting, suspending, and securing against mechanical hazards.

- Power interface: wiring diagrams & requirements, PS requirements, appropriate current connections, insulation integrity limits.

- Protection interface: internal and connector wiring diagrams and heater-PS requirements.

- Diagnostic interface: sensor/transducer information, internal and connector wiring diagrams; internal and connector wiring labels and/or specifications)

- Diagnostics calibration tables.

- Pre-shipment magnet diagnostic performance and Q-A measurements.

# DIVISION OF RESPONSIBILITIES:

The “Magnet-Connection Interface” documentation is expected to be prepared by the magnet builder, in collaboration with personnel at the proposed test facility, in accordance with the guidelines in this document, including the following sequence and division of responsibilities:

- Magnet mfg: Follow the applicable guidelines herein described, while designing and building the magnet, paying especial attention to those mechanical, electrical, and diagnostic constraints that may impact the ability to safely test the magnet, or jeopardize the validity of the measurements or their interpretation.

- Magnet mfg. and test facility:  Jointly create a goal-directed diagnostic plan that specifies the number, kind, and location of magnet diagnostics (voltage taps, temperature sensors, strain gauges, heaters, etc) that are consistent with test facility constraints, and the test goals for the magnet.

- Magnet mfg:  Create all required magnet electrical schematics, up to (but not including) the interface connectors, with accurate, unambiguous labels for individual devices (V-tap, temp sensor etc.), and their associated wiring, according to the standards described in these guidelines.

- Test facility:  Review the above documentation, and, after agreement, specify the wiring and connector pin assignment portion of the documentation, and return the associated documentation to the magnet builder for implementation.

- Magnet builder:  After the diagnostics devices, and their intended interconnection has been approved by all parties, generate the diagnostic calibration document file and transmit it to the test facility, so the test facility can complete their signal interface documentation before the magnet test is scheduled to begin.

- Magnet builder:  Perform all the relevant Q-A and diagnostics measurements before shipping the magnet.

- Magnet builder:  Bundle a hard copy and electronic copy of the accumulated interface documentation, for arrival with the magnet at the test facility.

- Test Facility:  All documentation generated for signal creations and acquisition, or generated after magnet arrival at the test facility.  The guidelines for this documentation are documented in a separate document.

# Summary of MAGNET & Magnet-test Sub-document

The “Summary of Magnet…” sub-document is expected to contain information that will orient its readers to the general information, background, and R&D goals of the associated magnet.  The following information is expected to be included in this summary:

- Summary of “Magnet Production Report”:

  - Magnet ID/name.

  - Description of magnet (e.g., first 1 meter long prototype, 90 mm bore, accelerator-quality, Cos-Theta quadrupole magnet for upgrading the LHC’s interaction region…).

  - List of noteworthy magnet features (e.g., aluminum shell-supported Nb3Sn coils; Coils assembled with 300K pre-stress that should develop 180MPa at 4.5K; …).

  - List of noteworthy diagnostic features (e.g., first attempt to measure onset of coil-separation via coil-mounted strain-gauges; Protection heater only in the outer layer; …).

  - List, description, and significance of noteworthy production anomalies (e.g., ramp area de-cabled in coil-C; Coil-D had unusual fluffiness in its return-end; Potting voids in Coil-A’s lead-end; …).

  - Short-sample I(B) plot(s) of un-degraded strand performance at requested test temperature(s) (EXCEL format).

  - Peak-field load-line B(I) plot.

  - Anticipated RRR, dR/dl(300K), 4.2K joint resistance(s).

  - Anticipated ramp-rate sensitivity features (if available).

  - Table of calculated B-field harmonics at fields where measurements are desired.

- List of primary purposes of magnet test (e.g. compare magnet training and magnetic-field harmonics of a shell-type support-structure, and its end-support system, against a collared structure and its end-support system)...

- List of principle test goals (e.g., strain-gauge ratcheting before first training quench, quench-training curve, ramp-rate sensitivity, magnetic-field harmonic analysis, comparison of SG-ratcheting inside and outside the coils, comparison of SG-ratcheting with slip-stick frequency).

- Hipot limitations (voltage, duration, and expected safety procedures).

  - Coil to heater(s).

  - Coil to insulated support structure.

  - Coil to insulated diagnostics (assumed to be at diagnostic reference potential).

  - Coil to cryostat (and anti-cryostat, if applicable).

  - Heater to insulated support structure.

  - Heater to insulated diagnostics.

  - Heater to cryostat (and anti-cryostat, if applicable).

  - Insulated support structure to insulated diagnostics.

  - Insulated support structure to cryostat (and anti-cryostat, if applicable).

  - Insulated diagnostics to cryostat (and anti-cryostat, if applicable).

- Testing limitation information:

  - “Out-of-spec” conditions (e.g., coil Q4 had a short to its winding island; coil Q2 broke-over to its strain-gauges at 200 V).

  - Maximum allowed conductor temperature (MIIT’s, deg.C).

  - Maximum protection-heater delay (MIIT’s).

  - Maximum energy extraction delay (MIIT’s)

  - Maximum energy extraction voltage.

# MECHANICAL INTERFACE Sub-document

The “Mechanical–Interface” sub-document is expected to contain all the mechanical information that is relevant for safely and efficiently lifting, transporting, mounting, securing the magnet for testing, demounting, and repacking for its return to the manufacturer.  While the test facility is expected to supply the means for holding and securing the magnet from a cryostat’s top-plate, the magnet supplier is expected to supply the following information in a manner that facilitates timely, cost-effective magnet testing (i.e., no “last-minute” changes; or significant alteration of previous agreements, without substantial justification):

- Lifting and attachment specifications and pictures:

  - Drawings that will facilitate reliable, secure attachments at the test facility.

  - Magnet weight.

  - Recommended lifting points and requirements.

  - Recommended tilting methods and procedures.

  - Special care instructions.

  - Description of particular mechanical hazards.

- Transportation requirements:

  - Crate requirements

  - Packing and unpacking checklists.

  - Special care instructions.

  - Transportation tracking and monitor requirements.

  - Recommended transportation.

- Mechanical requirements for attaching and securing the magnet during testing:

  - Bolt patterns.

  - Hole sizes.

  - Bolt requirements.

  - Cryostat-support rigidities.

  - Maximum dimensions.

  - Minimum bore diameter

- Misc. mechanical:

# POWER INTERFACE Sub-document

The “Power Interface” sub-document is expected to contain all power-related information that is relevant for safe and efficient magnet testing, including the following:

- Intra-coil wiring diagram.

- Maximum expected voltage & current required.

- Inductance and maximum stored energy

- Maximum MIITS delay before attempting energy-extraction.

- Maximum safe energy-extraction voltage.

- Terminal voltage polarity and reference requirements/preferences.

- Splicing requirements, instructions, and/or limitations.

- Table of insulation integrity limitations.

# PROTECTION INTERFACE Sub-document

The “Protection Interface” sub-document is expected to contain all magnet-protection-related information that is relevant for safe and efficient magnet testing, including the following:

- Desired (or delivered) intra-heater and connector wiring diagram(s), and Heater-PS redundancy

- Maximum expected current for quenching at 30% of Iss.

- Expected heater-circuit resistance(s)

- Maximum-permitted heater terminal voltages.

- Polarity and voltage-reference requirements or preferences (if relevant).

- MIIT’s curve for the magnet.

- Maximum MIITS delay before heater discharge

# Diagnostics INTERFACE Sub-document

The “Diagnostics Interface” sub-document is expected to contain all diagnostics-related information that is relevant for safe and efficient magnet testing, including the following:

- Listing(s) of physical devices:

  - Type (v-tap, strain-gauge, RTD, spot-heaters, ect.)

  - Description of physical location

  - Device-ID names (preferably coded via type and physical location).

- Wiring tables, relating the above information with:

  - Wire-ID and/or color

  - Interface- connector’s ID and pin #.

- Maps of all physical device locations (e.g., coil, layer, turn-#, location relative to significant coil features).

- Detailed electrical schematic (showing electrical significance of above information).

  - Intra-coil diagnostics.

  - Structural diagnostics .

  - Environmental diagnostics

# DIAGNOSTIC DEVICE CALIBRATIONs Sub-document

- RTD: calibration curves, look-up tables, excitation limits.

- Strain-gauge: gauge-factors, calibration curves and curve fits (quadratic and cubic)

# PRE-SHIPMENT Q-A MEASUREMENTS Sub-document

- V-tap:  300K resistance sequence (wrt indicated reference).

- Temperature sensors: 4-wire resistance & temperature, full accuracy.

- Strain gauges: 4-wire resistance & strain, full accuracy.

- Heaters: 4-wire resistance

- Magnet’s R, L, Q (100Hz, 1kHz)

- Recent 300K hipot results.

- All of the pre-assembly strain-gauge results.

# Magnet-Diagnostic Interface Conventions

Magnet/coil instrumentation is connected to the test facility’s DAQ system via several blocks of Hypertronic connectors that reliably mate with blocks located within the cryostat’s magnet chamber.  Subsequent sections describe each connector in detail, with tables and schematics that document the relationships between devices, device-locations, and the assignment of device wires to connector pins.  The magnet-half of each connector module is built from Hypertronic modules that are wired to the magnet’s diagnostic and protection wires.  All modules in a block of modules are assumed to have the same sex, in a manner that the dominant power/information flow flows out of the female, into the male:

- Voltage-tap output (female): 5-pin modules, #LAFST, maximum of 20 V-taps/coil.

- Heater input and monitor (male):

  - Quench-heater: 2-pin modules, #LCMSTH??.

  - Spot-heater (& heater-status): 5-pin modules #LAMST, 10 pins/coil.

- Low-voltage diagnostics output (female): 17-pin modules, # LDFSTAH.

- Multi-pin, multi-module sequence and voltage-polarity conventions:

  - Multi-Pin voltage-sequence: When paired (or otherwise related) lower numbered pins are more positive.

  - Modules supplying external power:  When one or more pin-groups are powered, the first two pins in the group supply the power (I+, I-).

  - Multi-module sequence: Lower numbered modules are more positive.

  - 4-wire pin-sequence: I+, I-, V+, V- (in order of increasing pin-#).

- Wires & connectors: uniquely labeled/coded according to the supplied schematics.

- Pin-diagrams and labels unambiguously describe the orientation of each connector.

- No connector can be plugged in backward.

# RECOMMENDED DIAGNOSTIC WIRING PRACTICE

- Basic Connection-Reliability Rule: Provide sufficient strain-relief that will ensure connection integrity despite differential cool-down displacements and accidental cable tugging and pulling

- Basic Low-Noise Rule: Route each monitor wire along whatever path minimizes the net enclosed noise-flux between it and its reference wire.

- Low-noise voltage-tap wire-routing:  Whenever possible, route a V-tap monitor wire toward its reference V-tap, as close as possible to the coil-segment it is intended to monitor.  Upon reaching the reference V-tap, route as close as possible to its reference wire (twisting whenever possible) until reaching the associated diagnostic-attachment plug that connects to the test-facility’s DAQ-system.

- The lowest noise signals result from choosing a magnet-connector pin-sequence whose cryostat and test-facility twisted pair wiring matches the magnet’s twisted pairs.  As the twisted-pair sequencing usually starts with every connector’s #1 pin (check with the test-facility), one can choose a magnet connector pin-sequence that matches the magnet’s twisted-pairs by careful chose of pin-pairs.  For example, test-facility connector pin-pair 15-16 connect to a twisted pair, while pair 16-17 are each connected to a member of adjacent twisted-pairs in the test facility.  In order to transmit all low-noise signals along twisted pairs in the test facility, one may occasionally need to parallel adjacent pins on the magnet side of the connector.

- Splice and lead monitor wiring should be tightly routed along their respective leads (taped to the lead if needed), in a manner to minimize the loop area.

- Reliable, low-noise resistance measurements:

  - Minimize the flux-coupling area between V-sensing wires.

  - Minimize the flux-coupling area between I-excitation wires.

  - Match magnet twisted-pairs with test-facility twisted-pairs.

  - Bundle and route “low-voltage” wires well-separated from “high-voltage” wires, using extra insulation if required).

- Low-noise heater wiring: As heaters are usually pulsed with voltages between 50-500V, they generate considerable noise in nearby circuits. For minimum inductive and capacitive coupling:

  - Twist the heater excitation wires and heater monitor wires as tightly as practical, for as far as practical, matching the test-facility twisted-pairs in the manner described above  (#5).

  - Group all heater wires together, and maximize their distance from all other diagnostic wires.

- High-voltage reliability of Hypertronic V-tap and heater monitor connectors is improved by wrapping Kapton tape on the connector’s side-rails.

- Strain-relieve wires with shrink tubing at each connector-pin’s soldering-cups.

- Magnet connector conventions:

  - Manufacturer: Hypertronics, gold-plated pins.

  - Modularity:

    - Each coil has its own connector

    - Facilitates R&D coil swapping w/o altering any connectors).

  - Sex:

    - All modules in a connector have the same sex (determined by the connector’s dominant Poynting/information vector, out the female).

    - V-tap connector: Female.

    - Heater connector: Male (even for heater monitor/status)

    - Low-voltage diagnostic modules: Female.

- Multiple Hypertronic module-ordering and connector alignment:

  - Module/pin sequence follows the 17-pin Hypertronics pin-sequence:

    - Bottom to top within columns

    - Female: column # increases to the right (solder-cup view).

    - Male: column # increases to the left (solder-cup view).

  - Connector alignment pins:

    - An alignment pin with the same sex as the connector is installed nearest to pin-1, module-1 (to facilitate locating pin-1, module-1 during trouble-shooting and low-visibility conditions).

    - A second, opposite sex alignment pin is diagonally opposite (to facilitate reliable connector mating low-visibility conditions).

# Heater ConnectorS

The LARP convention (power and information out the female) requires that the magnet’s heater connections be male, while voltage-tap and low-voltage diagnostic connectors be female. The magnet side of the Heater-interface” connector (Fig.1) is representative of the type of high-voltage interface connector that must be matched by the magnet’s coil connector-modules (when they exist).

Table 1 illustrates the voltage-tap and heater pin assignment table that is provided by the magnet mfg.  Accompanying circuit schematics (e.g., Fig. 4) are expected to unambiguously show the electrical relationship between all devices and all nearby relevant magnet elements, and the interface connector pins.  Physical layout maps (e.g., Fig. 5, 6) and/or photos are expected to unambiguously show the mechanical relationship between all devices and all nearby relevant magnet elements.  How the pins are associated with test-facility signals is determined by the test facility and is documented in a subsequent signal tables.  Explicit signal requests must be negotiated with the test-facility.

FIGURE 1: FNAL’s Heater Connector Layout.  Magnet side uses male (pins).  Modules are numbered according to the Hypertronics convention.

TABLE 1: A Hypertronic Magnet-Plug Pin Assignment Table.  Wires are grouped into pairs whenever possible, with the lower pin-# of a pair being the (+) signal voltage/current, the second being its reference or return.

| Schematic Wire label | Magnet tap/wire color/label ID | Wire Color | Module # | Pin # |
| --- | --- | --- | --- | --- |
| HAouter.spotI+ | S.heater Coil-A Outer I+ | Red-Y | 1 | 1 |
| HAouter.spotI- | “ “ I- | Blk-Y | 1 | 2 |
| HAouter.spotV+ | “ “ V+ |  | 1 | 3 |
| HAouter.spotV- | “ “ V- |  | 1 | 4 |
|  |  |  |  | 5 |
| HAouterI+ | Q.heater Coil-A Outer I+ | Red | 5 | 1 |
| HAouterI- | “ “ I- | Blk | 5 | 2 |
| HBouterI+ | Q.heater Coil-B Outer I+ | Red-W | 6 | 1 |
| HBouterI- | “ “ I- | Blk-W | 6 | 2 |

# Voltage-Tap Connectors

Figure 2 illustrates a coil-module’s V-tap (female) connector layout.  The magnet builder is expected to wire each connector’s pins and modules in a manner whereby higher pin-#’s and higher module-#’s always have higher (or equal) 300K resistance relative to the magnet’s reference (negative) terminal.  The modular arrangement shown in Fig. 2 allows up to 20 (or 40) V-taps/coil.  Excess coil V-taps may spill into an adjacent coil’s connector, provided the sequential voltage rule is obeyed.  All departures from the sequential voltage rule must be explicitly noted and negotiated with the relevant test facility, as they may increase the cost of test preparation and signal debugging.

Table 2 shows a typical voltage-tap wiring table that explicitly details the wiring path from voltage-tap in the magnet to a pin in a module of an interface connector.  Accompanying circuit schematics (e.g., Fig. 4) are expected to unambiguously show the electrical relationship between all devices and all nearby relevant magnet elements, and the interface connector pins.  Physical layout maps (e.g., Fig. 5, 6) and/or photos are expected to unambiguously show the mechanical relationship between all devices and all nearby relevant magnet elements.  The method of monitoring “½-magnet”, “¼ magnet”... imbalance voltages is at the discretion of the test facility (unless otherwise negotiated beforehand).

FIGURE 2: Voltage Tap Connector Layout (solder-cup view of female)

TABLE 2: TQS01 Coil-A’s Voltage-Tap Connector Wire/Pin Assignment Table

| Schematic Wire Label | V-Tap Location/Description | Wire Color | Module # | Pin # |
| --- | --- | --- | --- | --- |
| VA101 | Coil-A Layer-1,Splice-Ti | Blk | 1 | 1 |
| VA102 | Coil-A Layer-1,Splice-Sn | Wht | 1 | 2 |
| VA103 | “” Layer-1 Turn-16 Lead-side | Red | 1 | 3 |
| VA104 | “” Layer-1 Turn-1 Lead-end | Blu | 1 | 4 |
| VA105 | “” Layer-1 Turn-1 Ret-end | Yel | 1 | 5 |
| VA106 | “” Layer-1 Turn-1 Ramp | Grn | 2 | 1 |
| VA206 | “” Layer-2 Turn-1 Ramp | Blk-W | 2 | 2 |
| VA205 | Layer-2 Turn-1 Ret-end | Wht-Bk | 2 | 3 |
| VA204 | Layer-2 Turn-1 Lead-end | Red-W | 2 | 4 |
| VA203 | Layer-2 Turn-21 Lead-side | Blu-W | 2 | 5 |
| VA202 | Coil-A Layer-1,Splice-Sn | Yel-Bk | 3 | 1 |
| VA201 | Coil-A Layer-1,Splice-Ti | Grn-W | 3 | 2 |
|  |  |  | 3 | 3 |
|  |  |  | 3 | 4 |
|  |  |  | 3 | 5 |
|  |  |  | 4 | 1 |
|  |  |  | 4 | 2 |
|  |  |  | 4 | 3 |
|  |  |  | 4 | 4 |
|  |  |  | 4 | 5 |

# Low Voltage (Strain-Gauge & RTD) Connectors

All low voltage connectors are constructed from one or more 17-pin Hypertronic module(s) (Fig. 3).  The magnet side is female.  Each 17-pin module accommodates 8-twisted pairs, with the 17th pin available to continue any shield (if desired).  Each unit (each module and each connector) is expected to be unambiguously labeled with a name that indicates it coil origin and relative sequence (if there is more than one module/connector per coil).  If power is supplied by a module, pins 1 & 2 are reserved for I+ & I-, respectively.  If additional power is supplied, the first two pins in the new set supply the power, and are appropriately indicated in the wire-table (Table 3).  Accompanying circuit schematics (e.g., Fig. 4) are expected to unambiguously show the electrical relationship between all devices and all nearby relevant magnet elements, and the interface connector pins.  Physical layout maps (e.g., Fig. 5, 6) and/or photos are expected to unambiguously show the mechanical relationship between all devices and all nearby relevant magnet elements.  External (e.g., series/parallel) connections are the responsibility of the test facility.

FIGURE 3: Strain Gauge Connector Layout

TABLE 3: Example of a Strain-Gauge Connector Wire/Pin Assignment Table

| Schematic Wire Label | Functional location-description | Wire Color | Module # | Pin # |
| --- | --- | --- | --- | --- |
| SG0101I+ | I+ for coil-01 SG-01 | Blk | 1 | 1 |
| SG0101I- | I- for coil-01 SG-01 | Wht | 1 | 2 |
| SG0101V+ | V+ for coil-01 SG-01 | Red | 1 | 3 |
| SG0101V- | V- for coil-01 SG-01 | Blu | 1 | 4 |
| SG0102I+ | I+ for coil-01 SG-01 | Blk | 1 | 5 |
| SG0102I- | I- for coil-01 SG-01 | Wht | 1 | 6 |
| SG0102V+ | V+ for coil-01 SG-01 | Red | 1 | 7 |
| SG0102V- | V- for coil-01 SG-01 | Blu | 1 | 8 |
| TS0101I+ | I+ for coil-01 TS-01 | Blk | 1 | 9 |
| TS0101I- | I- for coil-01 TS-01 | Wht | 1 | 10 |
| TS0101V+ | V+ for coil-01 TS-01 | Red | 1 | 11 |
| TS0101V- | V- for coil-01 TS-01 | Blu | 1 | 12 |
| TS0102I+ | I+ for coil-01 TS-01 | Blk | 1 | 13 |
| TS0102I- | I- for coil-01 TS-01 | Wht | 1 | 14 |
| TS0102V+ | V+ for coil-01 TS-01 | Red | 1 | 15 |
| TS0102V- | V- for coil-01 TS-01 | Blu | 1 | 16 |
| EMPTY |  |  | 1 | 17 |
| SG0201I+ | I+ for coil-02 SG-01 | Blk | 2 | 1 |
| SG0201I- | I- for coil-02 SG-01 | Wht | 2 | 2 |
| SG0201V+ | V+ for coil-02 SG-01 | Red | 2 | 3 |
| SG0201V- | V- for coil-02 SG-01 | Blu | 2 | 4 |
| SG0202I+ | I+ for coil-02 SG-01 | Blk | 2 | 5 |
| SG0202I- | I- for coil-02 SG-01 | Wht | 2 | 6 |
| SG0202V+ | V+ for coil-02 SG-01 | Red | 2 | 11 |
| SG0202V- | V- for coil-02 SG-01 | Blu | 2 | 10 |
| TS0201I+ | I+ for coil-02 TS-01 | Blk | 2 | 9 |
| TS0201I- | I- for coil-02 TS-01 | Wht | 2 | 8 |
| TS0201V+ | V+ for coil-02 TS-01 | Red | 2 | 7 |
| TS0201V- | V- for coil-02 TS-01 | Blu | 2 | 12 |
| Ts0202I+ | I+ for coil-02 TS-01 | Blk | 2 | 13 |
| Ts0202- | I- for coil-02 TS-01 | Wht | 2 | 14 |
| Ts0202V+ | V+ for coil-02 TS-01 | Red | 2 | 15 |
| Ts0202V- | V- for coil-02 TS-01 | Blu | 2 | 16 |
| EMPTY |  |  | 2 | 17 |

# Schematics AND MAPS

Figure 4 illustrates a coil wiring sequence, and how voltage-taps are wired to an interface connector at FNAL.  Similar schematics are needed for the all devices, so the appropriate meaning and significance can be determined and assigned to the signals that are recorded.

FIGURE 4: Fixed Voltage Taps, Twisted Pair & Connector-Pin Assignments at FNAL.

Pictures and maps with readable, reliable labels are worth a thousand words of description (Fig.5, 6).

FIGURE 5: Physical map of voltage taps, and heater layout for layer-2 of LBNL TQ01 quadrupole coil.

FIGURE 6: Enlarged physical map of lead-end voltage-taps shown in Fig. 5.

# Text boxes

Twisted

Pairs