##### QADATA Summary – August 2008

The Inorganic Blind Sample Project (IBSP) of the Branch of Quality Systems reviews about 90 inorganic analytical lines at the National Water Quality Laboratory (NWQL) using double-blind, quality-control samples.  A summary of this review, the QADATA Summary, is a short report commenting on only those analytical lines for which the quality-control charts indicate or have indicated a bias and/or higher than expected variability.  Overall, the vast majority of the inorganic methods at the NWQL demonstrate acceptable variability and little or no bias.

The data reviewed in this summary is for the time period 01/01/08 through 07/31/08; however, longer-range charts may be used when they add clarity to the problem.  The attached quality-control charts contain data from samples that have been subject to the same laboratory handling procedures as the samples submitted by the USGS district offices, and the data have the same level of quality-control review as the data released to the districts.  Therefore, any factors that affect the data displayed on the IBSP quality-control charts should also affect the districts’ data in a similar manner.  Data from samples known to have been affected by a problem independent of the NWQL are removed from the charts.

In this report, the “Problem” is determined and written by the IBSP and then sent to the NWQL.  The NWQL then gets two weeks to respond to the issue by writing an “NWQL Comment” and a “Corrective Action.”  Comments from the NWQL are incorporated into the QADATA Summary as received.  The most recent comments and responses are in bold followed by comments and responses from previous QADATA Summaries.  Approximately one month after the QADATA Summary is distributed, the NWQL and IBSP personnel usually meet to follow-up on corrective actions and on how the corrective actions are affecting subsequent data.

Current charts (updated twice weekly) may be accessed at  http://bqs.usgs.gov/bsp/Controlcharts.htm.  Laboratory and district personnel may have their own custom quality-assurance data retrieval to parallel their objectives by contacting Ted Struzeski (303.236.1872, struzesk@usgs.gov).  Charts pertaining specifically to this report can be accessed electronically by following the underlined hyperlinks embedded within the report and associated with each analysis discussed.

Historical QADATA Summaries may be accessed at http://bqs.usgs.gov/bsp/QAData.htm.

Tedmund M. Struzeski

struzesk@usgs.gov

Table of Contents

Report Highlights 	Page

One-Page Summary	3

Detailed Description by Analyte

Trace Metals

Iron, Filtered,  ICP, Labcode 0645, Parmcode 01046PLA11	4

Manganese, Filtered,  ICP, Labcode 0648, Parmcode 01056PLA11	4

Nickel, Unfiltered,  cICP-MS, Labcode 3131, Parmcode 01067PLM11	5

Strontium, Unfiltered,  ICP-MS, Labcode 2387, Parmcode 01082PLM48	5

Uranium, Filtered,  ICP-MS, Labcode 1797, Parmcode 22703PLM43	5

Vanadium, Unfiltered,  cICP-MS, Labcode 3135, Parmcode 01087PLM11	6

Vanadium, Unfiltered,  ICP, Labcode 2370, Parmcode 01087PLM15	6

Zinc, Filtered,  cICP-MS, Labcode 3138, Parmcode 01090PLM10	7

Report Highlights

Note:  Click the underlined text to access the control chart associated with the comment.

- Iron, filtered, ICP, parmcode 1046PLA11– Sharp downward trend and increased variability starting late February 2008.

- Nickel, unfiltered, cICP-MS, parmcode 01067PLM11 – Blind blanks exhibit 33% false positive rate starting April 2008.

- Strontium, unfiltered, ICP-MS, parmcode 01082PLM48 – Blind blanks show 30% false positive rate starting March 2008.

- Vanadium, unfiltered, cICP-MS, parmcode 01087PLM11 – Blind blanks demonstrate 65% false positive rate starting November 2007.

- Vanadium, unfiltered, ICP, parmcode 01087PLM15 – Blind blanks show 24% false positive rate starting mid November 2007.

- Zinc, filtered, cICP-MS, parmcode 01090PLM10 – Positive bias for more than 18 months.  Current bias is 11%.

NWQL PERFORMANCE ON IBSP SAMPLES

Note:  Click the underlined text in the left-hand column to access the control chart for the associated comment.

| Trace Metal Determinations | Problem | NWQL Comment | Corrective Action |
| --- | --- | --- | --- |
| Iron, Filtered ICP labcode 0645 parmcode 1046PLA11 | The results continue to demonstrate a negative bias and high variability. Since the last assessment, the bias is 16%, on average, below the target value; and 15 out of 22 results (68%) are outside of -2 f-pseudosigmas. The same reference material is submitted for analysis for unfiltered iron by ICP (pc 1045PLA15), and the data show no bias. | Dilutions on our third-party check showed no bias, but there had been maintenance performed on the instrument prior to the test. New calibrants have also been made. These actions have brought iron back within the acceptance range, but blind results are still showing a low bias on BQS charts and no bias on internal charts. Internal charts and BQS charts are based on samples at different concentrations (500 ug/L for the internal charts versus 38 – 208 ug/L for the BQS charts) and this may indicate a bias at the lower concentrations. | NWQL will use blank water from different sources and use this water to make new calibration standards. We will also test the source of current blank water by a different technology (method). |
| Previous comments for iron, filtered, labcode 0645, parmcode 01046PLA11 |
| From QADATA Summary June 2008 | Results received since the last assessment indicate a sharp negative trend and increased variability. As of late February 2008, the bias is 7.7%, on average, below the target value; and nine out of 30 results (30%) are outside of -2 f-pseudosigmas. This variability and bias are not evident for unfiltered iron by ICP (pc 1045PLA15). | On-line charts do not show this trend, but the references used are at higher concentrations than those used by BQS. | We will dilute one of our third-party-check samples to match the concentrations commonly found in the BQS check samples and perform multiple analyses for comparison. We will also check our calibration standards by a different analytical method. Outcomes of these tests will determine our next steps. |
| Manganese, Filtered ICP labcode 0648 parmcode 01056PLA11 | A slight negative bias has developed as of July 2008. Since then, the bias is 5.7%, on average, below the target value. | Internal charts show no bias, but BQS samples are at a much lower concentration (250 ug/L for the internal charts vs 8 - 38 ug/L for the BQS charts). This may indicate a bias at the lower concentrations. | NWQL will use blank water from different sources and use this water to make new calibration standards. We will also test the source of current blank water by a different technology (method). |
| Nickel, Unfiltered cICP-MS labcode 3131 parmcode 01067PLM11 | Since April 2008, a false positive rate of 33% is indicated for blind blanks. No false positives are evident prior to April 2008. The results for non-blank reference materials show acceptable bias and variability. | We performed maintenance on the instrument that initially appears to have improved this analysis. | NWQL will continue to monitor this analysis. |
| Strontium, Unfiltered ICP-MS labcode 2387 parmcode 01082PLM48 | Since March 2008, a false positive rate of 30% is indicated for blind blanks. No false positives are evident prior to March 2008. | We performed maintenance on the instrument that initially appears to have improved this analysis. | NWQL will continue to monitor this analysis. |
| Uranium, Filtered ICP-MS labcode 1797 parmcode 22703PLM43 | Corrective actions taken by the NWQL appear to have improved the negative bias; however, a slight positive bias is now present. Analysis should be closely monitored until the next QADATA Summary is distributed. |  | NWQL will continue to monitor this analysis. |
| Previous comments for uranium, filtered, labcode 1797, parmcode 22703PLM43 |
| From QADATA Summary June 2008 | The data demonstrate a slight negative bias. As of mid March 2008, results are on average 5.3% below the target value, and all results are less than the target value. | On-line charts show the same trend and the analyst has already made new calibration standards that appear to have corrected this slight bias. | New standards have been made and the analysis will be monitored to ensure that the noted bias has been corrected. |
| Vanadium, Unfiltered cICP-MS labcode 3135 parmcode 01087PLM11 | High variability and a positive bias are still shown for blind blanks. Since November 2007, a 64% false positive rate is demonstrated for blind blanks. The IRL is currently listed as 0.1 ug/L. | We performed maintenance on the instrument that initially appears to have improved this analysis, but data indicate that the IRL of 0.1 ug/L may be changed to an LT-MDL of 0.5 ug/l for the next water year. | The IRL may be changed to an LT-MDL of 0.5 ug/l for the next water year. |
| Previous comments for vanadium, unfiltered, labcode 3135, parmcode 01087PLM11 |
| From QADATA Summary June 2008 | The high variability and positive bias are still demonstrated for blind blanks. Since November 2007, 20 out of 31 results (65%) are above the IRL. | After consultation with the BQS the Interim Reporting Level for this analysis will be changed to 0.5 ug/L to more accurately reflect actual performance. | IRL will be changed to 0.5 ug/L |
| From QADATA Summary April 2008 | High variability and a positive bias are shown for blind blanks. Since November 2007, a 72% false positive rate is demonstrated for blind blanks. | The IRL of 0.1 ug/L was an estimate when the analytical system was put on-line in September of 2006 and most likely incorrect. This IRL may need to be raised. | We will work with BQS to evaluate the IRL and change it if appropriate. |
| From QADATA Summary February 2008 | High variability and a positive bias are shown for blind blanks. Since November 2007, 62% of the blind-blank results are greater than the IRL. | We are seeing this with our blanks as well. We believe that the variability is instrumental and have contacted the manufacturer to have maintenance performed. If the variability is related to the instrument aging, we may need to raise the IRL for this element. | We will monitor the instrument closely after the maintenance and take appropriate steps based on the information gathered at that time. |
| Vanadium, Unfiltered ICP labcode 2370 parmcode 01087PLM15 | For blind blanks, all results received since the last summary are below the LTMDL. Analysis should be closely monitored until the next QADATA Summary is distributed. |  | NWQL will continue to monitor this analysis. |
| Previous comments for vanadium, unfiltered, labcode 2370, parmcode 01087PLM15 |
| From QADATA Summary June 2008 | Although the variability for blind blanks has slightly improved, the overall false positive rate since mid November 2007 is 24%. No false positives are evident prior to November 2007. | In November of 2007 periodic maintenance was performed and a number of adjustments were made, including rebuilding the torch. It is possible that the PM caused a change in the ability to analyze for vanadium. | We will contact the service representative for assistance. |
| From QADATA Summary April 2008 | Since mid November 2007, a false positive rate of 29% is indicated for blind blanks. No false positives are evident prior to November 2007. | On-line blanks show no change, but we will review the sample digestion and analysis processes to look for possible causes for the change in the blind-blank results. | We will review the sample digestion and analysis processes to look for possible causes and we will continue to monitor this analysis. |
| Zinc, Filtered cICP-MS labcode 3138 parmcode 01090PLM10 | The high variability has improved to normal levels with all results received since the last assessment within +/- 2 f-pseudosigmas. However, the long-term positive bias is still present. The most recent data show a positive bias of 11.3%, on average, above the target value. The same reference material is submitted for analysis for unfiltered zinc by cICP-MS (pc 1092PLM11), and the data do not show this bias. | Charts indicate that variability has been reduced and the method is operating within acceptance criteria. BQS blind samples indicate a high bias, on-line charts show a slight low bias. Since the BQS blind RA samples do not show a similar bias the NWQL will analyze some BQS RA samples using the FA method and some FA samples by the RA method. Results should help narrow the possible causes. | Since the BQS blind RA samples do not show a similar bias the NWQL will analyze some BQS RA samples using the FA method and some FA samples by the RA method. Results should help narrow the possible causes. |
| Previous comments for zinc, filtered, labcode 3138, parmcode 01090PLM10 |
| From QADATA Summary June 2008 | This method has demonstrated a positive bias for over 16 months. Although the positive bias has slightly improved, the variability has increased since the last assessment. Since the last assessment, four out of the last 28 results (14%) are outside of control limits. The overall positive bias is 11%. Results for unfiltered zinc by cICP-MS (pc 1092PLM11) are not exhibiting this bias. | On-line third-party-check samples show a slight low bias. Blind blanks and on-line blanks do not show any bias. | We will purchase a NIST traceable standard from a different vendor and test BQS samples and the current on-line check samples to determine where the discrepancy is. |
| From QADATA Summary April 2008 | The positive bias and high variability are showing some improvement. Results since the last assessment exhibit a 13% positive bias on average and only one result is outside of control limits. Results for unfiltered zinc by cICP-MS (pc 1092PLM11) are not exhibiting this bias. | On-line charts show similar improvement. | We will continue to monitor this line to ensure that the improvement continues. |
| From QADATA Summary February 2008 | The positive bias continues to increase. Since the last assessment, almost half of the results are outside of control limits and the overall bias is 21%. Results for unfiltered zinc by cICP-MS (pc 1092PLM11) are not exhibiting this bias or variability. | We have looked at a number of possible causes, but none of the attempted corrections has been proven to correct the issue yet. The most recent modification to the calibration procedure is showing promise, but we want to see more data to be sure. | It will take 2-3 more weeks of data to ensure that this solution is working. We will continue to monitor this line closely and if the line is still having problems we will request input from the Quality Assurance Section of the NWQL on March 11, 2008. |
| From QADATA Summary December 2007 | The positive bias displayed since January 2007 has increased even more since the last assessment. The most recent data show a positive bias greater than 13%, and one out of every four results is outside of control limits. Results for unfiltered zinc by cICP-MS are not exhibiting this problem. | On-line QC charts indicate the bias does not appear to be related to standards. The bias also does not seem to be related to carry-over or contamination. The analyst is investigating the possibility of interference, however, the instrument is currently undergoing repair. More research needs to be conducted as soon as the instrument is operational again. | The NWQL will continue to monitor the instrument for bias. |
| From QADATA Summary October 2007 | This analytical line has shown a positive bias since January 2007. As of then, the results are on average 13% above the target value, and 89 of the 95 results are greater than the target value. Results for unfiltered zinc by cICP-MS (pc 1092PLM11) come from the same reference materials but do not exhibit this bias. | On-line QC was showing no bias so the analyst purchased another standard from a different source. Initial analyses show a bias similar to the one on the BQS charts. | The NWQL will make new calibration and check standards from new stock solutions. |
| From QADATA Summary August 2007 | Although there is some improvement, a positive bias is still present. Results received since the last assessment are on average 14% above the target value with most results greater than the target value. A positive bias is not evident for unfiltered zinc by standard ICP-MS (pc 1092PLM48) or by cICP-MS (pc 1092PLM11). | On-line QC charts indicate an equal number of points above and below the MPV at 10 and 100 ug/L. At 10 ug/L there is a difference between the 2 instruments with one trending high and one trending low. | We will monitor the line to see if new results continue to improve the charts. |
| From QADATA Summary June 2007 | The positive bias is increasing. Results received since the last assessment are on average 18% above the target value with three out of 22 results (14%) outside of +/-2 f-pseudosigmas. A positive bias is not evident for unfiltered zinc by cICP-MS (pc 1092PLM11). |  | We are making new standards and we will continue to closely monitor the analysis for bias. |
| From QADATA Summary April 2007 | The data demonstrate a positive bias. As of January 2007, results are on average 11% above the target value, and most results are greater than the target value. | NWQL QC charts indicated a similar bias and new calibration standards were prepared in the end of March 2007. We are still gathering data to see if the bias has improved. | We will continue to closely monitor the analysis for bias. |