## Documentation
Documentation
2008-04-16 | This spreadsheet was originally created by staff in the NIST Weights and Measures Division in the 1980's.
It contains the equations from SOP 5 and has been tested in numerous seminars and against the data in SOP 5.
Disclaimer
These Spreadsheets and associated materials have been prepared with a high degree of expertise and professionalism, and it is believed that the Spreadsheets provide a useful and accurate approach for metrology calculations. However, NIST WMD, and any other organization involved in their development, collectively and individually, do not warrant these Spreadsheets for any purpose, nor do they make any representations regarding their fitness for any use or purpose whatsoever. Each User agrees to decide if, when and how to use the Spreadsheets, and does so at his or her sole risk. When using the tools provided on the NIST WMD website, you agree that you are not entitled to rely on any information generated using these worksheets. You further agree to hold NIST WMD, and any of their partners in the creation of the tools, harmless for loss you might suffer arising out of any inaccuracies in numbers generated by the worksheets. Under no circumstances shall NIST WMD, or any of their partners that helped create the tools, be liable for any damages, including incidental, special or consequential damages, arising from the use of these Spreadsheets or an inability to use them.
If you distribute these tools through any means other than the NIST WMD website at www.nist.gov/labmetrology, you should check the website to ensure the tool being provided is the latest version available, and provide information to users on how to check for updates and revisions to the tools.

## 3-1 Design, SOP 5
3-1 WEIGHING DESIGN, 7/21/2000 [Press Ctrl W to use entry form.]
Procedure: SOP 5, NIST Handbook 145, Uncertainty Conforms to ISO Guide.
Densities used to compute air buoyancy corrections. CIPM 81/91 Air Density.
** Initial Parameters **
Date: |  |  | 2000-07-21
Balance used: |  |  | AT 1005
Load (g): |  |  | 1000
Temperature (C): |  |  | 20
Pressure (mmHG): |  |  | 753.5
Relative Humidity (%): |  |  | 45
True Mass of SW (mg): |  |  | 50.086
Density of SW (g/cm3): |  |  | 8.41
Nominal Value of Weights (g): |  |  | 1000
Mass of S (g): |  |  | 999.9985
Uncertainty of Mass of S (mg): |  |  | 0.032667 | (Enter "one sigma" value)
Density of S (g/cm3): |  |  | 8
Accepted Value for Mass of Sc (g): |  |  | 1000.0023 | (From control chart or NIST reference value.)
Density of Sc (g/cm3): |  |  | 8
Density of X (g/cm3): |  |  | 7.84
** Tare Weights **
Mass of Ts (g): |  |  | 0
Density of Ts (g/cm3): |  |  | 1
Mass of Tx (g): |  |  | 0
Density of Tx (g/cm3): |  |  | 1
Mass of TSc (g): |  |  | 0
Density of TSc (g/cm3): |  |  | 1
** Standard Deviation Values **
Pooled ST SD (mg), s(w): |  |  | 0.023 | Process s.d. for F test
Pooled Balance SD (mg), s(p): |  |  | 0.1 | From control chart, for check std
Other standard deviations, u(o): |  |  | 0.05 | Any other factors to be included in uncertainty calculations
** Observations **
Double Substitution
| S vs X | S vs Sc | X vs Sc
Obs-1 | 10 | 10.3 | 15.5 | Additionally needed for Mass Code:
Obs-2 | 15.3 | 14 | 14.1 | --cubical coefficient of expansion for volume correction to 20 C
Obs-3 | 65.3 | 64.1 | 64 | --between time standard deviation, s(b)
Obs-4 | 60.1 | 60.4 | 65.6
** Calculated Values ** |  |  |  | OLD: SOP 2 | for comparison purposes only
Air Density (mg/cm3) (CIPM 81/91): |  |  | 1.1894 | 1.1894
A1: |  |  | -5.2583
A2: |  |  | -3.6984
A3: |  |  | 1.5054
Short Term SD (mg): |  |  | -0.031426
F - Statistic: |  |  | 1.8669 | Passed F-test
dsc (mg): |  |  | 3.7166
Mass of Sc (g): |  |  | 1000.0022
Uncertainty (mg): |  |  | 0.23296 | Root Sum Square, k=2
¦Observed - Accepted¦ (mg): |  |  | 0.082849 | Check Std Passes: Bias < U, k=1
dx (mg): |  |  | 5.2401
Nominal Value (g): |  |  | 1000
Mass Correction for X (mg): |  |  | 6.7757
AMx vs. Brass Correction (mg): |  |  | -3.275
AMx vs. 8.0 Correction (mg): |  |  | 3.714
Uncertainty (mg): |  |  | 0.23296

## AIR
|  | Differences |  | Air Density Calcs |  |  | TN 577: | HB 145: | CIPM 1981/91 | SOP 2, intm step | M(a) | R | Kelvin (T) | Pa | x(v) | Z
|  | CIPM-H 145 | CIPM-TN 577 | P (mm Hg): | Temp (C): | RH (%): | mg/cm3 | mg/cm3 | mg/cm3 | e(s): |  |  | 273.15+C | 101325/760 |  |  | a(0) | a(1) | a(2) | b(0) | b(1) | c(0) | c(1) | d | e | p(sv) | p(sv)/P | f
0.0001 | -0.0001 | 8.4323e-06 | -0.0025453 | 750.0615 | 20 | 50 | 1.186 | 1.1835 | 1.1835 | 17.5357 | 28.9635 | 8.3145 | 293.15 | 100000 | 0.011742 | 0.99962 | 1.5812e-06 | -2.9331e-08 | 1.1043e-10 | 5.707e-06 | -2.051e-08 | 0.00019898 | -2.376e-06 | 1.83e-11 | -7.65e-09 | 2339.1632 | 0.023392 | 1.004 | 0.023485 | 1 | 0.00038269 | 2.0067e-06
0.0001 | -0.0001 | 2.4795e-05 | -0.002614 | 825.0677 | 20 | 10 | 1.3092 | 1.3066 | 1.3066 | 17.5357 | 28.9635 | 8.3145 | 293.15 | 110000 | 0.0021357 | 0.99961 | 1.5812e-06 | -2.9331e-08 | 1.1043e-10 | 5.707e-06 | -2.051e-08 | 0.00019898 | -2.376e-06 | 1.83e-11 | -7.65e-09 | 2339.1632 | 0.021265 | 1.0043 | 0.021357 | 1 | 0.00039429 | 2.5717e-06
0.0001 | -0.0001 | 9.462e-05 | -0.0039844 | 750.0615 | 15 | 90 | 1.2064 | 1.2023 | 1.2024 | 12.802 | 28.9635 | 8.3145 | 288.15 | 100000 | 0.015411 | 0.99955 | 1.5812e-06 | -2.9331e-08 | 1.1043e-10 | 5.707e-06 | -2.051e-08 | 0.00019898 | -2.376e-06 | 1.83e-11 | -7.65e-09 | 1705.6691 | 0.017057 | 1.0039 | 0.017123 | 1 | 0.00044703 | 1.9852e-06
0.0001 | -0.0001 | -9.0386e-05 | -0.0014019 | 450.0369 | 25 | 50 | 0.69556 | 0.69425 | 0.69416 | 23.7676 | 28.9635 | 8.3145 | 298.15 | 60000 | 0.02649 | 0.99977 | 1.5812e-06 | -2.9331e-08 | 1.1043e-10 | 5.707e-06 | -2.051e-08 | 0.00019898 | -2.376e-06 | 1.83e-11 | -7.65e-09 | 3169.757 | 0.052829 | 1.0029 | 0.05298 | 1 | 0.00023193 | 5.2371e-07
0.0001 | -0.0001
0.0001 | -0.0001 | 6.2847e-06 | -0.0025393 | 753.5 | 20 | 45 | 1.192 | 1.1894 | 1.1894 | 17.5357 | 28.9635 | 8.3145 | 293.15 | 100458.4284 | 0.01052 | 0.99962 | 1.5812e-06 | -2.9331e-08 | 1.1043e-10 | 5.707e-06 | -2.051e-08 | 0.00019898 | -2.376e-06 | 1.83e-11 | -7.65e-09 | 2339.1632 | 0.023285 | 1.004 | 0.023378 | 1 | 0.00038082 | 2.0496e-06
0.0001 | -0.0001 | -6.4011e-05 | -0.0020227 | 620 | 22 | 45 | 0.97258 | 0.97062 | 0.97056 | 19.8283 | 28.9635 | 8.3145 | 295.15 | 82659.888 | 0.014451 | 0.99969 | 1.5812e-06 | -2.9331e-08 | 1.1043e-10 | 5.707e-06 | -2.051e-08 | 0.00019898 | -2.376e-06 | 1.83e-11 | -7.65e-09 | 2645.1797 | 0.032001 | 1.0035 | 0.032112 | 1 | 0.00030694 | 1.31e-06
0.0001 | -0.0001 | -5.4719e-05 | -0.0021073 | 650 | 22 | 45 | 1.0199 | 1.0178 | 1.0178 | 19.8283 | 28.9635 | 8.3145 | 295.15 | 86659.56 | 0.013785 | 0.99968 | 1.5812e-06 | -2.9331e-08 | 1.1043e-10 | 5.707e-06 | -2.051e-08 | 0.00019898 | -2.376e-06 | 1.83e-11 | -7.65e-09 | 2645.1797 | 0.030524 | 1.0036 | 0.030634 | 1 | 0.00031996 | 1.4523e-06
0.0001 | -0.0001 | -4.039e-05 | -0.0022182 | 690 | 22 | 45 | 1.083 | 1.0808 | 1.0808 | 19.8283 | 28.9635 | 8.3145 | 295.15 | 91992.456 | 0.012988 | 0.99966 | 1.5812e-06 | -2.9331e-08 | 1.1043e-10 | 5.707e-06 | -2.051e-08 | 0.00019898 | -2.376e-06 | 1.83e-11 | -7.65e-09 | 2645.1797 | 0.028754 | 1.0038 | 0.028863 | 1 | 0.00033737 | 1.6524e-06
0.0001 | -0.0001
0.0001 | -0.0001 | 9.806e-06 | -0.0025563 | 760 | 20 | 45 | 1.2023 | 1.1997 | 1.1998 | 17.5357 | 28.9635 | 8.3145 | 293.15 | 101325.024 | 0.01043 | 0.99962 | 1.5812e-06 | -2.9331e-08 | 1.1043e-10 | 5.707e-06 | -2.051e-08 | 0.00019898 | -2.376e-06 | 1.83e-11 | -7.65e-09 | 2339.1632 | 0.023086 | 1.004 | 0.023179 | 1 | 0.00038384 | 2.0868e-06
0.0001 | -0.0001 | -3.9501e-07 | -0.0024732 | 760 | 21 | 45 | 1.1978 | 1.1954 | 1.1954 | 18.6507 | 28.9635 | 8.3145 | 294.15 | 101325.024 | 0.011095 | 0.99963 | 1.5812e-06 | -2.9331e-08 | 1.1043e-10 | 5.707e-06 | -2.051e-08 | 0.00019898 | -2.376e-06 | 1.83e-11 | -7.65e-09 | 2488.0592 | 0.024555 | 1.004 | 0.024655 | 1 | 0.00037577 | 2.0597e-06
0.0001 | -0.0001 | -9.9891e-06 | -0.002407 | 760 | 22 | 45 | 1.1934 | 1.191 | 1.191 | 19.8283 | 28.9635 | 8.3145 | 295.15 | 101325.024 | 0.011795 | 0.99963 | 1.5812e-06 | -2.9331e-08 | 1.1043e-10 | 5.707e-06 | -2.051e-08 | 0.00019898 | -2.376e-06 | 1.83e-11 | -7.65e-09 | 2645.1797 | 0.026106 | 1.0041 | 0.026212 | 1 | 0.00036795 | 2.0313e-06
0.0001 | -0.0001 | -1.4544e-05 | -0.0023804 | 760 | 22.5 | 45 | 1.1912 | 1.1888 | 1.1888 | 20.4415 | 28.9635 | 8.3145 | 295.65 | 101325.024 | 0.01216 | 0.99964 | 1.5812e-06 | -2.9331e-08 | 1.1043e-10 | 5.707e-06 | -2.051e-08 | 0.00019898 | -2.376e-06 | 1.83e-11 | -7.65e-09 | 2726.9423 | 0.026913 | 1.0041 | 0.027023 | 1 | 0.00036414 | 2.0166e-06
0.0001 | -0.0001 | -1.5435e-05 | -0.0023756 | 760 | 22.6 | 45 | 1.1907 | 1.1884 | 1.1884 | 20.5661 | 28.9635 | 8.3145 | 295.75 | 101325.024 | 0.012234 | 0.99964 | 1.5812e-06 | -2.9331e-08 | 1.1043e-10 | 5.707e-06 | -2.051e-08 | 0.00019898 | -2.376e-06 | 1.83e-11 | -7.65e-09 | 2743.5572 | 0.027077 | 1.0041 | 0.027187 | 1 | 0.00036338 | 2.0136e-06
0.0001 | -0.0001
0.0001 | -0.0001 | -3.602e-05 | -0.0024197 | 760 | 22 | 10 | 1.1975 | 1.1951 | 1.1951 | 19.8283 | 28.9635 | 8.3145 | 295.15 | 101325.024 | 0.0026212 | 0.99966 | 1.5812e-06 | -2.9331e-08 | 1.1043e-10 | 5.707e-06 | -2.051e-08 | 0.00019898 | -2.376e-06 | 1.83e-11 | -7.65e-09 | 2645.1797 | 0.026106 | 1.0041 | 0.026212 | 1 | 0.00034474 | 2.1506e-06
0.0001 | -0.0001 | -2.9614e-05 | -0.0024171 | 760 | 22 | 20 | 1.1963 | 1.194 | 1.1939 | 19.8283 | 28.9635 | 8.3145 | 295.15 | 101325.024 | 0.0052424 | 0.99965 | 1.5812e-06 | -2.9331e-08 | 1.1043e-10 | 5.707e-06 | -2.051e-08 | 0.00019898 | -2.376e-06 | 1.83e-11 | -7.65e-09 | 2645.1797 | 0.026106 | 1.0041 | 0.026212 | 1 | 0.0003505 | 2.132e-06
0.0001 | -0.0001 | -1.4326e-05 | -0.0024095 | 760 | 22 | 40 | 1.194 | 1.1916 | 1.1916 | 19.8283 | 28.9635 | 8.3145 | 295.15 | 101325.024 | 0.010485 | 0.99964 | 1.5812e-06 | -2.9331e-08 | 1.1043e-10 | 5.707e-06 | -2.051e-08 | 0.00019898 | -2.376e-06 | 1.83e-11 | -7.65e-09 | 2645.1797 | 0.026106 | 1.0041 | 0.026212 | 1 | 0.00036411 | 2.0576e-06
0.0001 | -0.0001 | 4.2509e-06 | -0.0023985 | 760 | 22 | 60 | 1.1916 | 1.1892 | 1.1892 | 19.8283 | 28.9635 | 8.3145 | 295.15 | 101325.024 | 0.015727 | 0.99962 | 1.5812e-06 | -2.9331e-08 | 1.1043e-10 | 5.707e-06 | -2.051e-08 | 0.00019898 | -2.376e-06 | 1.83e-11 | -7.65e-09 | 2645.1797 | 0.026106 | 1.0041 | 0.026212 | 1 | 0.00038049 | 1.9337e-06
0.0001 | -0.0001 | 2.0329e-05 | -0.0023882 | 760 | 22 | 75 | 1.1898 | 1.1874 | 1.1875 | 19.8283 | 28.9635 | 8.3145 | 295.15 | 101325.024 | 0.019659 | 0.99961 | 1.5812e-06 | -2.9331e-08 | 1.1043e-10 | 5.707e-06 | -2.051e-08 | 0.00019898 | -2.376e-06 | 1.83e-11 | -7.65e-09 | 2645.1797 | 0.026106 | 1.0041 | 0.026212 | 1 | 0.0003946 | 1.8083e-06
0.0001 | -0.0001
0.0001 | -0.0001
0.0001 | -0.0001
0.0001 | -0.0001