Fermilab

DES Lab A LN2 Test

Operating Procedures

| Rev | Date | Description | Originated by | Approved by |
| --- | --- | --- | --- | --- |
| None | Aug 7, 2008 | Original issue | T. Tope H. Cease |  |

Reviewed by: ________________________

Date:  ______________________________

DES Lab A LN2 Test

Operating Procedures

Table of Contents

1.0	Introduction	x

2.0	Flow schematic	x

3.0	Required Training Prior to Operating the System	x

4.0	Fill Procedure	x

5.0	Drain Procedure	x

6.0	Heat Exchanger Cool Down Procedure	x

7.0	Liquid Circulation Start up	x

8.0	Abnormal Conditions Procedures	x

DES Lab A LN2 Test

Operating Procedures

1.0	Introduction

The Lab A LN2 System is used to perform preliminary tests of the cooling system and to cool the Multi-CCD Test Vessel. The MCCDTV is a prototype camera vessel for the Dark Energy Survey Camera.  The LN2 reservoir and valvebox are located outside lab A. A triple jacketed hose that contains supply, return and vacuum jacket is used a transfer line from the valve box to the heat exchanger inside Lab A. The heat exchanger is a simple tube heat exchanger and is used to cool the MCCDTV.

2.0 Flow Schematic

The flow schematic for the cryostat is shown in Figure 2.1.  The flow schematic is also available at http://des-docdb.fnal.gov:8080/cgi-bin/ShowDocument?docid=1537 as “Process and Instrument Diagram - ME-436389REVA”. Valves and instruments are referenced by number in the operating procedures.

Figure 2.1: DECAM Lab A LN2 test flow schematic.

3.0	Required Training Prior to Operating the System

A general overview of the Lab A LN2 Test system and familiarity with the FESHM chapters are required prior to operating the system. A general overview can be obtained with on the job training and reviewing all of the documents associated with the cryogenic safety review which can be obtained at http://des-docdb.fnal.gov:8080/cgi-bin/ShowDocument?docid=1537 or near the Lab A LN2 Test controls desk.

As described in FESHM 5032 “Cryogenic personnel should attend the Fermilab safety courses, "Oxygen Deficiency Hazards," "Cryogenic Safety" and "Pressurized Gas Safety" or their equivalents.  Other courses, as described in Chapters 4010 and 4030, may also be appropriate.  General training in cryogenic principles may also be beneficial, particularly to personnel involved in operations.”

FESHM 5031, Pressure Vessels, includes a general description of what requirements are expected of a system that contains a pressure vessel. The operator should be familiar with the information contained within FESHM 5031.

FESHM 5101, Personal Protective Equipment (PPE), outlines protective equipment needed while working with cryogens. Cryogenic gloves and a full face shield shall be worn at all times while operating manual valves or working around the safety relief devices. Long pants and closed toe shoes are recommended while working with cryogenics.

Before any of the following procedures are performed, get training from a person experienced with the complete system.

4.0 Fill Procedure

The 200L Vessel is filled using 160L portable dewars.  Station the 160L dewar near the fill line on the 200L vessel. The first time the 200L vessel is filled, two 160L dewars are needed to completely fill the system.

Prior to performing a fill,

Don face shield, and cryogenic gloves.

Check that the 160L fill dewar is pressurized above 200psig,

Attach an all metal transfer line with the proper flare fittings, and have wrenches available to tighten the fittings.

- Open fill valve MV-6204

- Open Fill try-cock valve MV-6208

- Slowly open the valve on the 160L dewar to start transferring liquid into the 200L vessel Note: The 200L is full when liquid vents from the fill try-cock valve.

- Slowly close the valve on the 160L dewar .

- Close fill valve MV-6204.

- Close fill try-cock valve MV-6208.

- Wait a few minutes, using the wrenches slowly open the flare fittings and remove the transfer line.

5.0	Drain Procedure

When the cooling system is stopped, most of the liquid drains back into the 200L vessel. If maintenance is needed on the vessel, the liquid is drained from the system.

Prior to draining the dewar, don face shield, cryogenic gloves. The fog produced when venting the vessel is extremely cold and creates a low oxygen pressure area. Ensure that this operation takes place slow enough that a hazard is not created.

- Slowly open the fill try-cock valve MV-6208 to vent excess pressure from the 200L vessel.

- Slowly open valve MV-6205 and let the liquid vent from the vessel. MV-6208 is also used to control the pressure in the dewar and the venting rate.

- When all of the liquid has been vented, close valves MV-6205 and MV-6208.

6.0 Heat Exchanger Cool Down Procedure

This procedure is used to cool down all of the piping in the system prior to operations. The 200L vessel is pressurized, liquid from the vessel is then pushed thru the piping system. The boil-off gas is vented using MV-6307. Prior to operating manual valves, don face shield, and cryogenic gloves.

- Close valves MV-6307, 6208, 6205, 6204, 6310, 6302, 6303.

- Open valves MV-6216, 6201.

- Let the system pressurize to the normal operating pressure, pressure can be built up using heater 6102 and setting the set-point on TE-6211 to 100K.

- Slowly Open return line vent valve MV-6307 and allow the piping system to cool down. Approximately 20 Liters are needed to boil-off to cool all of the piping down.

- Once the system is cold, slowly open MV-6303 and close MV-6307

7.0	Liquid Circulation Start up

Liquid circulation can start up once the system has been pressurized and cooled down. The pump is initially started by circulating all of the liquid thru the bypass valve MV-6310. Once the pressure stabilizes in the 200L vessel then liquid can slowly be circulated through the rest of the piping system until full design flow is achieved. The heater in the heat exchanger is not used until the full design flow is achieved. Prior to operating manual valves, don face shield, and cryogenic gloves.

- Pressurize the system and cool down all of the piping using the Cool Down Procedure.

- Close valves MV-6302, 6307, Valve MV-6303 is left open ¼ turn so that a trapped volume in the piping system is not created

- Open control valve MV-6310 half way

- Slowly start pump 6103, increase speed until design flow rate is achieved. Confirm flow rate using PDI-6313 and the pump flow curve.

- Slowly open valve MV-6302

- Once the pressure in the vessel has stabilized, slowly open valve MV-6303

- Adjust flow rate to the design flow rate using control valve MV-6310 and confirming the flow rate using FT-6301 and PDI-6313.

- Once the flow rate and pressure in the vessel are stable, the heat exchanger can be operated.

8.0	Abnormal Conditions Procedures

Small leak, Alert other people, Evacuate area,

Large Leak, Alert other people, Evacuate area, call safety -3131

Dewar failure, Alert other people, Evacuate area, call safety -3131