DESIGN CRITERIA DOCUMENT

WBS 1.4.2

Drift Tube Linear Accelerator

Date:  October 2000

SNS-104020000-DC0001-R00

DESIGN CRITERIA DOCUMENT

WBS 1.4.2

Drift Tube Linear Accelerator

# Tom Ilg

# Date Published: August 2000

DESIGN CRITERIA DOCUMENT

WBS 1.4.2

Drift Tube Linear Accelerator

# August 2000

Linac Systems Senior Team Leader

Linac Systems Mechanical Group Leader_____________________________________

DTL Systems Lead Engineer_______________________________________________

Contents

1.0 Scope	5

2.0 Functions and system description	6

2.1 RF Structure (WBS 1.4.2.2)	7

2.1.1 Post Coupler Assembly	8

2.1.2 RF Slug Tuners	8

2.1.3 Endwalls	8

2.1.4 Waveguide Iris	8

2.1.5 Monitor Loops	8

2.2 DTL Drift Tubes (WBS 1.4.2.3)	8

2.3 Vacuum System  (WBS 1.4.2.4)	8

2.4 Water System (WBS 1.4.2.5)	9

2.5 Mechanical Support (WBS 1.4.2.6)	9

2.6 Alignment System (WBS 1.4.2.6.2)	9

2.7 Assembly (WBS 1.4.2.7)	9

2.8 Installation (WBS 1.4.2.8)	9

3.0 Design Requirements	9

3.1 DTL Requirements	9

3.1.1 RF Structure (WBS 1.4.2.2)	10

3.1.1.1 	RF surfaces	11

3.1.1.2	Vacuum seals	11

3.1.1.3	Permanent RF seals	11

3.1.1.4	Demountable RF seals	11

3.1.2	DTL Tank	11

3.1.2.1 	Tank diameter	11

3.1.2.2 	RF performance	11

3.1.2.3	Tank material	12

3.1.2.4	Vacuum load effects	12

3.1.2.5	 Vacuum pumping ports	12

3.1.3	Slug tuner	12

3.1.3.1	Slug tuner locations	12

3.1.3.2	Vacuum integrity	13

3.1.4 	Post couplers	13

3.1.4.1	Post coupler locations	13

3.1.4.2	Post coupler assembly	13

3.1.4.3	Vacuum integrity	13

3.1.5	Endwalls	13

3.1.5.1	Endwall locations	14

3.1.5.2	Vacuum integrity	14

3.1.6	Waveguide iris	14

3.1.6.1	Waveguide iris performance	14

3.1.6.2	Input power feed connection	14

3.1.6.3	Waveguide iris location	14

3.1.6.4	Iris dimensions	14

3.1.6.5	Vacuum integrity	14

3.1.7	Monitor Loop	15

3.1.7.1	Monitor Loop Performance	15

3.1.7.2	Dimensions of Monitor Loop Interface	15

3.1.7.3	Monitor Loop Assembly	15

3.1.7.4	Vacuum Integrity	15

3.1.2 Drift Tubes (WBS 1.4.2.3)	15

3.1.2.1	Drift-Tube Stems	15

3.1.2.2	Drift-Tube Locations	16

3.1.2.3	Quadrupole Magnets	16

3.1.2.4	Quadrupole Gradients	16

3.1.2.5	Alignment Fiducials	16

3.1.2.6	Dimensions of Drift-Tube/Magnet Assemblies	16

3.1.2.7	Steering dipoles	16

3.2 Vacuum System (WBS 1.4.2.4)	17

3.3 Water System (WBS 1.4.2.5)	18

3.4 Mechanical systems (WBS 1.4.2.6)	19

3.4.1	Beam centerline location	19

3.4.2	Travel range	19

3.4.3	Adjustments	19

3.4.4	Support stand fabrication requirements	19

3.4.5	Installation of stand	19

3.4.6	Electrical grounding	20

3.4.7	Dimensions of Support Stand	20

3.5 Alignment system (WBS 1.4.2.6.2)	20

3.5.1 Alignment requirements	20

3.5.2 Alignment environment	20

3.6 Structural requirements	20

3.6.1	Seismic requirements	20

3.6.2	Vibration and stability	21

3.6.3	Floor loading	21

3.7 Assembly	21

3.7.1 Assembly	22

3.7.2 Alignment	22

3.7.3 Low power RF tuning	22

3.7.3.1 Low power adjustable endwalls	22

3.7.3.2 Low power adjustable post couplers	22

3.7.3.3 Low power adjustable slug tuners	22

3.7.3.4 Bead-perturbation measurements	22

3.7.4 Assembly and tuning environment	22

3.7.5 Shipping	22

3.7.6 Pre-installation check out	23

3.8 Installation	23

4.0 System Interfaces	24

4.1 Interfaces within Linac Systems	24

4.1.1	RF Power	24

4.1.2	Coupled-Cavity Linac	24

4.1.3	Diagnostics	24

4.2 Interfaces With Other SNS Systems	24

4.2.1 Front-End Systems	24

4.2.2 Conventional Facilities	24

4.2.3 Control Systems	24

5.0 Quality Assurance	25

5.1 General	25

5. 2 Tests	25

5.2.1	Engineering Tests	25

5.3 Quality Conformance Inspections	25

5.3.1	Certifications	25

6.0  Safety Requirements	25

7.0 Documentation	26

7.1 Specifications	26

7.2 Engineering Drawings and Associated Lists	26

7.3 Technical Manuals and Procedures	26

7.3.1	Procedures	26

7.3.2	Manuals	26

7.4.4 	Documentation Numbering	26

8.0 Operations and Maintenance Requirements	26

8.1 Operating Modes	26

8.2 Surveillance and in-service inspection requirements	26

8.3 General Maintenance	26

8.4 Abnormal Events and Recovery Procedures	26

9.0 Installation Requirements	27

10.0 Acceptance Testing Requirements	27

11.0 Commissioning Requirements	27

12.0 Codes and Standards	27

12.1 SNS documents	27

12.2 Non-SNS documents and standards	27

Appendix A	28

Drift tube cell geometry table	28

Appendix B	29

DTL Magnet, Post coupler, and Diagnostics schematic Layout	29

List of Tables

Table 3.1.2 DTL RF Structure Summary	10

Table 3.1.3 Tank RF Power	11

Table 3.1.4 DTL slug tuner configuration	12

Table 3.1.5 Post coupler configuration	13

Table 3.1.6 Waveguide Iris Performance Requirements	14

Table 3.1.7 Monitor Loop Performance Requirements	15

Table 3.1.8 Drift Tube/Magnet Assembly Alignment Tolerances	16

Table 3.2.1 Vacuum pressure limits along the beamline in the DTL.	17

Table 3.2.2 Summary of the DTL vacuum system design parameters.	17

Table 3.3.3 DTL drift tubes nominal water cooling system design and operation parameters.	18

Table 3.5.1 DTL alignment budget	20

DESIGN CRITERIA DOCUMENT

WBS 1.4.2

.0-1	Specification Tree;

DRIFT TUBE LINEAR ACCELERATOR

OF THE

SPALLATION NEUTRON

SOURCE

# 1.0 Scope

This document establishes the performance, design, development, and test requirements for the Drift Tube Linear Accelerator (DTL). The linac follows the MEBT and accelerates the H( beam from an energy of 2.5 MeV to 970 MeV. Two types of RF structure perform this task. The first structure is a 402.5-MHz DTL that accelerates the beam to 86.8 MeV. Figure 1.1 shows a general schematic of the Spallation Neutron Source (SNS) Front End and Linac System Layout.

Figure 1.1 SNS Front End and Linac System Layout

# 2.0 Functions and system description

The DTL is the first of two accelerators that comprise the Linear Accelerator of SNS. It accelerates the beam from 2.5 MeV to 86.8 MeV. Figure 2.1 is a general layout of the 86.8 MeV DTL system and figure 2.2 is a cut-away view of DTL tank 1 assembly showing the functional components.

Figure 2.1 86.8 MeV DTL system layout

Figure 2.2 DTL tank 1 assembly partial cut-away view

1. Drift tube assembly  2. Drift tube mount assembly  3. Upstream endwall 4. Endwall PMQ 5. Drift tube coolant fitting   6. Endwall coolant fitting  7. Drift tube stem  8. Drift tube alignment target hole  9. Slug tuner  10.Vacuum pumping slots  11. 300 l/s ion pump  12. Ion pump spool piece 13. 6” gate valve with 8”conflat flanges  14. 300 l/s Turbo pump  15. Turbo pump spool piece with instrumentation ports  16. Endwall alignment target hole  17. MEBT gate valve flange  18. Endwall RF and vacuum seals  19. Endwall coolant channels  20. Drag support strut  21. Welded support structure  22. Tank wall coolant channel  23. Tank section interface flange  24. Drag strut support fitting  25. Waveguide iiris assembly  26. 402.5 MHz window interface  27. Iris waveguide vacuum pumping slots  28. . Iris waveguide turbo pump  29. Iris waveguide vacuum  pump spool piece  30. Iris waveguide NEG vacuum pump  31. Post coupler assembly  32. Monitor loop  33.Veritcal support strut  34. Veritcal support fitting  35. Downstream endwall  36. Lateral support strut and fitting  37. Iris waveguide mount flange  38. Tank 1b section  39. Tank 1a section. 40. Tank section RF and vacuum seals  41. Tank section alignment bushing  42. Tank inside surface or diameter  43. Support stand alignment target hole  44. Waveguide iris turbo pump gate valve. 45. Intertank beam box  46. Intertank diagnostics  47. Intertank beam box vacuum pump  48. Endwall alignment pin  49.  Waveguide iris ion gauge.    (water system not shown)

The 402.5-MHz DTL will be based on the Alverez design. Cells are of length (( and the transverse focusing period has a length of 6(( at 402.5 MHz. A focusing period spans six RF gaps and has a focusing permanent  magnet quadrupole

(PMQ) inside each of the first two drift tubes, followed by a empty drift tube, followed by defocusing quadrupoles inside the next two drift tubes, and again followed by a empty drift tube. This layout is called a FFODDO lattice. In some cases the empty drift tubes will be instrumented with steering electromagnet dipoles (EMD) and beam position monitors (BPM) diagnostics.  The last four empty drift tubes of each DTL tank will incorporate an EMD. Starting with DTL tank 2 the first 2 empty drift tubes in each tank will incorporate a BPM. See appendix B for general magnet locations. The DTL will have post couplers to stabilize and tune the longitudinal field distribution.

## 2.1 RF Structure (WBS 1.4.2.2)

The DTL RF structure WBS consists of all items, except for the drift tube, which is covered in WBS 1.4.2.3, that will see RF fields (tanks, end walls, tuners, post couplers, and the waveguide iris). The RF Structure provides a stable platform for an array of drift-tube assemblies, post coupler assemblies and RF slug tuners. The structure, or tank, also provides support and interfaces to other beam path components, the RF system, the water system, and the vacuum system. The structure contains passages to flow temperature controlled water to regulate the temperature for frequency control.

### 2.1.1 Post Coupler Assembly

Post couplers provide longitudinal stability to the tank electric field by providing a transverse field that couples with the drift tube body. The number of post couplers required per DTL Tank varies, see appendix B for general layout of post couplers per tank.

### 2.1.2 RF Slug Tuners

One type of static tuners are used to initially tune the tank. Slug tuners, accessible from outside and underneath the tank, are used to provide static frequency adjustment to the RF cavity. Slug tuner penetration length is determined by Aluminum dummy slug tuners used in the low power tuning process. Adjusting the slug tuner penetration length is not required during operation.

### 2.1.3 Endwalls

Each tank has two endwalls located at the upstream and downstream end of the DTL assembly. A total of twelve endwalls are required. Endwalls provide a vacuum enclosure and RF surface terminations at each end of the DTL tank assemblies.  Endwalls provide for upstream and downstream cell geometry and incorporate a half drift tube geometry and shape. The endwalls are part of the FFODDO lattice structure and are required to provide housings for PMQ’s. Toriod diagnostics will be housed in empty endwalls where available per the FFODDO lattice layout.

### 2.1.4 Waveguide Iris

Each tank has one waveguide iris to couple the RF power to the tank from the waveguide. Each iris consists of the iris coupling slot, a tapered waveguide transition, a half height WR2100 waveguide flange, and an vacuum pumping system located near the 402.5 MHz window interface to ensure the proper vacuum pressure in the region.

### 2.1.5 Monitor Loops

Each vessel has several monitor loops to sample the cavity field strength at several points along the tank for feedback control and RF phase and amplitude measurements.

## 2.2 DTL Drift Tubes (WBS 1.4.2.3)

The DTL is predominately comprised of a series of drift tubes, which the proton beam passes through. In the DTL the beam undergoes periods of acceleration separated by periods of drifting. Oscillating the electrical charge of the drift tubes generates an electric filed between the drift tubes which accelerates the beam particles, and the particles drift at a constant velocity while inside the tubes because the electric fields are shielded from the interior of the tubes. Thus, the tubes are referred as drift tubes.  The PMQ’s are located inside the drift tubes and provide transverse beam focusing within the DTL.  The PMQ’s must withstand the operational radiation environment. Twenty-four steering EMD’s will be used to steer the beam, and they will be located in the last four open drift tubes in each tank. Two empty drift tubes downstream of the EMD drift tubes will have BPM’s installed.

A total of 210 drift tubes assemblies are unevenly distributed within the six tanks in a FFODDO lattice. Each drift-tube assembly consists drift-tube body, a cylindrical copper can, and a stem that supports the drift-tube from the tank and allows water coolant to enter and return from the drift tube body.

## 2.3 Vacuum System  (WBS 1.4.2.4)

The function of the vacuum system is to minimize interactions between the beam and residual gases within the DTL.  The vacuum system for each tank consists of a set of ion pumps to maintain the vacuum of each tank at the proper operating level. Each tank has a valve and connection for a roughing system, that will evacuate the tank from atmospheric pressure to a level appropriate to start the ion pumps. Each tank has a valve and connection for backfilling the tank. Instrumentation is provided for monitoring valve status and vacuum pressure.

## 2.4 Water System (WBS 1.4.2.5)

The cooling system must keep the DTL at an operating temperature of about 80( F to 85( F. Temperature control and stability must satisfy the requirements for maintaining cavity resonance control.  The system interfaces with the RF control system.

## 2.5 Mechanical Support (WBS 1.4.2.6)

The mechanical support system is required to support the DTL tanks and ancillary equipment including vacuum pumps, cooling water manifolds, and electrical cables. The Support structure is required to provide adequate stability that meets drift tube stability criteria as well as provide seismic restraint to the DTL system.

## 2.6 Alignment System (WBS 1.4.2.6.2)

The alignment system is required to measure and properly align each drift tube within each DTL tank. External alignment features, i.e. optical target tooling holes, will be provided on each DTL tank for proper alignment within the LINAC tunnel. The external alignment features will be referenced to individual DTL endwall surfaces and/or PMQ magnetic centers.

## 2.7 Assembly (WBS 1.4.2.7)

The DTL assembly process is responsible for assembling all components from the RF structure, drift tube assemblies, mechanical support, and portions of the water system. Drift tubes shall be alignment per WBS 1.4.2.6.2 followed by the low power tuning process. The low power tuning process will provide DTL frequency measurements and adjustments, and field measurements and adjustments. The WBS is required to provide the shipping of each assembled and tuned DTL to ORNL. The shipping method shall protect the contents from all damages, scratches, and marks etc. during the packaging and shipping process. Contents shall be stabilized against movement, and prevented from any mechanical or chemical damage during transportation.

## 2.8 Installation (WBS 1.4.2.8)

The installation process is responsible for installing and locating all 6 DTL tanks from the assembly process (WBS 1.4.2.7) into the LINAC tunnel including providing longitudinal and transverse alignment, with respect to a gobal coordinate system or to the appropriate component(s). All remaining DTL components from WBS 1.4.2.6 Mechanical systems, WBS 1.4.2.5 Water systems, and WBS 1.4.2.6 DTL Vacuum systems will be install.

# 3.0 Design Requirements

## 3.1 DTL Requirements

Table 3.1 is a summary of the major DTL requirements.  Appendix A contains the DTL drift-tube cell geometer requirements.

Table 3.1.1 DTL Requirements

| Parameters | Value |
| --- | --- |
| Input Energy | 2.5 MeV |
| Structure frequency (MHz) | 402.5 |
| Output Energy (MeV) | 86.8 |
| Average beam current | 2.08 mA |
| Peak macropulse current | 52 mA |
| Average macropulse current | 36 mA |
| Macropulse repetition rate | 60 Hz |
| Macropulse beam period | 1.0 msec |
| RF Structure Beam duty factor | 7.02 % |
| Physical length (m) | 36.5 |
| Number of tanks | 6 |
| Total number of cells | 214 |
| Bore radius (cm) | 1.25 |
| Focusing lattice | FFODDO |
| Lattice period length | 6(( |
| Quadrupole Gradient (kG/cm) | 3.5 |
| First quadrupole for H- in the vertical plane | Focusing |
| Number of PMQs | 147 |
| Accelerating gradient Eo (MV/m) | 3.6 |
| Ave. acceleration rate (MeV/m) Beam Height | 2.38 50 inches |

### 3.1.1 RF Structure (WBS 1.4.2.2)

Table 3-2 is a summary of the parameters of the RF structure.

Table 3.1.2 DTL RF Structure Summary

| Total Number of Drift tubes | 210 |
| --- | --- |
| Tank #1 Drift tubes + endwalls | 59 + 2 |
| Tank #1 Length | 4.2 m |
| Tank #2 Drift tubes + endwalls | 47 + 2 |
| Tank #2 Length | 6.0 m |
| Tank #3 Drift tubes + endwalls | 33 + 2 |
| Tank #3 Length | 6.3 m |
| Tank #4 Drift tubes + endwalls | 27 + 2 |
| Tank #4 Length | 6.4 m |
| Tank #5 Drift tubes + endwalls | 23 + 2 |
| Tank #5 Length | 6.3 m |
| Tank #6 Drift tubes + endwalls | 21 + 2 |
| Tank #6 Length | 6.3 m |
| Magnet Type | Permanent Magnet Quadrupole, Sm2CO17 |
| Quad Gradient | 3.7 kG/cm |
| Lattice | FFODDO |
| Tank #1 Post couplers | 20 |
| Tank #2 Post couplers | 23 |
| Tank #3 Post couplers | 16 |
| Tank #4 Post couplers | 27 |
| Tank #5 Post couplers | 23 |
| Tank #6 Post couplers | 21 |
| Total DTL Length | 36.5 m |
| Drift-Tube Bore Radius | 2.5 cm |
| Tank Inside Diameter | 43.472 cm for tanks 1 & 2, 45 cm for tanks 3 thru 6 |
| Kinetic Energy after Tank #1 | 7.523 MeV |
| Kinetic Energy after Tank #2 | 22.838 MeV |
| Kinetic Energy after Tank #3 | 39.789 MeV |
| Kinetic Energy after Tank #4 | 56.579 MeV |
| Kinetic Energy after Tank #5 | 72.495 MeV |
| Kinetic Energy after Tank #6 | 86.828 MeV |

.1.1-1	Diagram of DTL Tanks;

#### 3.1.1.1 	RF surfaces

All RF surfaces shall have a minimum electrical conductivity of 95% IACS (International Annealed Copper Standard). Sharp corners shall be smoothed with a minimum radius of .06” to .09”. Surface roughness shall not exceed 60-inch, and buffing or abrasive polishing is not permitted. Plating, if required, shall be copper electro-deposited from an acid bath to a minimum thickness of .002” to .003” (50 m to 80m). Intermediate nickel plating strikes are permissible. Zinc strikes shall not be used. Plating thickness shall be inspected using a calibrated eddy-current type tester or equivalent.

#### Vacuum seals

All vacuum seals shall be of either metallic or elastomer construction providing the following requirements are meet.

- The cumulative radiation dose limit for selected materials do not exceed 2.16 x 106 Rads for thirty years.

- The total leak rates and outgassing rates shall be minimized to meet the required operating base pressure of 1.84(10-7 Torr (N2).

- Dual-function (RF and vacuum) seal design is allowed providing items 1 and 2 above are meet.

#### Permanent RF seals

RF joints shall be flush with the RF surface and made by welding or brazing with high conductivity alloys.

#### Demountable RF seals

All RF seals shall be of metallic construction. Demountable RF seals shall be used only where permanent joints are impractical. Demountable seals shall be designed around a sealing element with stored spring energy. Dual-function (RF and vacuum) seal design is allowed. RF seal surfaces are to be plated with ether sliver, copper, or gold.

### DTL Tank

#### 3.1.2.1 	Tank diameter

The inside diameter of tanks 1 and 2 shall be 17.115(.010 inches. The inside diameter of tanks 3 thru 6 shall be 17.717(.010 inches. The inner surface of each tank shall meet a cylindricity tolerance of  (.010”. Should a tank, for ease of manufacture or assembly, be constructed of more than one beamline section, the total concentricity tolerance of the assembled sections of the tank shall be (.005” or less.

.1.1-1	DTL Tank Summary;

#### 3.1.2.2 	RF performance

Each tank shall resonate at a frequency of 402.5 MHz at the nominal operating temperature and vacuum pressure. The RF power required for each tank at accelerating gradient shall be less than or equal to the values in Table 3-3.

.2.1.1.1-1	Tank RF Power;

Table 3.1.3 Tank RF Power

| Tank No. | Accelerating gradient (MeV/m) | RF Power in MW |
| --- | --- | --- |
| 1 | Ramped field, 1.13 to 2.97 | 0.53 |
| 2 | Ramped field, 2.98 to 3.6 | 1.64 |
| 3 | 3.6 | 1.93 |
| 4 | 3.6 | 1.93 |
| 5 | 3.6 | 1.87 |
| 6 | Ramped field, 3.6 to 3.77 | 1.88 |

##### 3.2.1.6.5	 Performance parameters

The performance parameters of the tanks are as specified in the SNS parameters list, SNS100000000-PL-0001-R03. Each DTL tank shall be water cooled to provide resonance frequency control.

#### 3.1.2.3	Tank material

The tanks shall be constructed of low carbon steel and the endwalls of OFE copper. The design shall allow a safety factor in accordance with the applicable sections of the ASME Boiler and Pressure Vessel Code, Section VIII. The tanks shall withstand repeated high vacuum cycling and/or temperature cycling without deleterious effects on the tank inner surfaces, vacuum integrity, RF characteristics, or structural rigidity.

#### 3.1.2.4	Vacuum load effects

Allowable deflections of the tank wall and endwalls shall be determined by analysis. Specifications shall be maintained for the following: longitudinal field distribution, drift-tube alignment and tank resonant frequency.

#### 3.1.2.5	 Vacuum pumping ports

The apertures of the vacuum pumping ports shall be slotted to minimize the RF fields in the pumping ports. The slot orientation shall be in the same direction as the RF currents. The total vacuum pumping conductance per port shall be 1000 liters/sec or higher to allow for adequate pumping to achieve the operational base pressure.

### Slug tuner

Slug tuners shall have enough tuning range to correct for all possible frequency errors resulting from manufacturing tolerances. Minimum required tuning range for all slug tuners in each tank should be ( 2 MHz. Slug tuner diameter shall be (3.75 inches. Aluminum dummy slug tuners used in the low power tuning process determines slug tuner penetration length. The slug tuners will carry significant RF currents and will require water-cooling. The slug tuner assembly shall be constructed of OFHC copper and allow for water coolant to enter an exit the tuner assembly.

#### 3.1.3.1	Slug tuner locations

Tank 1 shall have 8 slug tuners. Tanks 2 thru 6 shall have 12 slug tuners. Slug tuners shall be approximately equally spaced along the longitudinal axis of each tank. See table 3.1.4 for the slug tuner configuration.

Table 3.1.4 DTL slug tuner configuration

| Tank number | Number of slug tuners | Location |
| --- | --- | --- |
| 1 | 8 | Approximately equally spaced |
| 2 | 12 | Approximately equally spaced |
| 3 | 12 | Approximately equally spaced |
| 4 | 12 | Approximately equally spaced |
| 5 | 12 | Approximately equally spaced |
| 6 | 12 | Approximately equally spaced |

#### 3.1.3.2	Vacuum integrity

The total leak rates and outgassing rates shall be minimized to meet the required operating base pressure of 1.84(10-7 Torr (N2).

### 3.1.4 	Post couplers

The post-stabilizers shall be capable of changing the field distribution to meet the SNS parameters list, SNS100000000-PL-0001-R03.

#### 3.1.4.1	Post coupler locations

The number locations of post couplers required per DTL Tank varies, see table 3.1.5 and appendix B for general location of post couplers. The post couplers shall be nominally located on the drift-tube centerlines. Post coupler axis orientation and stem penetration shall be determined experimentally in the DTL cold model.

Table 3.1.5 Post coupler configuration

| Tank number | Number of post couplers | Location |
| --- | --- | --- |
| Tank 1 | 19 | Every third drift tube starting on drift tube #3 of tank 1, alternating sides of tank thereafter |
| Tank 2 | 23 | Every two drift tubes starting on drift tube #2 of tank 2, alternating sides of tank thereafter |
| Tank 3 | 16 | Every two drift tubes starting on drift tube #2 of tank 3, alternating sides of tank thereafter |
| Tank 4 | 27 | Every drift tube starting on drift tube #1 of tank 4, alternating sides of tank thereafter |
| Tank 5 | 23 | Every drift tube starting on drift tube #1 of tank 5, alternating sides of tank thereafter |
| Tank 6 | 21 | Every drift tube starting on drift tube #1 of tank 6, alternating sides of tank thereafter |

All post couplers start on right side of each DTL tank

#### 3.1.4.2	Post coupler assembly

The nominal stem diameter of the post coupler shall be .50”. The post coupler shall have an appropriately sized "tab" determined by the cold model. The post couplers will carry significant RF currents and will require water-cooling. The post coupler assembly shall be constructed of OFHC copper and allow for water coolant to enter an exit the post coupler stem.

#### 3.1.4.3	Vacuum integrity

The total leak rates and outgassing rates shall be minimized to meet the required operating base pressure of 1.84(10-7 Torr (N2).

### 3.1.5	Endwalls

The endwalls shall be constructed of OFE copper. The endwalls shall withstand repeated high vacuum cycling and/or temperature cycling without deleterious effects on the endwall inner surface, vacuum integrity, or RF characteristics. Endwall nose deflections due to thermal gradients and vacuum loading shall be minimized to reduced end cell frequency shifts. Maximum half-cell frequency shift shall be -500 kHz. The endwalls shall be water cooled to control the nose deflections due RF heating.

#### Endwall locations

Endwalls are located at each end of the each DTL tank. Each endwall shall have the appropriate half drift geometry required for the respective cell location. See appendix A for drift tube geometry information.

#### 3.1.5.2	Vacuum integrity

The total leak rates and outgassing rates shall be minimized to meet the required operating base pressure of  1.84(10-7 Torr (N2).

### 3.1.6	Waveguide iris

#### 3.1.6.1	Waveguide iris performance

The iris shall have the performance specified in Table 3.1.6. An appropriately sized vacuum pumping system shall be incorporated to ensure a pressure of 1.84 x10-8 torr or less in the region between the window and the iris.

.2.1.3.1-1	Input Power Coupler Performance Requirements;

Table 3.1.6 Waveguide Iris Performance Requirements

| Frequency | 402.5 Megahertz |
| --- | --- |
| Peak Power | 2.0 Megawatts |
| Pulse Width | 1.6 milliseconds |
| Repetition Rate | 60 Hertz |
| Max. duty factor | 7.02% |
| Average power transmitted | 140 kW |

#### 3.1.6.2	Input power feed connection

The waveguide iris shall mate with the 402.5 MHz window using a half-height WR2100 waveguide flange.

#### 3.1.6.3	Waveguide iris location

The irises shall be located at approximately the tank midpoints. The coupler shall be positioned 45o from the top, clockwise as viewed from the beam input end of the tank. The iris shall be removable from the tank to allow inspection, repair and replacement.

#### Iris dimensions

The iris slot size shall be nominally .079 inches wide by 3.188 inches long. Final iris dimensions shall be determined during the low power RF tuning process.

#### 3.1.6.5	Vacuum integrity

The total leak rates and outgassing rates shall be minimized to meet the required operating base pressure of  1.84(10-7 Torr (N2).

### 3.1.7	Monitor Loop

#### 3.1.7.1	Monitor Loop Performance

The monitor loops shall have the performance specified in Table 3.1.7

.2.1.3.2-1	Monitor Loop Performance Requirements;

Table 3.1.7 Monitor Loop Performance Requirements

| Type | H type loop |
| --- | --- |
| Output Power at full field | 2 Watts |
| Pulse Width | 1 millisecond |
| Repetition Rate | 60 Hertz |
| Output Impedance | 50 Ohms |

#### 3.1.7.2	Dimensions of Monitor Loop Interface

There shall be six (6) monitor loops per tank equally spaced along the length of the tank.  They shall be located 58.5o from the horizontal centerline, counterclockwise as viewed from the beam input end of the tank. The loops shall be removable from the tank to allow inspection, repair and replacement.

#### 3.1.7.3	Monitor Loop Assembly

The following materials shall be used:

Conductors:	OFHC copper

Insulators:	>96% Al2O3

The monitor loop shall be flanged to allow removal. The monitor loop shall mate with a type N connector.

#### 3.1.7.4	Vacuum Integrity

The total leak rates and outgassing rates shall be minimized to meet the required operating base pressure of  1.84(10-7 Torr (N2).

### 3.1.2 Drift Tubes (WBS 1.4.2.3)

The drift-tubes shall provide a vacuum tight enclosure for the quadrupole magnet. The vacuum integrity of the drift-tubes shall be verified using helium leak detection. Unproven materials or processes shall be thoroughly tested for sparking, vacuum compatibility, and mechanical and electrical properties. The drift-tubes shall have the dimensions as specified in Appendix A of this document.

#### 3.1.2.1	Drift-Tube Stems

The stems shall be constructed of stainless steel. The Drift-tube stem area exposed to RF and vacuum shall be copper plated. They shall have a water passage in them to maintain the drift-tubes at the tank operating temperature. Nominal diameter shall be 20 mm.

#### 3.1.2.2	Drift-Tube Locations

The drift-tube assemblies shall be located according to the locations specified in Appendix A of the Linac SRD, SNS104000000-SR0001-R00. The magnetic centers shall be aligned using an in situ magnet alignment system such as a voltage-pulsed-wire apparatus, Reference "Taut Wire Alignment of Multiple Permanent Magnet Quadrupoles". The Drift-Tube/ Magnet Assembly Alignment Tolerances shall be as listed in Table 3-6.

#### 3.1.2.3	Quadrupole Magnets

Quadrupole magnets shall be of the 16 piece permanent magnet quadrupole design. Pole tip fields of up to 1.2 T are required. Magnetic field direction and pole position shall be fixed, reference Linac SRD,  and mechanical alignment techniques shall be provided for assembly into the drift-tube bodies. The quadrupole rotational axis shall be held to 0.5o with respect to the horizontal and vertical alignment. All magnets shall be characterized for harmonic content and shall be measured at 80% of aperture for N=3 @1% of total and N=4 through 6 @ 1/2% of total. Magnetic center shall be concentric to the drift-tube body to 0.15 mm or 0.3 mm Total Indicator Reading.  Magnets shall be axially centered in the drift-tube body along the beam direction to 0.15 mm.

#### 3.1.2.4	Quadrupole Gradients

The quadrupole magnets shall have the gradients as shown in Appendix A of this document. The plus (+) gradient indicates that the quadrupole is X-focusing and the minus (-) gradient indicates that the quadrupole is Y-focusing.  The tolerance on the gradient times length of a quadrupole is 1%.

.2.2.2.1.1.7-1	Drift-Tube/ Magnet Assembly Alignment Tolerances;

Table 3.1.8 Drift Tube/Magnet Assembly Alignment Tolerances

| Description of Error | Tolerance |
| --- | --- |
| Quadrupole transverse displacement | (0.07 mm |
| Quadrupole tilt | (5.8 mrad |
| Quadrupole roll | (2.5 mrad |
| Quadrupole-gradient error | (0.3% |

#### .2.2.2.1.1.7-2	Drift-Tube/ Magnet Assembly Tolerance Budget;

#### 3.1.2.5	Alignment Fiducials

Three precision holes for mounting alignment targets shall be located on each fixed end flange. Each hole and spotface shall be inspected and the locations documented with respect to the magnetic center of the end wall quadrupoles, and the RF surface of the low-energy end wall.

#### 3.1.2.6	Dimensions of Drift-Tube/Magnet Assemblies

The alignment of the quadrupole magnet, drift-tube, and stem shall be as shown TBD. For the end assemblies, the alignment of the magnet and drift-tube in the flange shall be as shown in TBD. The tolerances on the alignments shall be

#### 3.1.2.7	Steering dipoles

Each Steering Magnet shall have a BL product of 1600 Gauss-cm. All Steering Magnets must fit within an existing Drift Tube design.  In order to avoid a requirement for “soft vacuum”, the insulation system shall not be hygroscopic. The magnets must be able to withstand furnace braze temperatures with no adverse affect in performance. There will be a minimum of two horizontal and two vertical magnets per module. See appendix A for Dipole magnets locations.

## 3.2 Vacuum System (WBS 1.4.2.4)

The primary requirement for the DTL vacuum systems is to provide sufficient pumping to overcome the surface outgassing of vacuum facing components and maintain a beam tube pressure that is below the values required for a 2-mA beam operation. Table 3.4 shows the Vacuum pressure limits along the beamline in the DTL. Table 3.5 shows the summary of SNS LINAC vacuum system design parameters.

See the DTL and Coupled cavity vacuum systems design criteria document, SNS-104020400-TR0002-R01, for a further detail regarding the DTL vacuum systems requirements.

Table 3.2.1 Vacuum pressure limits along the beamline in the DTL.

| Linac Section | DTL Tank | Exit Beam Energy | Max Allowable Pressure |
| --- | --- | --- | --- |
|  | No. | MeV | Torr |
| DTL | 1 | 7.5 (starts at 2.5) | 1.84(10-7 |
| DTL | 2 | 22.8 | 1.84(10-7 |
| DTL | 3 | 39.8 | 1.84(10-7 |
| DTL | 4 | 56.6 | 1.84(10-7 |
| DTL | 5 | 72.5 | 1.84(10-7 |
| DTL | 6 | 86.8 | 1.84(10-7 |

Table 3.2.2 Summary of the DTL vacuum system design parameters.

| Parameter | DTL |
| --- | --- |
| Operating Base Pressure (after initial outgassing and RF conditioning) | Sum of weighted partial pressures of residual gases < 1.84(10-7 Torr Weighting factors: He = 0.125 H2 = 0.15 H2O = 0.66 N2 = 1.0 CO = 1.0 O2 = 1.0 CO2 = 1.5 |
| RF Structure Gas Load (During conditioning period) | Initial outgassing before RF conditioning ( 2.5(10-9 Torr-L/s/cm2 Outgassing during RF conditioning = TBD |
| (During steady-state operation following all conditioning) | Outgassing during steady-state operating = 1.0(10-10 Torr-L/s/cm2 |
| Number of structures | 6 DTL tanks with independent vacuum systems |
| Number of Pumps or Pumping Stations | Three, 300 L/s ion pumps per tank One 300 L/s turbo pump and one fore pump (turbo/roughing station) per tank One, 1000 L/s NEG pumps and one 70 L/s turbo pumps per tank for the one RF window |
| Valves | Turbo/roughing stations should be able to be isolated with valves to provide for installation and maintenance issues. Isolation valves will be provided between DTL tanks and before tank 1 and after tank 6. Note that clearance problems prohibit a valve between tanks 1 and 2. |
| RF Window Gas Load | Peak gas load during RF conditioning are TBD |
| Control System Domain | One vacuum control system (PLC, computer, electronics rack, etc) per DTL tank |
| Vacuum Interfaces | Vacuum level between DTL and MEBT is 4 x 10-7 Torr |
| Other | Possible water-cooled grid in pumping port to reduce RF losses Permanent magnet surface outgassing rate = TBD Torr/L/s/cm2 |

## 3.3 Water System (WBS 1.4.2.5)

The DTL water cooling system removes the waste heat from the copper RF structure and maintains resonance through active temperature control. A representative set of data for the heat load, water flow rates, water temperature, and resonance control parameters for the DTL drift tubes is given in Table 3.9. Further detail and discussions regarding the water system requirements and design can be found in the DTL and Coupled cavity water system design criteria document, SNS-104020400-TR0001-R01.

Table 3.3.3 DTL drift tubes nominal water cooling system design and operation parameters.

| Parameter | Value | Comments/References |
| --- | --- | --- |
| Nominal heat load per drift tube Nominal average operating temperature of drift tubes | 0.05-1.85 kW 26.6(C (79.9oF) | Heat load is different for each drift tube Average drift tube operating temperature is identical to average CCL copper temperature. |
| Nominal water flow rate per drift tube Water flow rate accuracy per drift tube Water flow rate resolution and stability per drift tube Flow resistance across drift tubes | 0.13-3.22x10-4 m3/s (0.2-5.1 gpm) ( 6.3x10-6 m3/s (( 0.1 gpm) ( 6.3x10-6 m3/s (( 0.1 gpm) 1.06x1012-3.22x1014 Pa/m6s (0.61-185.92 psi/gpm2) | Flow rate is adjusted for each individual drift tube. A uniform frequency shift for all cells within a tank is obtained by balancing the flow rate and tailoring the cooling channels for each individual drift tube. Flow resistances for each drift tube are estimated from standard pipe flow correlations. |
| Drift tube water inlet temperature prior to RF power and during steady-state, full RF power Temperature range of water Nominal temperature rise of cooling water through 1 drift tube Temperature accuracy Temperature resolution Temperature stability | 20.0(C (68.0 oF) 8.3oC (14.9oF) about mean inlet temperature 0.6-2.6oC (1.1-4.7 oF) ( 0.1oC ( 0.1oC ( 0.4oC | Initial drift tube water inlet temperature chosen to equal mean desired operating temperature. Drift tube water inlet temperature must remain constant as RF power is introduced. An approximate (8.3oC band about this mean temperature will be required for full resonance control (i.e. to get (50 kHz frequency adjustment). A 0.4oC change in water temperature corresponds to a change in RF frequency of 2.4 to 2.8 KHz. |
| Frequency shift per change in water inlet temperature Range of frequency control Stability in frequency control | 6-7 kHz/1oC ( 50 kHz ( 2.5 kHz | Resonance is maintained by dynamically adjusting the water temperature in the drift tube circuit and the tank circuit simultaneously. |
| Chilled water supply temperature Chilled water supply temperature stability | 7.2oC (45.0oF) (0.28(C (( 0.5oF) |  |

## 3.4 Mechanical systems (WBS 1.4.2.6)

### 3.4.1	Beam centerline location

Each stand shall locate the nominal beam centerline of each tank at a height of 50 inches from the floor.

### 3.4.2	Travel range

Support system for technical components must be able to provide adjustments over a 2-inch vertical range.

### Adjustments

The DTL tanks shall be adjustable in any direction by .50”, with respect to the stand. Positioning precision shall be .005”.  The support stand height shall be adjustable in the vertical direction by (1.50” using shims at the floor and stand interface.

### Support stand fabrication requirements

All welding shall conform to AWS D1.1-2000. Structural steel tubing shall conform to ASTM A-500 grade 2 or better.

### 3.4.5	Installation of stand

Each stand shall have provision for permanently securing it to a concrete floor, once the necessary adjustments have been achieved. All mounting hardware, brackets, etc., shall be contained within the envelope of the stand proper. The mounting design shall allow the use of flat washers and lock washers.

### 3.4.6	Electrical grounding

The DTL tank and its support stand must be securely grounded. A pair of "earth" ground connections shall be provided on each end of the stand. These connections shall be .500-13unc-2b threaded holes, with at least five (5) complete threads. All connections shall be on the left side of the structures, looking "downstream". The pair of holes shall be about 4.00 inches apart.

### 3.4.7	Dimensions of Support Stand

The width shall be limited to 36” and the length shall be limited to that of the tank it supports.

## 3.5 Alignment system (WBS 1.4.2.6.2)

The alignment system is required to measure and properly align each drift tube within each DTL tank. External alignment features, i.e. optical target tooling holes, will be provided on each DTL tank for proper alignment within the LINAC tunnel. The external alignment features will be referenced to individual DTL endwall surfaces and/or PMQ magnetic centers.

### 3.5.1 Alignment requirements

Table 3.3.4 shows the required alignment budget for drift tubes and DTL tanks.

Table 3.5.1 DTL alignment budget

| Alignment Parameter | Value (inches) |
| --- | --- |
| PMQ (drift tube) magnetic centers | (.005 transverse (.005 longitudinal |
| Dipole (drift tube) magnetic centers | (.005 transverse (.005 longitudinal |
| Empty drift tubes geometric centers | (.010 transverse (.010 longitudinal |
| DTL tank to tank Using endwall PMQ magnetic centers and/or Endwall geometric centers | (.005 transverse (.010 longitudinal |

### 3.5.2 Alignment environment

Drift tube alignment shall be performed in a temperature and vibration controlled environment. Daily temperature variations in the DTL during alignment shall not exceed 2 degrees Fahrenheit. Alignment shall not be performed in proximity to vibration sources such as reciprocating machinery, i.e. pumps, motors etc.

## 3.6 Structural requirements

### 3.6.1	Seismic requirements

The linac and associated systems must survive a PC-2 earthquake, 0.55g peak acceleration per DOE-STD-1020-94.

### 3.6.2	Vibration and stability

Performance of the DTL shall not degrade when exposed to the natural levels of vibration expected at the SNS site. Vibratory motion of drift tubes with internal PMQs or EMQs shall not have deleterious effects on beam quality. Drift tube stability criteria = TBD (-inch transverse RMS

Figure x-x below shows the estimated baseline vibratory environment to be used.

Figure 3.6.1 Estimated baseline vibratory environment

### 3.6.3	Floor loading

Each stand, with its DTL tank installed, shall have a floor load of less than TBD pounds per linear foot on a TBD foot x TBD foot grid or 2 line loads of TBD pounds per linear foot separated by 3 feet or 2 each TBD pound loads on a 3 foot by 3 foot grid.

## 3.7 Assembly

The DTL assembly process is responsible for assembling all components from the RF structure, drift tube assemblies, mechanical support, and portions of the water system. Drift tubes shall be alignment per WBS 1.4.2.6.2 followed by the low power tuning process. The low power tuning process will provide DTL frequency measurements and adjustments, and field measurements and adjustments. The WBS is required to provide the shipping of each assembled and tuned DTL to ORNL. The shipping method shall protect the contents from all damages, scratches, and marks etc. during the packaging and shipping process. Contents shall be stabilized against movement, and prevented from any mechanical or chemical damage during transportation.

### 3.7.1 Assembly

### 3.7.2 Alignment

See section 3.5 of this document.

### 3.7.3 Low power RF tuning

Low power RF tuning shall provide DTL frequency measurements and adjustments, and field measurements and adjustments. Bead-perturbation measurements shall be used to aid in the tuning process. Adjustable tuning components or dummy endwalls, post couplers, and slug tuner shall be used to aid in the field and frequency adjustments. All low power RF tuning measurements shall meet parameters specified in the SNS parameters list, SNS-10000000-PL001-R03.

#### 3.7.3.1 Low power adjustable endwalls

Each low endwall drift tube nose shall have an adjustable range of (0.120 inches. The endwall nose shall not come in contact with the drift tubes. Each endwall nose shall have the required geometry per appendix A of this document. The low power endwalls shall be made from Aluminum 6061-T6. No copper plating is required. RF seasl are required per sections 3.1.1.3 and 3.1.1.4 of this document.

#### 3.7.3.2 Low power adjustable post couplers

Each low post coupler shall have 360( rotational adjustments and a full tank penetration adjustment. The rotational adjustment shall have an accuracy of (0.50(. The penetration adjustment shall have an accuracy of (0.020 inches. The post coupler tip shall not come in contact with the drift tubes. The low power post coupler shall be made from Aluminum 6061-T6. No copper plating is required. Low power post couplers quantities and locations shall be per section 3.1.4 and appendix A of this document. RF seals are required per sections 3.1.1.3 and 3.1.1.4 of this document.

#### 3.7.3.3 Low power adjustable slug tuners

Each low power slug tuner shall have a tank penetration adjustment of 3.00 inches (0.020 inches. The low power slug tuner shall be made from Aluminum 6061-T6. No copper plating is required. Low power slug tuner quantities and locations shall be per section 3.1.3 of this document. RF seals are required per sections 3.1.1.3 and 3.1.1.4 of this document.

#### 3.7.3.4 Bead-perturbation measurements

The bead-perturbation measurements measure the field tilt sensitivity. Post couplers will be adjusted to set the field distribution. The bead-perturbation measurements shall be performed on each DTL tank individually. The bead-perturbation measurement will use a 0.25-inch diameter hollow aluminum bead attached to an 8-pound test nylon fishing line. The bead shall pass thru the endwalls and drift tubes of each tank at a constant rate of approximately 5 inches per second. The nylon line shall not deflect more than 0.002 inches.

### 3.7.4 Assembly and tuning environment

Assembly and RF tuning shall be performed in a temperature and vibration controlled environment. Daily temperature variations shall not exceed 2 degrees Fahrenheit. Alignment and tuning shall not be performed in proximity to vibration sources such as reciprocating machinery, i.e. pumps, motors etc.

### 3.7.5 Shipping

The shipping method shall protect each DTL assembly and all attached components from all damage, scratches, and marks etc. during the packaging, shipping, and un-packaging processes. Each DTL assembly shall be stabilized against movement, and prevented from any mechanical or chemical damage during transportation.  A temperature between 65( F and 90(F shall be maintained during the packaging, shipping, and un-packaging processes.

Further detail will be provided during title II (final design) phase of this project.

### 3.7.6 Pre-installation check out

The pre-installation check out shall consist of the following steps:

- Un-pack the DTL tank assemblies and components

- Check for external damage

- Using the bead-perturbation measurements, verify RF stability and distribution has not changed during the shipping process.

- Using the optical alignment system, verify the drift transverse alignment has not changed during the shipping process.

- Further detail will be provided during title II (final design) phase of this project

## 3.8 Installation

The installation process is responsible for installing and locating all 6 DTL tanks from the assembly process (WBS 1.4.2.7) into the LINAC tunnel including providing longitudinal and transverse alignment, with respect to a gobal coordinate system or to the appropriate component(s). All remaining DTL components from WBS 1.4.2.6 Mechanical systems, WBS 1.4.2.5 Water systems, and WBS 1.4.2.6 DTL Vacuum systems will be install.

Further detail will be provided during title II (final design) phase of this project.

# 4.0 System Interfaces

The DTL has interfaces with a number of other WBS elements as listed below.  The location of the interface, details of the hardware that forms the physical interface, and the parameters that define the function interface, between the DTL and these respective systems are outlined below and will be completely defined in separate interface control documents developed during detail design.

## 4.1 Interfaces within Linac Systems

### 4.1.1	RF Power

a.  RF Power

Function:	402.5 MHz power is provided by the RF system to accelerate the beam.

Physical:       Half height WR2100 waveguide flange below the 402.5 MHz Rf window including an interface

flange for an arc detector located on the iris waveguide.

b.  RF Controls

Function:	The DTL monitor loop provides cavity field information.

Physical:       At the N-type coax connector provided on the DTL.

### 4.1.2	Coupled-Cavity Linac

Function:	Continuation of accelerating structure past the DTL.

Physical:       The beam flange past the DTL-supplied gate valve with a flexible bellows connection for mechanical compliance.

### 4.1.3	Diagnostics

Function:	Diagnostics are incorporated throughout the DTL beamline to provide beam physics information.

Physical:       See appendix B for general location of beam diagnostics.

## 4.2 Interfaces With Other SNS Systems

### 4.2.1 Front-End Systems

Function:	Medium-energy beam transport between the radio-frequency quadrupole and the DTL.

Physical:       See drawing 155Y504100, MEBT/DTL interface control drawing.

### 4.2.2 Conventional Facilities

Functional interfaces to Conventional Facilities (WBS 1.8) are identified in the Systems Requirements Document for the Title I Design of the Front End Building, the Linac Tunnel, and the Klystron Building, SNS10830000SR0001.

### 4.2.3 Control Systems

See the vacuum system design criteria document, SNS-104020400-TR0002-R01, for the vacuum system control requirements.

See the water system design criteria document, SNS-104020400-TR0001-R01, for the water system control requirements. See the control systems design criteria document, document number TBD, for the general control system requirements..0	QUALITY ASSURANCE;

# 5.0 Quality Assurance555555.5555555555555555

## 5.1 General

All activities associated with the DTL will be conducted in accordance with the requirements contained in SNS-QA-P01, SNS QA Plan; and accepted engineering standards and practices.  Good engineering practices within the established design process will ensure future safe and reliable operation of the SNS, and mitigate conditions that could pose a threat to success.  In all cases, consensus standards will be used to accomplish design activities.  Where existing standards do not adequately control an activity, appropriate administrative controls will be considered.

Thoughtfully derived and properly “graded” controls for design, manufacture, installation, testing, and operation of the remote handling systems will be established prior to execution of each of these activities.  Additional controls will only be proposed in cases where they enhance success.

Successful QA program performance will be verified through validation activities such as design reviews, surveillance activities, inspections, tests, and readiness reviews.

## 5. 2 Tests

### 5.2.1	Engineering Tests

Tests used primarily for the purpose of acquiring data to support the design and development shall be specified in the individual development and product specifications.

## 5.3 Quality Conformance Inspections

DTL tank design and performance requirements identified in this specification and referenced specifications shall be verified by inspection, analysis, demonstration, similarity, test, or a combination of these.

### 5.3.1	Certifications

A copy of inspection and test records and certifications relating to the requirements associated with the following parameters, as a minimum shall be prepared:

- PMQ magnet mapping

- Endwall alignment mapping

- Drift tube alignment measurements

- Final electric field distribution and tilt sensitivity tests

- Frequencies measurements

- Vacuum leak checks

- Hydrostatic leak checks

- Inspection reports on fabricated components

# 6.0  Safety Requirements

The DTL shall meet all applicable Department of Energy and other Federal safety regulations, plus those applicable State, Local, and SNS safety requirements, and the SNS Environment, Safety, and Health (ES&H) Plan. A hazard analysis shall be conducted in conjunction with the design of the DTL tanks. Particular emphasis shall be placed on restricting the use of hazardous materials where possible; grounding of electrical systems; RF radiation emission; cleanliness. RF leakage shall conform to the requirements of ANSI C95.1-1982.

# 7.0 Documentation

Technical documentation shall be delivered, sufficient in level of detail, type and quality to allow re-procurement and satisfy the SNS documentation standards and practices. Electronic, PDF files shall be provided. The technical documentation package shall include the following.

## 7.1 Specifications

The following specifications shall be developed:

a. Fabrication Specification establishing the requirements for manufacture and acceptance of the DTL tanks and components.

b.	Process Specifications defining specific process involved in manufacturing techniques.

c.	Material specifications defining raw materials and semi-fabricated material (e.g. electrical cable, piping)

d.	Interface Control Documents specifying the interfaces for beam input from the MEBT.

## 7.2 Engineering Drawings and Associated Lists

Drawings and associated lists shall be provided. SNS Drawing Numbering System, and Global Drafting Manual, supplemented by procedures and instructions pertinent to SNS only, shall be used as guides.

## 7.3 Technical Manuals and Procedures

### 7.3.1	Procedures

Procedures shall be provided for, at minimum,

a) initial installation and set-up of equipment

b) normal operation of equipment

c) normal and/or preventative maintenance

d) a troubleshooting guide for any anticipated potential malfunctions

### 7.3.2	Manuals

Manuals shall be provided for, at minimum, all purchased components and subsystems delivered to SNS.

### 7.4.4 	Documentation Numbering

The SNS document numbering procedure shall be followed.

# 8.0 Operations and Maintenance Requirements

## 8.1 Operating Modes

The DTL normal operating mode is with beam on. The water system has multiple operating modes, see the DTL and Coupled cavity water system design criteria document, SNS-104020400-TR0001-R01.

## 8.2 Surveillance and in-service inspection requirements

No design requirements have been identified.

## 8.3 General Maintenance

No design requirements have been identified.

## 8.4 Abnormal Events and Recovery Procedures

No design requirements have been identified.

# 9.0 Installation Requirements

TBD

# 10.0 Acceptance Testing Requirements

TBD

# 11.0 Commissioning Requirements

TBD

# 12.0 Codes and Standards

The DTL systems shall be designed in accordance with the codes and standards of record as evoked by the Systems Requirements Document for the SNS Linear Accelerator and the following:

## 12.1 SNS documents

SNS Environment, Safety, and Health (ES&H) Plan,” (Reference Doc No: 102030000ES0001R00, 12-1-99)

SNS-QA-P01, SNS QA Plan; Quality Assurance Plan for the Target System of the Spallation Neutron Source Project, Title I (Preliminary Engineering Design).

SNS Systems Requirements Document for Equipment, Device and Signal Naming," SNS 102000000 SR0001-R00

System Requirements Document for Title I Design of the Front-end Building, the Linac Tunnel and the Klystron Building,” SNS 108030000SR000

## 12.2 Non-SNS documents and standards

American Welding Society AWS D1.1-2000, Structural Welding Code - Steel

American Society of Mechanical Engineers, ASME NOG-1

490 Specification Practices

Global Drafting Standard

Dimensioning and Tolerancing per ASME Y14.5M-1994

ASME Boiler and Pressure Vessel Code. Section VIII

ANSI C95.1-1982 American National Safety Standard Levels with respect to Human Exposure to Radio Frequency Electromagnetic Fields (300 kHz-100GHz)

Structural Steel Tubing per ASTM A-500

# Appendix A

## Drift tube cell geometry table

Face angles are in degrees

All other parameters are in inches

Tk = DTL tank number

DT = Drift tube number in respective tank

CellLen = Cell Length

CellLoc = Cell location from start of MEBT

Zstem = center of drift tube from tank upstream endwall

FA_1 = Face angle of upstream drift tube

FA_2 = Face angle of downstream drift tube

Flat_1 = Distance between tangencies of Rin_1 and Rout_1 on upstream drift tube

Flat_2 = Distance between tangencies of Rin_2 and Rout_2 on downstream drift tube

Rin_1 = Drift tube inside bore radius on upstream drift tube

Rin_2 = Drift tube inside bore radius on downstream drift tube

Rout_1 = Radius between flat and face angle on upstream drift tube

Rout_2 = Radius between flat and face angle on downstream drift tube

DT_Len = Drift tube body length

Rb_1 = Inside bore tube radius of upstream drift tube

Rb_2 = Inside bore tube radius of downstream drift tube

Rc_1 = Outside corner radius of upstream drift tube

Rc_2 = Outside corner radius of downstream drift tube

Diam = Outside diameter of drift tube

GapLen = Gap length, distance between drift tube noses

GL = Focusing gradient (tesla)

Figure A-1 Drift tube cell geometry

# Appendix B

## DTL Magnet, Post coupler, and Diagnostics schematic Layout

Table of symbols

# Text boxes

QV

Tank 1

Tank 1 1

Tank 2

7.5 MeV

2.5 MeV

22.8 MeV

Tank 2

Tank 3

Tank 4

Tank 5

Tank 6

86.8 MeV

2.5 MeV

QH

QV

QV

QH

QV

10

DV

QV

QV

30

DV

DV

QH

40

30

QV

QH

QV

20

QV

QH

QV

QH

QH

QH

QV

QH

DH

DH

DH

DH

QH

10

QH

QV

QH

QV

DV

5

Wire Scanner

210

12

Dipole, Vertical

Dipole, Horizontal

DV

DH

6

WS

Half drift tube

QV

QH

FC

1

Quadrupole

5

5

6

10

12

12

75

74

Drift tube

Number of units

Name

Symbol

6

Number of units

Name

Symbol

PMQ, Horizontal

PMQ, Vertical

Gate Valve

Faraday cup (Beam stop)

Energy degrader

Current Monitor (CM)

BPM / Phase detector in DT

DTL Tank

DH

DH

DV

DV

20

10

QV

QV

50

40

30

20

10

QV

QH

QH

QH

QV

QV

QH

QH

QH

QV

QV

QV

QV

QH

QH

QV

QH

QH

QV

QV

QH

QH

QV

QV

QH

QH

QV

QH

QV

QH

QH

QV

QH

QH

QV

QV

QH

QV

QV

QH

QH

QH

QV

QV

QH

QH

QV

QV

QH

QH

QV

QV

QH

Beam

60

FC

WS

Tank 2

0

0

Tank 1

QH

QV

QV

QH

QH

QV

QV

QH

QH

QV

QH

QH

QV

QV

QH

QV

QV

QH

QH

QH

QV

QV

QH

QV

QV

34

Tank 4

WS

FC

Tank 3

0

48

WS

FC

QV

DH

DH

DH

DH

DH

DH

DV

DV

DV

DV

DV

DV

20

10

20

10

20

QV

QV

QH

QH

QV

QV

QH

QH

QV

QV

QH

QH

QV

QV

QH

QH

QV

QH

QH

QV

QV

QH

QH

QV

QV

QH

QH

QV

QV

QH

QH

QH

QH

QV

QV

QH

QH

QV

QV

QH

QH

Seg. 2

FC

Seg. 1

CCL Module # 1

0

0

22

28

WS

FC

WS

FC

Tank 6

24

Tank 5