Specifications for

Testing Coefficient of Thermal Expansion (CTE)

# Section A.  Requirements For Testing Ultra Low Expansion (ULE) Samples

- Temperature Measurement

  - Thermal distortion tests (“distortion” is used here to represent either expansion or contraction, as appropriate) are to be run between a low temperature of no warmer than 30 K (Kelvin) and a high temperature of no warmer than room temperature (nominally 293 K).

  - Calibration data for the temperature sensors used during testing shall be a deliverable item.

  - A minimum of two temperature sensors shall be mounted on each sample.  Vendor shall provide evidence that the temperature sensors and their mounting methods will not adversely affect the thermal distortion behavior of the samples.

  - Vendor shall assure no temperature gradients greater than 1.0 K exist in the test samples when distortion data are recorded.

  - Vendor shall be able to measure sample temperature with an accuracy of +/- 0.5 K or better.  Vendor shall state the accuracy to which sample temperature can be measured.

  - If other materials are used as a comparative reference during testing, temperature of these references shall be recorded whenever distortion and sample temperature data are recorded.

- Room Temperature Length Measurement

  - Vendor shall be able to measure sample length (in measurement direction, as shown in Figure 2) at room temperature with an accuracy of +/- 0.0001 inch or better.  Vendor shall state the accuracy to which sample length can be measured at room temperature.

- Distortion Measurement

  - Thermal distortion shall be measured in the 1.0-inch direction of the 1.0 x 1.0 x 0.175 inch samples (see Figure 2).

  - The distortion of each sample shall be measured over the temperature range given in 1a above with an accuracy goal of +/- 5x10-6 inch or better.

  - Vendor shall state, in microinches, to what accuracy the distortion can be measured of a ULE sample which is 1.0 inch long in the measurement direction over the temperature range given in 1a above using the proposed equipment and methods.

  - Vendor shall state whether the accuracy of distortion measurements varies with sample length and, if so, quantify the accuracy versus sample length.

  - Vendor shall state whether the accuracy of distortion measurements varies with temperature and, if so, quantify the accuracy versus temperature.

- Data Reporting

  - Vendor shall record the room temperature length of each sample before and after it is cryo cycled.

  - Vendor shall record temperature and distortion data no less often than at every 5 K temperature change between the lowest temperature reached and 100 K, and no less often than at every 10 K  change between 100 K and room temperature.

  - As a minimum, temperature and distortion data shall be taken either during the cool-down phase or during the warm-up phase. If the vendor chooses, temperature and distortion data may be recorded during both the cool-down phase and the warm-up phase.

  - Vendor shall deliver a final report containing descriptions and illustrations of test setup as appropriate, test procedures, results, and conclusions.

- Other

  - Vendor shall provide any available data on tests which have been performed on calibrated CTE standards (NIST, NBS, etc.) using the same test apparatus and test procedures used for this testing.

  - Vendor shall provide any available data on tolerances and error contributions of the specific test apparatus and procedures used for this testing.

  - To assure repeatability in the test apparatus and procedures, the first sample or set of samples shall be tested completely, removed from the test apparatus, reinstalled, and tested completely again.  By comparing the results the vendor shall show the level of repeatability which can be achieved with the test apparatus.

  - Alternate approaches or requirement levels may be proposed and will be assessed in light of the evaluation criteria and requirements listed here.

  - Shall use applicable ASTM Standard Test Method for Linear Thermal Expansion

# Deliverable Items

- Temperature sensor calibration data

- Analysis of temperature sensor effects on test sample

- Statement of temperature measurement accuracy as a function of temperature between room temperature and 30 K

- Statement of room temperature sample length measurement accuracy

- Statement of distortion measurement accuracy for a 1.0-inch ULE sample, in microinches, for a temperature change from room temperature to 30 K.

- Statement of how distortion measurement accuracy varies with sample length, if any

- Statement of how distortion measurement accuracy varies with sample temperature, if any

- Pre- and post-cryo cycle room temperature length of each specimen

- Distortion vs. temperature data for each test of each specimen, including any temperature data for comparative references if they are used

- Any available calibration of the test apparatus against standards (NIST, NBS, etc.)

- Any available tolerance and error contributions of test apparatus and procedures

- Final report: ULE to be tested first (before Be) due to sample availability.  Final ULE report (summary of tabulated data, charts, graphs, textual summary with findings and conclusion) due two weeks following completion of final test run.  Additionally, a weekly preliminary report (tabulated data sufficient) is due to MSFC by Tuesday each week for runs completed by the previous Friday (email, fax, etc.).

- Return of all samples

| ULE CTE Testing Evaluation Criteria | Points |
| --- | --- |
| 1) Temperature Measurement |  |
| Low temperature no warmer than 30 K (colder is better) | 2 |
| Provide calibration data for temperature sensors | 5 |
| Minimum two temperature sensors per sample (more is better if there is no effect on sample behavior) | 5 |
| Analysis/evidence that temperature sensors do not affect thermal distortion behavior of samples | 7 |
| No temperature gradients in samples greater than 1.0 K (smaller is better) | 8 |
| Measure sample temperature with accuracy of +/- 0.5 K or better (smaller is better) | 8 |
| 2) Room Temperature Length Measurement |  |
| Measure sample at room temperature with accuracy of +/- 0.0001 inch or better (smaller is better) | 2 |
| 3) Distortion Measurement |  |
| Measure distortion with accuracy goal of +/- 5x10-6 inch or better (smaller is better) | 30 |
| Quantify distortion measurement accuracy versus sample length variation | 8 |
| Quantify distortion measurement accuracy versus temperature variation | 8 |
| 4) Data Reporting |  |
| Record temperature/distortion data both warm-to-cold and cold-to-warm (optional) | 4 |
| Record temp/distortion at least every 5 K for coldest to 100 K, at least every 10 K for 100 K to room temp (more often is better) | 3 |
| 5) Other |  |
| Test data on calibrated CTE standards | 5 |
| Tolerances/error contributions of test apparatus and procedures | 5 |
| Total | 100 |

# ULE CTE Test Specimen Details - Rev. D, 17 May 02

(1/4 inch thick, with chamfers)

# Section B.  Requirements For Testing Beryllium 0-30H (Be)

- Temperature Measurement

  - Thermal distortion tests (“distortion” is used here to represent either expansion or contraction, as appropriate) are to be run between a low temperature of no warmer than 30 K (Kelvin) and a high temperature of no warmer than room temperature (nominally 293 K).

  - Calibration data for the temperature sensors used during testing shall be a deliverable item.

  - A minimum of two temperature sensors shall be mounted on each sample.  Vendor shall provide evidence that the temperature sensors and their mounting methods will not adversely affect the thermal distortion behavior of the samples.

  - Vendor shall assure no temperature gradients greater than 1.0 K exist in the test samples when distortion data are recorded.

  - Vendor shall be able to measure sample temperature with an accuracy of +/- 0.5 K or better.  Vendor shall state the accuracy to which sample temperature can be measured.

  - If other materials are used as a comparative reference during testing, temperature of these references shall be recorded whenever distortion and sample temperature data are recorded.

- Room Temperature Length Measurement

  - Vendor shall be able to measure sample length (in measurement direction, as shown in Figure 1) at room temperature with an accuracy of +/- 0.0001 inch or better.  Vendor shall state the accuracy to which sample length can be measured at room temperature.

- Distortion Measurement

  - Thermal distortion shall be measured in the 1.5-inch direction of the 1.5 x 1.75 x 0.125 inch samples (see Figure 1).

  - The distortion of each sample shall be measured over the temperature range given in 1a above with an accuracy goal of +/- 5x10-6 inch or better.

  - Vendor shall state, in microinches, to what accuracy the distortion can be measured of a beryllium sample which is 1.5 inch long in the measurement direction over the temperature range given in 1a above using the proposed equipment and methods.

  - Vendor shall state whether the accuracy of distortion measurements varies with sample length and, if so, quantify the accuracy versus sample length.

  - Vendor shall state whether the accuracy of distortion measurements varies with temperature and, if so, quantify the accuracy versus temperature.

- Data Reporting

  - Vendor shall record the room temperature length of each sample before and after it is cryo cycled.

  - Vendor shall record temperature and distortion data no less often than at every 5 K temperature change between the lowest temperature reached and 100 K, and no less often than at every 10 K  change between 100 K and room temperature.

  - As a minimum, temperature and distortion data shall be taken either during the cool-down phase or during the warm-up phase. If the vendor chooses, temperature and distortion data may be recorded during both the cool-down phase and the warm-up phase.

  - Vendor shall deliver a final report containing descriptions and illustrations of test setup as appropriate, test procedures, results, and conclusions.

- Other

  - Vendor shall provide any available data on tests which have been performed on calibrated CTE standards (NIST, NBS, etc.) using the same test apparatus and test procedures used for this testing.

  - Vendor shall provide any available data on tolerances and error contributions of the specific test apparatus and procedures used for this testing.

  - To assure repeatability in the test apparatus and procedures, the first sample or set of samples shall be tested completely, removed from the test apparatus, reinstalled, and tested completely again.  By comparing the results the vendor shall show the level of repeatability which can be achieved with the test apparatus.

  - Alternate approaches or requirement levels may be proposed and will be assessed in light of the evaluation criteria and requirements listed here.

  - Shall use applicable ASTM Standard Test Method for Linear Thermal Expansion

# Deliverable Items

- Temperature sensor calibration data

- Analysis of temperature sensor effects on test sample

- Statement of temperature measurement accuracy as a function of temperature between room temperature and 30 K

- Statement of room temperature sample length measurement accuracy

- Statement of distortion measurement accuracy for a 1.5-inch Be sample, in microinches, for a temperature change from room temperature to 30 K.

- Statement of how distortion measurement accuracy varies with sample length, if any

- Statement of how distortion measurement accuracy varies with sample temperature, if any

- Pre- and post-cryo cycle room temperature length of each specimen

- Distortion vs. temperature data for each test of each specimen, including any temperature data for comparative references if they are used

- Any available calibration of the test apparatus against standards (NIST, NBS, etc.)

- Any available tolerance and error contributions of test apparatus and procedures

- Reports Required: Be testing to follow ULE testing and, as required before, a final report (summary of tabulated data, charts, graphs, textual summary with findings and conclusion)  due to MSFC two weeks following completion of final test run.  As with ULE, a weekly preliminary report (tabulated data sufficient) is due to MSFC by Tuesday each week for runs completed by the previous Friday (email, fax, etc.).

- Return of all samples

| Be CTE Testing Evaluation Criteria | Points |
| --- | --- |
| 1) Temperature Measurement |  |
| Low temperature no warmer than 30 K (colder is better) | 2 |
| Provide calibration data for temperature sensors | 5 |
| Minimum two temperature sensors per sample (more is better if there is no effect on sample behavior) | 5 |
| Analysis/evidence that temperature sensors do not affect thermal distortion behavior of samples | 7 |
| No temperature gradients in samples greater than 1.0 K (smaller is better) | 8 |
| Measure sample temperature with accuracy of +/- 0.5 K or better (smaller is better) | 8 |
| 2) Room Temperature Length Measurement |  |
| Measure sample at room temperature with accuracy of +/- 0.0001 inch or better (smaller is better) | 2 |
| 3) Distortion Measurement |  |
| Measure distortion with accuracy goal of +/- 5x10-6 inch or better (smaller is better) | 30 |
| Quantify distortion measurement accuracy versus sample length variation | 8 |
| Quantify distortion measurement accuracy versus temperature variation | 8 |
| 4) Data Reporting |  |
| Record temperature/distortion data both warm-to-cold and cold-to-warm (optional) | 4 |
| Record temp/distortion at least every 5 K for coldest to 100 K, at least every 10 K for 100 K to room temp (more often is better) | 3 |
| 5) Other |  |
| Test data on calibrated CTE standards | 5 |
| Tolerances/error contributions of test apparatus and procedures | 5 |
| Total | 100 |

### Beryllium CTE Test Specimen Details - Rev. C, 19 April 02

(1/4 inch thick)

# Period of Performance

This procurement shall be referred to as the first-round of CTE testing for Be and ULE.  The period of performance for the first-round of testing shall be from effective date of award through November 29, 2002.

If performance proves acceptable, it is possible that funding might be available for additional CTE testing of the same types of materials or different materials.  This will be referred to as the second-round of testing.  The period of performance for the second-round of testing shall be from the end of the first-round through November 30, 2003.