ARM TR-057

New Microwave Radiometer Profiler (MWRP) Handbook

January 2005

J.C. Liljegren

Work supported by the U.S. Department of Energy,

Office of Science, Office of Biological and Environmental Research

Contents

1.	General Overview	1

2.	Contacts	1

3.	Deployment Locations and History	1

4.	Near-Real-Time Data Plots	1

5.	Data Description and Examples	1

6.	Data Quality	2

7.	Instrument Details	3

# General Overview

The microwave radiometer profiler (MWRP) provides vertical profiles of temperature, humidity, and cloud liquid water content as a function of height or pressure at approximately 5-minute intervals for nearly all weather conditions.  The profiles are derived from measurements of absolute microwave radiances (expressed as “brightness temperatures”) obtained at twelve frequencies in the range of 22-30 GHz and 51-59 GHz.  The data are useful for input to numerical weather forecast models and others that require continuous, high temporal resolution profiles.  The microwave radiance measurements are useful for testing models of microwave radiation transfer used to derive the profiles.

# Contacts

## Mentor

Jim Liljegren

Decision and Information Sciences Division

Argonne National Laboratory, Bldg. 900

Argonne, IL 60439

Phone:  630-252-0540

E-mail:  jcliljegren@pnl.gov

## Instrument Developer

# Deployment Locations and History

# Near-Real-Time Data Plots

# Data Description and Examples

This section is not applicable to this instrument.

## Data File Contents

### Primary Variables and Expected Uncertainty

#### Definition of Uncertainty

### Secondary/Underlying Variables

### Diagnostic Variables

### Data Quality Flags

### Dimension Variables

## Annotated Examples

## User Notes and Known Problems

## Frequently Asked Questions

# Data Quality

## Data Quality Health and Status

## Data Reviews by Instrument Mentor

## Data Assessments by Site Scientist/Data Quality Office

## Value-Added Procedures and Quality Measurement Experiments

# Instrument Details

## Detailed Description

### List of Components

### System Configuration and Measurement Methods

### Specifications

## Theory of Operation

The frequency-dependent microwave radiance, reported as a brightness temperature TB is the fundamental quantity measured by the radiometer. The accuracy of the brightness temperature measurements is critical to obtaining accurate retrievals of atmospheric state parameters. The measurement accuracy depends on the quality and stability of the hardware as well as on the accuracy and precision of the calibration, which can be difficult to separate. In this section the measurement accuracy is evaluated by analysis of the calibration data and through direct and indirect comparisons with other instruments. A detailed investigation of the calibration procedures and algorithms is beyond the scope of this study.

At each frequency the signal voltage Vsky resulting from a sky measurement is transformed to a brightness temperature according to the expression

TB = Tref -G-1 Vref -Vsky

(	)

(1)

where Tref is the (near ambient) temperature of an internal blackbody reference target, Vref is the

†

signal voltage resulting from a measurement of the reference target, and G is the radiometer gain

given by

G = Vhi -Vlo

(	)/ Thi -Tlo

(	)

(2)

where Thi and Vhi are the temperature and signal voltage associated with a high temperature

†

reference and Tlo and Vlo are associated with a low temperature reference. In the case of the

(3)

MWRP, the low temperature reference is provided by the internal blackbody target such that Tlo=Tref. The high temperature reference is achieved by adding a noise injection temperature Tnd from a calibrated noise diode such that Thi = Tref + Tnd. The gain is then calculated as

G = Vref +nd -Vref

(	)/Tnd

(4)

G = Vref -Vcold

(	)/ Tref -Tcold

(	)

where Vref+nd – Vref is the difference in signal voltages recorded when viewing the internal target

†

with and without the noise diode switched on. The value of Tnd is calibrated as follows. First the

gain is determined using

where Tcold and Vcold correspond to the temperature and signal voltage of a cold external reference.

†

Then (3) and (4) are combined to yield

(5)

Ê	ˆ

(	) Vref +nd -Vref

.

Tnd = Tref -Tcold Á	˜

Ë Vref -Vcold ¯

Because the K-band (22-30 GHz) and V-band (51-59 GHz) receivers operate over different

†

ranges of brightness temperature, different methods of providing the cold external reference

temperature are utilized.

K-band channels (22-30 GHz)

The water vapor resonance centered at 22.235 GHz is weakly absorbing so the optical thickness increases linearly with increasing water vapor amount. Consequently, the tipping curve method is used to provide the cold reference temperatures in the 10-90 K range necessary to calibrate the K-band channels. The tipping curve method exploits the linear relationship between optical thickness and atmospheric path length that holds when the sky is free of liquid water clouds and strong horizontal water vapor gradients to derive the zenith brightness temperature, which serves as Tcold in (5), from which Tnd is determined.

The MWRP acquires tipping curve measurements continuously, at about 15-minute intervals. These are screened for valid sky conditions based on the linearity of the regression of optical thickness on path length (“air mass”) by requiring that the correlation coefficient exceed 0.99. Since the MWRP was redeployed to the SGP in June 2001, over 26,000 valid tipping curves have been acquired. A statistical summary of the values of Tnd derived from these tipping curves is graphically depicted in Figure 5. The filled circles indicate the monthly means for 2001 (black) and 2002 (grey); the error bars indicate twice the standard deviation, which represents about 96% of the range of values.

The statistics depicted in Figure 5 address the issue of the long-term stability of the Kband receiver and the calibration. The inter-annual variation in the monthly means is generally very small (<1%) and in no case exceeds the spread for the month. It is also apparent that the values of Tnd vary seasonally, which reflects a weak dependence on temperature. To account for this, temperature correction coefficients have been determined for each channel based on a linear fit of Tnd to Tref. The corrections are used to adjust the individual values of Tnd to a nominal reference temperature (290 K). The adjusted values are low-pass filtered: yi = axi + (1–a)yi–1 where a = 0.1. The most recent low-pass filtered values of Tnd are then adjusted to the current value of Tref prior to being used in (3) to calculate the gain.

\.

Figure 5.	Monthly mean noise injection temperatures Tnd derived from tipping curves for 2001 (black circles) and 2002 (grey circles). The error bars indicate twice the standard deviation or 96% of the range of values. Individual values of Tnd derived from liquid nitrogen calibrations are indicated by crosses.

The small crosses on days 122, 179, and 340 in Figure 5 indicate values of Tnd determined using a liquid nitrogen-filled external target to provide the cold temperature reference in (5). For the most part these values are consistent with the results of the tipping curve calibrations. However, the results for the 23.035 GHz channel on day 179 in 2001 are more than two standard deviations away from the mean of the tipping curve-derived values for that channel; the results for the 22.235 and 23.835 GHz channels on this day are more than one standard deviation from the mean. This upward bias could have resulted from condensation forming on the bottom of the liquid nitrogen target, which would have caused the effective target temperature to be greater than that used to compute Tnd; however, no such errors were observed in the calibration data for the V-band (temperature sensing) channels. Also, the spread of the liquid nitrogen-derived results is relatively large, comparable to the monthly standard deviation of the tipping curvederived results in most cases. Although this suggests that the liquid nitrogen-filled target is inferior to the tipping curve procedure for providing a cold temperature calibration reference for the K-band channels, the number of liquid nitrogen-derived values is insufficient to draw a firm conclusion in this regard.

To assess the accuracy of the brightness temperatures measured by the MWRP, the measurements at 23.835 and 30.0 GHz were compared with brightness temperatures measured by a two-channel microwave radiometer (MWR) at 23.8 and 31.4 GHz. The calibration of the MWR is described by Liljegren (2000). The differences between the MWRP and MWR brightness temperatures are presented as functions of the MWR brightness temperature in Figures 6 and 7. Because these comparisons are not at exactly the same frequencies, the predicted differences based on the Rosenkranz (1998) absorption models for oxygen and water vapor are also indicated. The statistics are summarized in Table 4.

Table 4	.	Brightness temperature difference ~TB statistics for the comparison of MWRP and

MWR during 2000-2002 under liquid-water-free sky conditions.

|  | 2000 | 2001 | 2002 | Model† |
| --- | --- | --- | --- | --- |
| Number of samples | 8650 | 7202 | 13,711 |  |
| 23.835 – 23.8 GHz |  |  |  |  |
| Slope of ~TB vs TB regression, K/K | 0.005 | 0.036 | 0.030 | -0.009 |
| Regression root-mean-square error, K | 1.03 | 0.70 | 0.67 |  |
| Mean ~TB (bias), K | 0.14 | 0.59 | 0.48 |  |
| Standard deviation ~TB, K | 1.03 | 0.93 | 0.82 |  |
| 30.0 – 31.4 GHz |  |  |  |  |
| Slope of ~TB vs TB regression, K/K | -0.016 | 0.089 | 0.070 | 0.048 |
| Regression root-mean-square error, K | 0.79 | 0.51 | 0.51 |  |
| Mean ~TB (bias), K | 0.08 | 0.43 | 0.32 |  |
| Standard deviation ~TB, K | 0.79 | 0.73 | 0.66 |  |

†Rosenkranz, 1998.

Figure 6.	Differences in measured brightness temperature between the MWRP at 23.835 GHz and the collocated two-channel MWR at 23.8 GHz at the SGP. The heavy black line indicates the differences between model-calculated brightness temperatures at these frequencies.

Figure 7.	Differences in measured brightness temperature between the MWRP at 30.0 GHz and the collocated two-channel MWR at 31.4 GHz at the SGP. The heavy black line indicates the differences between model-calculated brightness temperatures at these frequencies.

The results do not appear to be entirely consistent over all three years: for both the 23.8 and 31.4 GHz comparisons, the mean difference is near zero over the range of brightness temperatures in 2000 whereas for 2001 and 2002 the mean difference is near zero for the lower end of the brightness temperature range but shows increasing bias (MWRP higher) with increasing brightness temperature. The latter behavior appears to agree with the model-predicted trend at 31.4 GHz; however, the correlation with the trend at 23.8 GHz suggests this may also be a coincidence. The difference in behavior between 2000 and 2001-2002 may be due to modifications made to the radiometer by Radiometrics after it was returned from Barrow, Alaska in the spring of 2001. The explanation of the observed trends in the measured brightness temperature differences between the MWRP and MWR is not clear and will require a detailed analysis of the tipping curve reduction algorithms implemented in the MWRP to determine. Despite this, the mean brightness temperature difference is generally less than the claimed 0.5 K accuracy of the MWRP and the regression root-mean-square (due to variations in both the MWR and MWRP) is only slightly larger than #2 times the claimed repeatability of 0.25 K.

To assess the accuracy of all five K-band channels, comparisons were carried out with model-calculated brightness temperatures. Of course, such a comparison depends on the accuracy of the model as well as the accuracy of the profiles of pressure, temperature, and water vapor used to drive the model. The Rosenkranz (1998) absorption models were used in the brightness temperature calculations. To overcome the known dry bias of the radiosonde relative humidity sensors as well as to reduce the variability in their calibration, the ARM linearly scaled sounding product (sgplssondeC1.c1) was used to drive the model. In this product the measured radiosonde relative humidity is linearly scaled by the ratio of the integrated water vapor amount from a collocated two-channel MWR to that of the original sounding. Ratios of measured-to modeled brightness temperatures are presented in Figure 8 for the 2001-2002 deployment at the SGP. The blue circles indicate the case where the collision-broadened width parameter gC of the 22.235 GHZ water vapor line used in the model calculations was the original value supplied by Rosenkranz; the red circles indicate the case where the width was reduced to 95% of the original value (per conversation with S. A. Clough). Statistics corresponding to the results in Figure 8 are presented in Table 5.

Although the trend of increasing ratio with increasing brightness temperature reflects the earlier trend of the MWRP–MWR comparison, nevertheless the results clearly show that reducing the line width substantially improves the agreement between the measurements and the model. Not only is this an important result in its own right, but it also demonstrates the consistency of the measured brightness temperatures across the five K-band channels.

The scatter in these comparisons reflected in the standard deviation of the ratios is not affected by the reduction in line width. This scatter results from variations in the sounding data and the MWR data used to scale the soundings as well as the MWRP measurements. This suggests that the accuracy of the MWRP K-band channels is better than 2.5-3.0%, consistent with the MWRP–MWR comparisons summarized in Table 4.

Figure 8.	Ratios of measured and modeled brightness temperatures at the K-band frequencies for both the original value of the width parameter (blue) provided with the Rosenkranz (1998) water vapor absorption model and 95% of the original value (red).

Table 5. 	Statistics of the ratio TB, MWRP/TB, Model at the five K-band frequencies using the original and reduced collision-broadened line width parameter gC for 719 clear-sky comparisons at the SGP during 2001-2002.

|  | Mean |  | Std. Dev. |
| --- | --- | --- | --- |
| Frequency, GHz | 1.0 gC | 0.95 gC | 1.0 gC | 0.95 gC |
| 22.235 | 1.040 | 1.003 | 0.026 | 0.025 |
| 23.035 | 1.026 | 0.998 | 0.029 | 0.028 |
| 23.835 | 1.013 | 1.001 | 0.026 | 0.025 |
| 26.235 | 0.984 | 0.996 | 0.026 | 0.026 |
| 30.000 | 0.997 | 1.009 | 0.032 | 0.032 |

V-band channels (51-59 GHz)

The strongly absorbing oxygen resonances in the 51-59 GHz range cause the relationship between optical thickness and absorber amount or atmospheric path length to be non-linear, which precludes the use of the tipping curve method for calibration. Instead, because the lowest brightness temperature measured at the least absorbing frequency (51.25 GHz) is about 100 K, a liquid nitrogen-filled external target is used to provide the ~77 K cold temperature reference needed to determine Tnd in (5). As shown in Figure 9, the external target is comprised of a very low absorption thick-walled Styrofoam™ container with a blackbody target placed inside on the bottom and filled with 25-30 liters of LN2. Alternating measurements of the near ambient internal and cold external targets are performed for up to an hour, or until condensation forms on the bottom of the Styrofoam container. The heated blower normally used to prevent dew from condensing on the polycarbonate foam window of the radiometer is manually cycled on and off to delay the formation of condensation on the bottom of the container.

The brightness temperature measured by the MWRP when viewing the LN2 target is elevated ~2K above the boiling point of LN2, which depends on atmospheric pressure, due to a number of effects, principally reflection at the interface between the Styrofoam and LN2 (~1.7 K), the absorption of the Styrofoam (~0.2K), and the elevation of the boiling point of LN2 due to the hydrostatic pressure of the 20 cm column of LN2 above the bottom of the container (~0.2K) as discussed by Solheim (1993). Corrections for these effects are included in the calibration process. No attempt has been made in this evaluation to verify these corrections.

Since the V-band receiver was upgraded in the spring of 2001, four LN2 calibrations have been performed. The results of these calibrations are summarized in Table 6. The standard deviations of the LN2-derived values of Tnd over the 11-month period shown have generally been less than 1% of the mean, except for the 52.85 GHz channel, which was slightly above 1%. This demonstrates good calibration stability.

Figure 9.	The Microwave Radiometer Profiler (MWRP) at Barrow, Alaska with the LN2-filled Styrofoam target in place during calibration of the V-band channels in September 2000.

Table 6. 	Statistics of temperature-corrected Tnd at the seven V-band measurement frequencies

for the LN2 calibrations performed at the Radiometrics facility in Boulder, CO and at the SGP
during 2001-2002.

| Location | Date | 51.25 GHz | 52.28 GHz | 53.85 GHz | 54.94 GHz | 56.66 GHz | 57.29 GHz | 58.80 GHz |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Boulder | Jun 2001 | 242.65 | 231.05 | 210.10 | 194.26 | 196.76 | 202.81 | 217.23 |
| SGP | Jun 2001 | 243.68 | 229.87 | 209.42 | 194.67 | 196.79 | 203.44 | 216.76 |
| SGP | Dec 2001 | 244.29 | 235.66 | 212.91 | 196.58 | 196.87 | 204.18 | 217.60 |
| SGP | May 2002 | 243.83 | 234.87 | 210.52 | 194.71 | 197.61 | 205.24 | 218.47 |
| Mean |  | 243.61 | 232.86 | 210.74 | 195.05 | 197.01 | 203.92 | 217.51 |

| Standard Deviation | 0.69 | 2.83 | 1.52 | 1.04 | 0.41 | 1.04 | 0.73 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Std Dev/Mean (%) | 0.29 | 1.22 | 0.72 | 0.53 | 0.21 | 0.51 | 0.33 |

Because a direct comparison of measured brightness temperatures with another instrument was not possible, a comparison with model-calculated brightness temperatures was performed. Again the Rosenkranz (1998) absorption models were used with input profiles of pressure, temperature, and water vapor supplied by radiosondes. MWRP–Model brightness temperature differences are presented in Figure 10 as a function of the day of the year. The colors indicate the calibration in effect at the time of the measurement, including the deployment at Barrow, Alaska. Summary statistics are provided in Table 7.

Several features are evident in the comparison results. First, the measured TB values generally agree 1-2 K with the model-calculated values. These biases could be due to shortcomings in the model or may reflect the limit of the liquid nitrogen calibration procedure. Also, the results are generally very consistent over time. The exceptions are at 51.25 GHz for the SGP 2000 calibration and at 52.28 GHz for the SGP 2001A calibration, where substantial drift in the calibration is evident. In both of these instances, it appears that the drift did not continue after the instrument was recalibrated. Radiometrics recommends monthly calibrations; however, this may not be practical at remote locations. The data indicate that semi-annual calibrations are sufficient, but regular calibration checks by comparison against model-calculated values would be helpful to detect drifts that may require a recalibration to be performed.

34

Figure 10.	Brightness temperature differences (measured – modeled) for the V-band channels under clear-sky conditions during 2000-2002. The colors indicate the calibration in effect at the time.

Table 7.	Brightness temperature difference ~TB statistics (MWRP–Model) at the seven V-band

frequencies for 719 clear-sky comparisons at the SGP during 2000-2002.

|  |  | ~TB, K | TB, K |  | ~TB/TB |
| --- | --- | --- | --- | --- | --- |
| Frequency, GHz | Mean | Std. Dev. | Mean | Std. Dev. | Std. Dev. |
| 51.25 | 1.11 | 1.72 | 111 | 6.70 | 0.015 |
| 52.28 | -0.86 | 1.36 | 153 | 6.12 | 0.009 |
| 53.85 | 0.62 | 0.65 | 257 | 5.88 | 0.003 |
| 54.95 | -0.07 | 1.18 | 287 | 5.68 | 0.004 |
| 56.66 | 0.77 | 0.45 | 294 | 6.47 | 0.002 |
| 57.29 | 1.06 | 0.47 | 294 | 6.51 | 0.002 |
| 58.80 | 0.86 | 0.79 | 295 | 6.67 | 0.003 |

Although the standard deviation of the differences is generally small – mostly less than 1 K and considerably less than 1% of the mean – the 51.25 GHz channel appears to be noisy relative to the other channels, and the 54.94 and 58.80 GHz channels appear to have become considerably noisier in July 2002. This aspect of radiometer performance is quantified by the noise figure F in decibels, defined by Solheim (1993) as

Ê Tr + T0  ˆ

F =10log Á	˜ .	(6)

Ë T0 ¯

T0 is a reference temperature set to 290 K by convention. Tr is the effective noise temperature of

†

the receiver such that the brightness temperature TB corresponding to a signal voltage V measured by the radiometer is given by

TB = G-1 V -V0

(	)

(7)

and V0 is the voltage offset due to the receiver temperature (i.e., Tr = G-1V0). †

Substituting (2) for the gain into (7), Tr can be calculated as

Thi -Y Tlo

Tr =	,	(8)

Y -1

where Y = Vhi/Vlo. In the case of the MWRP for Tlo = Tref and Thi = Tref + Tnd,

Tnd

Tr =	-Tref

(9)

Y -1

and Y = Vref+nd / Vref.

†

Weekly averaged values of F are presented in Figure 11 for the 7 V-band channels during the SGP deployment in 2001-2002. The error bars indicate twice the standard deviation or 96% of the range of values. The dashed line indicates the value of F (5.55 dB) corresponding to Tr = 750K, a typical value for a radiometer of this design. It is clearly evident that the noise figure significantly exceeds 5.55 dB for the 51.25 GHz channel whereas all of others have noise figures at or below 5.55 dB until week 29 (mid-July) of 2002. The relatively greater scatter in the MWRP–Model comparison at 51.25 GHz in Figure 10 and Table 7 reflects the elevated noise figure for this channel. At week 29 the average value of F jumps above 5.55 dB at 54.94 GHz, and the spread of the weekly averages increases markedly at 51.25, 54.49, and 58.80 GHz. Radiometrics diagnosed this as a failing synthesizer.

Infrared Thermometer (IRT)

The accuracy of the IRT measurements was evaluated by comparison with a similar instrument mounted on the collocated MWR at the SGP, and by comparison with the Atmospherically Emitted Radiance Interferometer (AERI) spectrometer at both the SGP and NSA. The comparison with AERI was accomplished by first integrating the AERI spectral radiance L(n) over the IRT pass band f(n)

u2

L = L(u) f (u)du Ú

(10)

v1

then integrating the Planck radiance B(n, T) over this pass band

u2

L* = B(u,T) f (u)du Ú

v1

and iteratively adjusting the temperature until L* = L. The results of these comparisons are

†

presented in Figure 12 and summarized in Table 8.

Figure 11.	Weekly average values of noise figure for the V-band channels during 2001-2002 at the SGP. The error bars indicate twice the weekly standard deviation. The dotted line corresponds to an effective receiver noise temperature of 750 K.

Figure 12.	Comparison of IRT on MWRP with (a) IRT on MWR at SGP for 2002-2001; (b) AERI at SGP during June 2001; (c) AERI at NSA during February 2001.

Table 8.	Statistical summary of IR temperature comparison (MWRP – MWR) for 2001-2002 at
SGP for all sky conditions.

Number of comparisons 19,367 Mean difference (bias) -0.11

K Median difference -0.56 K Standard deviation 4.5 K

The agreement between the IRT on the MWRP and the IRT on the MWR shows negligible bias. However, the large scatter and corresponding large standard deviation is primarily due to the differences in field of view. The IRT on the MWRP has a 30° field of view whereas the field of view of the IRT on the MWR is less than 3°. Although the larger field of view permits greater sensitivity, it also causes the IRT to respond to clouds that are not within the field of view of the MWRP (Table 1). The agreement with AERI is also good, but shows increasing bias at low sky temperatures. This is consistent with the expected decrease in accuracy as the difference between the IRT housing temperature and the target temperature increases (Table 1).

Surface Meteorological Sensors

The comparison of the barometric pressure, temperature, and relative humidity sensors in the MWRP to the surface values of the radiosondes is summarized in Table 9.

Table 9.	Statistical summary of surface meteorological sensor comparison (MWRP – sonde) for
719 soundings at SGP, 2001-2002.

|  | ~P, mb | ~T, K | ~RH, % |
| --- | --- | --- | --- |
| Mean (bias) | -0.14 | 0.12 | -0.48 |
| Standard deviation | 0.66 | 0.71 | 3.5 |

These results are consistent with the combined uncertainty in the MWRP sensors and the radiosonde sensors.

## Calibration

### Theory

### Procedures

### History

## Operation and Maintenance

### User Manual

### Routine and Corrective Maintenance Documentation

### Software Documentation

### Additional Documentation

## Glossary

See the ARM Glossary.

## Acronyms

See the ARM Acronyms and Abbreviations.

## Citable References

# Comments

I listed the ARM Glossary even though it isn’t on the web version.

I listed the ARM Acronyms even though it isn’t on the web version.