Table 1.  Summary of 2002 accuracy data for QC solutions.

NOMINAL	ACTUAL		        NET MEAN

ANALYTE		UNITS		VALUE		VALUES		n	BIAS %

CLPH			SU		4.70		4.70 ( 0.02	96	 0.07

EQPH			SU		4.70		4.70 ( 0.01	216	 0.07

ANC			µeq/L		40		39.6 ( 1.91	96	-1.02

DIC			mg/L C		5.00		5.05 ± 0.16	107	 0.98

DOC			mg/L C		5.00		5.00 ± 0.14	226	 0.00

COLOR			PCU		10.0		10.08 ( 0.51	196	 0.77

COND (Jan to Mar)		µS/cm		10		9.69 ( 0.18 	13	-3.15

COND (Apr to Dec)		µS/cm		23.8		24.03 ± 0.55	110	 0.97

Ca			mg/L		1.00	 	0.98 ( 0.05	272	-2.00

Mg			mg/L		1.00		1.00 ( 0.05	272	-0.18

Na			mg/L		1.00		1.00 ( 0.05	272	 0.25

K			mg/L		1.00		0.99 ( 0.05	272	-1.06

Cl			µeq/L		50.0		48.5 ( 1.91	214	-2.97

NO3			µeq/L		25.0		24.72 ( 0.80	214	-1.13

SO4			µeq/L		50.0		48.7 ( 1.76	214	-2.54

Total P			µg/L		57.2		56.9 ( 3.25	110	-0.47

Total N			mg/L		0.559		0.546 ± 0.01	117	-2.38

Si			mg/L		2.09		2.09 ± 0.06	134	 0.20

NH4			mg/L		0.54		0.55 ± 0.02	111	 2.37

Al			µg/L 		200.0		199.1 ( 8.54	278	-0.44

Table 2. Summary of data for field blanks.

| Variable | Median | Range | QA Objective | n |
| --- | --- | --- | --- | --- |
| PH, equilibrated | NA | NA | NA | NA |
| PH, ARAS | NA | NA | NA | NA |
| ANC (µeq/L) | 0 | -13 to 6 | 10 | 11 |
| Conductivity (µS/cm) | 1.71 | 0.58 to 2.3 | 4 | 11 |
| DOC (mg/L) | 0.38 | 0.38 to 0.38 | 0.2 | 1 |
| DIC (mg/L) | NA | NA | NA | NA |
| Al, total dissolved (µg/L) | 1.40 | 0 to 3.4 | 10 | 11 |
| Ca (mg/L) | 0.02 | 0 to 0.03 | 0.02 | 11 |
| Mg (mg/L) | 0.00 | 0 to 0 | 0.02 | 11 |
| Na (mg/L) | 0.01 | 0 to 0.05 | 0.02 | 11 |
| K (mg/L) | 0.02 | 0.01 to 0.05 | 0.02 | 11 |
| NH4 (mg/L) | 0.01 | 0 to 0.03 | 0.05 | 11 |
| Cl (µeq/L) | 1.6 | 0.4 to 3.2 | 1.0 | 11 |
| NO3 (µeq/L) | 0 | 0 to 0.2 | 0.2 | 11 |
| SO4 (µeq/L) | 0.7 | 0 to 1.4 | 2.0 | 11 |
| Total P | NA | NA | NA | NA |
| Total N (mg/L) | 0.009 | 0 to 0.024 | 0.050 | 11 |
| Si (mg/L) | 0 | 0 to 0.03 | 0.1 | 11 |
| Apparent Color (PCU) | 0 | 0 to 4 | 0 | 11 |
| Chla | NA | NA | NA | NA |

Table 3.  Itemization of sample blank QC violations defined as two times the MDL.

| Date | Ca (mg/L) | K (mg/L) | Na (mg/L) | Cl (µeq/L) | NO3 (µeq/L) | Al Org (µg/L) | DOC (mg/L) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1/9/02 |  |  |  | 2.5 |  |  |  |
| 1/24/02 |  |  |  | 1.4 |  |  |  |
| 2/6/02 |  |  |  | 1.4 |  |  | 0.33 |
| 2/13/02 |  | 0.03 |  | 1.5 |  |  | 0.30 |
| 2/20/02 |  |  |  | 1.4 |  |  |  |
| 2/28/02 |  |  |  | 1.3 |  |  | 0.28 |
| 3/5/02 |  | 0.03 |  | 1.3 |  |  |  |
| 3/13/02 |  | 0.03 |  | 1.1 |  |  |  |
| 3/17/02 |  |  |  | 1.3 |  |  |  |
| 3/20/02 |  | 0.03 |  | 1.8 |  |  |  |
| 3/27/02 |  | 0.03 |  | 1.5 |  |  | 0.25 |
| 4/5/02 |  |  |  |  |  |  | 0.32 |
| 4/17/02 |  | 0.03 |  | 1.6 |  |  | 0.25 |
| 5/14/02 |  | 0.03 |  |  |  |  |  |
| 5/19/02 |  |  |  |  | 0.2 |  | 0.22 |
| 5/22/02 |  | 0.03 |  |  |  |  |  |
| 5/29/02 |  | 0.03 |  |  | 0.2 |  |  |
| 6/6/02 |  | 0.03 |  |  |  |  |  |
| 6/11/02 |  | 0.03 |  |  |  |  | 0.33 |
| 7/6/02 |  |  |  | 1.5 |  |  | 0.29 |
| 7/24/02 |  |  |  | 1.2 |  |  |  |
| 7/30/02 |  |  |  | 1.7 |  |  |  |
| 8/1/02 |  |  |  | 1.7 |  |  |  |
| 8/6/02 |  |  |  | 1.2 |  |  |  |
| 8/8/02 |  |  |  | 1.1 |  |  |  |
| 8/14/02 | 0.03 | 0.03 |  | 1.2 |  |  |  |
| 8/20/02 | 0.03 |  |  | 1.7 |  |  | 0.23 |
| 8/27/02 | 0.03 |  |  | 1.2 |  |  | 0.25 |
| 8/28/02 |  |  |  |  |  |  | 0.32 |
| 9/2/02 | 0.03 | 0.03 |  | 1.2 |  | 16.1 | 0.23 |
| 9/8/02 |  | 0.03 |  | 1.7 |  |  |  |
| 9/18/02 |  | 0.03 | 0.04 |  |  |  |  |
| 9/24/02 |  |  |  | 1.1 |  |  | 0.29 |
| 9/26/02 |  | 0.03 | 0.04 | 2.1 |  |  | 0.28 |
| 9/30/02 |  | 0.03 | 0.03 | 1.1 |  |  |  |
| 10/1/02 |  |  |  |  |  |  | 0.31 |
| 10/9/02 |  |  |  | 1.4 |  |  |  |
| 10/15/02 |  |  |  |  |  |  | 0.24 |
| 10/21/02 |  |  |  | 1.1 |  |  |  |
| 10/28/02 |  | 0.03 | 0.04 | 1.5 | 0.4 |  | 0.25 |
| 10/30/02 |  | 0.03 |  | 1.5 |  |  | 0.24 |
| 11/12/02 |  |  |  |  |  |  | 0.25 |
| 11/17/02 |  |  |  |  |  |  | 0.35 |
| 11/23/02 |  |  |  |  |  |  | 0.74 |
| 11/26/02 |  |  |  |  |  |  | 0.79 |
| 12/11/02 |  |  |  | 1.6 |  |  |  |
| 12/17/02 |  |  |  | 1.2 |  |  | 0.66 |
| 12/22/02 |  |  |  | 1.5 |  |  |  |
| 12/27/02 |  |  |  |  |  |  |  |

NOTE:  DOC values in red represent those samples filtered through a glass 0.7 µm filter.  Values in black

Represent those samples filtered through a 0.4 µm filter.

Table 4.  Summary of precision data for field replicates.

| Variable | Median | Range | QA Objective | n |
| --- | --- | --- | --- | --- |
| pH, closed cell (values ( 5.74) | 0.01 | -0.09 to .17 | ( 0.075 | 31 |
| pH, closed cell (values ( 5.74) | 0.01 | -0.12 to 0.15 | ( 0.15 | 35 |
| pH, equilibrated (values ( 5.74) | -0.01 | -0-04 to 0.01 | ( 0.075 | 21 |
| pH, equilibrated (values ( 5.74) | -0.01 | -0.18 to 0.07 | ( 0.15 | 45 |
| ANC (values (100 µeq/L) | 0.1 | -26.9 to 3.7 | ( 5 µeq/L | 59 |
| ANC (values (100 µeq/L) | 0.65% | 0 to 2.05% | ( 5% | 7 |
| Conductivity (values ( 50uS/cm) | 0 | -1.6 to 1.3 | ( 1 uS/cm | 61 |
| Conductivity (values ( 50uS/cm) | 0.47% | 0.24 to 0.60% | ( 2% | 5 |
| DOC (values ( 2 mg/L) | 0.02 | -0.13 to 0.16 | ( 0.1 mg/L | 19 |
| DOC (values ( 2 mg/L) | 1.59% | 0 to 19.9% | ( 5% | 47 |
| DIC (values ( 2 mg/L) | 0 | -2.4 to 0.2 | ( 0.1 mg/L | 57 |
| DIC (values ( 2 mg/L) | 2.89% | 1.26 to 4.16% | ( 5% | 5 |
| Al, total dissolved (values ( 100 µg/L) | 0.2 | -6.2 to 6.9 | ( 10 µg/L | 25 |
| Al, total dissolved (values ( 100 µg/L) | 1.18% | 0 to 6.2% | ( 10% | 32 |
| Ca (values ( 0.4 mg/L) | 0.01 | 0.01 to 0.01 | ( 0.02 mg/L | 1 |
| Ca (values ( 0.4 mg/L) | 0.99% | 0 to 11.9% | ( 5% | 65 |
| Mg (values ( 0.2 mg/L) | 0.01 | 0 to 0.01 | ( 0.01 mg/L | 2 |
| Mg (values ( 0.2 mg/L) | 0.50% | 0 to 10.4% | ( 5% | 64 |
| Na (mg/L) | 1.43% | 0 to 15.0% | ( 5% | 66 |
| K (values ( 0.8 mg/L) | 0 | -0.08 to 0.11 | ( 0.04 mg/L | 62 |
| K (values ( 0.8 mg/L) | 2.74% | 1.64 to 4.36% | ( 5% | 4 |
| NH4 (mg/L) | 0 | -0.02 to 0.02 | ( 0.02 mg/L | 65 |
| Cl (values ( 20 µeq/L) | 0 | -2 to 8.5 | ( 1.0 µeq/L | 10 |
| Cl (values ( 20 µeq/L) | 0.63% | 0 to 6.7% | ( 5% | 56 |
| NO3 (values ( 10 µeq/L) | 0 | -0.6 to 0.3 | ( 0.5 µeq/L | 58 |
| NO3 (values ( 10 µeq/L) | 0.26% | -0.2 to 0.54% | ( 5% | 8 |
| SO4 (ueq/L) | 0% | 0 to 2.9% | ( 5% | 66 |
| Total P (µg/l) | 0.3 | -3 to 7 | ( 1 µg/L | 33 |
| Total N (mg/l) | 2.58% | 0 to 68% | ( 5% | 42 |
| Si (values ( 1 mg/L) | 0 | -0.07 to 0.05 | ( 0.05 mg/L | 21 |
| Si (values ( 1 mg/L) | 0% | 0 to 23% | ( 5% | 45 |
| True Color (values ( 50 PCU) | 0 | -3 to 2 | ( 5 PCU | 58 |
| True Color (values ( 50 PCU) | 0% | 0 to 0% | ( 10% | 5 |

Table 5.  Itemization of field replicate precision violations.

| Site | Sample Date | Analyte | Duplicate Values | Precision |
| --- | --- | --- | --- | --- |
| Gould | 10-23-02 | DOC | 5.26, 4.65 | 8.71% |
| Gould | 10-23-02 | Total P | 17, 15 | 2.00 µg/L |
| Gould | 10-23-02 | Total N | 0.35, 0.378 | 5.44% |
| Black | 10-27-02 | DIC | 2, 1.8 | 0.20 mg/L |
| Black | 10-27-02 | Total P | 10, 7.1 | 2.90 µg/L |
| SANDB | 11-16-02 | Ca | 1, 0.93 | 5.13% |
| SANDB | 11-16-02 | DOC | 4.29, 3.99 | 5.12% |
| SANDB | 11-16-02 | Total P | 17, 10 | 7.00 µg/L |
| Hamil | 11-20-02 | Total N | 0.25, 0.182 | 22.26% |
| Rock Pond | 10-9-02 | Na | 0.88, 0.75 | 11.28% |
| Rock Pond | 10-9-02 | K | 0.4, 0.33 | 0.07 mg/L |
| Rock Pond | 10-9-02 | Cl | 13, 10 | 3.00 µeq/L |
| Greenwood Pd | 10-9-02 | Ca | 0.5, 0.43 | 10.64% |
| Greenwood Pd | 10-9-02 | DOC | 1.89, 2.11 | 7.78% |
| CB1 | 1-24-02 | Ca | 0.89, 0.96 | 5.35% |
| CB1 | 1-24-02 | Mg | 0.35, 0.38 | 5.81% |
| CB1 | 2-20-02 | Total N | 0.023, 0.028 | 13.86% |
| HB1 | 3-6-02 | ClpH | 5.48, 5.57 | -0.09 |
| CB1 | 3-13-02 | ClpH | 5.66, 5.57 | 0.09 |
| CB1 | 3-13-02 | DIC | 0.7, 0.9 | -0.02 mg/L |
| CB1 | 3-13-02 | Total N | 0.07, 0.053 | 19.55% |
| HB1 | 4-24-02 | Total P | 2.5, 1.1 | 1.40 µg/L |
| HB1 | 5-29-02 | ANC | 41.1, 68 | -26.90 µeq/L |
| HB1 | 5-29-02 | Ca | 1.47, 1.74 | 11.90% |
| HB1 | 5-29-02 | Mg | 0.44, 0.51 | 10.42% |
| HB1 | 5-29-02 | K | 0.25, 0.3 | -0.05 mg/L |
| HB1 | 5-29-02 | EqpH | 6.83, 7.01 | -0.18 |
| HB1 | 5-29-02 | Si | 2.8, 3.9 | 23.22% |
| HB1 | 5-29-02 | Total N | 0.054, 0.153 | 67.64% |
| CB1 | 6-19-02 | Total N | 0.077, 0.098 | 16.97% |
| HB1 | 7-6-02 | Ca | 1.89, 2.05 | 5.74% |
| HB1 | 7-6-02 | Mg | 0.55, 0.6 | 6.15% |
| HB1 | 7-6-02 | DIC | 1.1, 3.5 | -2.40 mg/L |
| HB1 | 7-6-02 | Total P | 2.7, 4.2 | -1.50 µg/L |
| HB1 | 7-6-02 | Total N | 0.12, 0.168 | 23.57% |
| Salmon E | 11-4-02 | DIC | 1.6, 1.4 | 0.20 mg/L |
| Newbert E | 7-24-02 | Si | 0.86, 0.93 | -0.07 mg/L |
| Abol E | 7-31-02 | DOC | 2.08, 2.76 | 19.87% |
| Abol | 10-15-02 | Cl | 20, 22 | 6.73% |
| Bean E | 8-1-02 | Cl | 12, 14 | -2.00 µeq/L |
| Bean E | 8-1-02 | DIC | 0.9, 0.7 | 0.20 mg/L |
| Bracey E | 5-16-02 | K | 0.3, 0.24 | 0.06 mg/L |
| Crystal E | 8-5-02 | ClpH | 5.44, 5.27 | 0.17 |
| Abol O | 4-18-02 | DIC | 1.7, 2 | -0.30 mg/L |
| Abol O | 4-18-02 | Total P | 9.5, 5.2 | 4.30 µg/L |
| Bean O | 4-30-02 | Cl | 19, 16.6 | 2.40 µeq/L |
| Bean O | 4-30-02 | Total P | 13, 9.7 | 20.56% |
| Bean O | 5-6-02 | Total N | 0.235, 0.262 | 7.68% |
| Mud O | 3-4-02 | K | 0.41, 0.36 | 0.05 mg/L |
| Mud O | 3-4-02 | DIC | 0.5, 0.7 | -0.20 mg/L |
| Mud O | 3-4-02 | Total N | 0.169, 0.149 | 8.89% |
| Mud O | 3-28-02 | DOC | 4.92, 4.53 | 5.84% |
| Mud O | 3-28-02 | Sp Cond | 33.50, 35.00 | -1.50 µS/cm |
| Partridge O | 4-22-02 | DIC | 0.6, 0.8 | -0.20 mg/L |
| Wiley O | 3-11-02 | Sp Cond | 28.00, 29.60 | -1.60 µS/cm |
| Round Pond | 7-15-02 | DIC | 1.4, 1.6 | -0.20 mg/L |
| Miller | 9-2-02 | Na | 0.99, 0.8 | 15.01% |
| Miller | 9-2-02 | K | 0.29, 0.18 | 0.11 mg/L |
| Miller | 9-2-02 | Cl | 18, 9.5 | 8.50% |
| Miller | 9-2-02 | Sp Cond | 20.1, 18.8 | 1.30 µS/cm |
| Round | 9-2-02 | DIC | 1.6, 2 | -0.40 mg/L |
| Baker | 6-6-02 | K | 0.37, 0.45 | -0.08 mg/L |
| Baker | 6-6-02 | Cl | 38, 41 | 5.37% |
| Baker | 6-6-02 | NO3 | 0, 0.6 | -0.60 µeq/L |
| EB8 | 4-5-02 | DOC | 2.36, 2.15 | 6.59% |
| EB8 | 4-5-02 | Total N | 0.094, 0.081 | 10.51% |
| EB8 | 4-30-02 | DOC | 2.04, 1.88 | 5.77% |
| EB8 | 4-30-02 | Total P | 1.4, 4.4 | -3.00 µg/L |
| EB8 | 4-30-02 | Total N | 0.085, 0.075 | 8.84% |
| EB8 | 5-28-02 | DOC | 2.08, 1.9 | 6.40% |
| EB8 | 5-28-02 | Total P | 4.7, 2.2 | 2.50 µg/L |
| WB3 | 2-5-02 | Total P | 3.1, 2 | 1.10 µg/L |
| WB3 | 2-26-02 | Total P | 1.1, 3 | -1.90 µg/L |
| WB3 | 3-20-02 | Total P | 4.4, 2 | 2.40 µg/L |
| WB3 | 5-14-02 | DOC | 1.93, 1.77 | 0.16 mg/L |
| WB3 | 5-14-02 | Total N | 0.181, 0.219 | 13.44% |
| WB3 | 6-11-02 | Total N | 0.113, 0.134 | 12.02% |

Table 6. Summary of precision data for laboratory duplicates.

| Variable | Median | Range | QA Objective | n |
| --- | --- | --- | --- | --- |
| pH, closed cell (values ( 5.74) | 0 | -0.85 to 0.07 | ( 0.075 | 44 |
| pH, closed cell (values ( 5.74) | 0.01 | -0.18 to 0.08 | ( 0.15 | 89 |
| pH, equilibrated (values ( 5.74) | 0 | -0.4 to 0.04 | ( 0.075 | 57 |
| pH, equilibrated (values ( 5.74) | 0 | -0.53 to 0.12 | ( 0.15 | 81 |
| ANC (values (100 µeq/L) | -0.09 | -10.77 to 9.01 | ( 5 µeq/L | 93 |
| ANC (values (100 µeq/L) | 0.39% | 0 to 3.45% | ( 5% | 58 |
| Conductivity (values ( 50uS/cm) | -0.04 | -0.71 to 0.41 | ( 1 uS/cm | 74 |
| Conductivity (values ( 50uS/cm) | 0.20% | 0 to 0.72% | ( 2% | 8 |
| DOC (values ( 2 mg/L) | -0.02 | -0.3 to 0.17 | ( 0.1 mg/L | 15 |
| DOC (values ( 2 mg/L) | 0.74% | 0 to 13.9% | ( 5% | 85 |
| DIC (values ( 2 mg/L) | 0 | -0.10 to 0.11 | ( 0.1 mg/L | 50 |
| DIC (values ( 2 mg/L) | 1.66% | 0 to 3.9% | ( 5% | 10 |
| Al, total dissolved (values ( 100 µg/L) | -0.03 | -4.9 to 4.4 | ( 10 µg/L | 52 |
| Al, total dissolved (values ( 100 µg/L) | 0.57% | 0 to 5.62% | ( 10% | 29 |
| Ca (values ( 0.4 mg/L) | 0 | -0.02 to 0.005 | ( 0.02 mg/L | 13 |
| Ca (values ( 0.4 mg/L) | 0.76% | 0 to 8.55% | ( 5% | 69 |
| Mg (values ( 0.2 mg/L) | 0 | -0.002 to 0.001 | ( 0.01 mg/L | 16 |
| Mg (values ( 0.2 mg/L) | 0.73% | 0 to 7.16% | ( 5% | 65 |
| Na (values ( 0.46 mg/L) | 0 | -0.002 to 0.02 | ( 0.02 | 12 |
| Na (values ( 0.46 mg/L) | 1.03% | 0 to 7% | ( 5% | 69 |
| K (values ( 0.8 mg/L) | 0 | -0.02 to 0.03 | ( 0.04 mg/L | 77 |
| K (values ( 0.8 mg/L) | 1.52% | 0 to 7.44% | ( 5% | 33 |
| NH4 (values ( 0.36 mg/L) | 0 | -0.01 to 0.01 | ( 0.02 mg/L | 83 |
| NH4 (values > 0.36 mg/L) | 0% | 0 to 0% | ( 5% | 1 |
| Cl (values ( 20 µeq/L) | 0.04 | -0.98 to 1.35 | ( 1.0 µeq/L | 39 |
| Cl (values ( 20 µeq/L) | 0.14% | 0 to 2.87% | ( 5% | 105 |
| NO3 (values ( 10 µeq/L) | 0 | -0.68 to 2.52 | ( 0.5 µeq/L | 120 |
| NO3 (values ( 10 µeq/L) | 0.23% | 0.05 to 0.83% | ( 5% | 24 |
| SO4 (values ( 20 µeq/L) | 0 | -0.23 to 0.20 | ( 1.0 µeq/L | 14 |
| SO4 (values ( 20 µeq/L) | 0.12% | 0 to 1.93% | ( 5% | 130 |
| Total P (values ( 20 µg/l) | 0.10 | -4.60 to 5.80 | ( 1 µg/L | 72 |
| Total P (values > 20 µg/l) | 5.41% | 0 to 63.23% | ( 5% | 30 |
| Total N (mg/l) | 0.85% | 0 to 8.47% | ( 5% | 73 |
| Si (values ( 1 mg/L) | 0 | -0.04 to 0.05 | ( 0.05 mg/L | 42 |
| Si (values ( 1 mg/L) | 1.10% | 0 to 5.83% | ( 5% | 42 |
| True Color (values ( 50 PCU) | 0 | -1.00 to 1.00 | ( 5 PCU | 98 |
| True Color (values ( 50 PCU) | 0% | 0 to 0% | ( 10% | 13 |