AFE8 Board Qualification Criteria

Fred Borcherding / Jerry Blazey

D0 Electronics Group

This note defines the criteria to be met by a single pair of left- and right-handed AFE8 prototype boards.  Production of the AFE8s requires satisfaction of these criteria.

# General Requirements.

AFE8 boards must: supply bias voltage to the VLPC chips, monitor and regulate the cassette temperature, supply discriminated signals to the trigger and readout, and supply analog hit information for calibration. The specifications in this document are given in fC delivered to the AFE8 board from the VLPC.  With a VLPC gain of about 40K and a nominal bias voltage setting, 7fC roughly translates to one PE.

## Bias Voltage.

The bias settings must have a resolution and stability of < 50mV.

## Temperature Control.

Temperatures stable to  <100 mK after calibration.

The following specifications must be performed  (1) on the LAB3 test cryostat, (2) with operating cryogenics and bias controls, (3) with an operating matching board (both a LHB and RHB must be present), and (4) with a single set of threshold voltages for all channels of a single SIFT.

## Discriminated Signals.

On a single AFE8, data will be collected from each MCM with a patterned light mask such that one channel out of four is exposed.  The exposed channels will be illuminated with an average of 1.5 PE.  Thresholds shall be set such that the discriminator rate for the unexposed channels on a single MCM averages 0.5%.   The variance of the 128 thresholds will be calculated and used to simulate the expected efficiency of the track trigger.  This efficiency must be acceptable for the boards to be approved.

With these fixed thresholds and after five days uninterrupted running the firing rates must not increase by more than a factor of two.

The performance must not degrade with the addition of a 20MHz background.

The transmission to the DFE must have an error rate of less then 0.1%.  That is discriminator output must be transmitted on the proper output link, line, and time slice with errors less than 0.1%.  It must do this with 53MHz link speed.

The discriminated data must be sampled, stored, and transferred to the sequencer for L3 readout with less than a 0.1% error rate per full data set.  The information transferred must imitate the SVX chip readout such that the sequencer correctly receives and transmits this to the VRB.  The information must include address and data.  The transfer of the complete set of bits of data plus address information must be less than 8us.

## Fiber Analog Information.

The AFE8 must be able to add channel amplitude information through the SVX chip to the sequencer readout.  This can be requested dynamically though a command from the online control system. The data must be of sufficient quality such that fitting routines can separate the first and second photoelectron peaks for greater than 50% of the channels for each MCM.  This must be done for low and high gain cassettes

The system must be able to transfer the data without readout errors less than 0.1%.   No information or should be lost, bits dropped, or signal variations introduced in the transfer.  Readout completion must occur before completion of the SMT readout.

.

## Mechanical Criteria.

Mechanical specifications must be met with a fully populated and operating LHB and a fully populated and operating RHB.  Before and after reaching thermal equilibrium both boards mounted together must (1) properly mate with a cassette including both the back plane connectors and the high-density connectors along the bottom edge and (2) the two boards must also be mounted without contact.   In addition (3) two adjacent boards each belonging to two different but adjacent cassettes must be mounted without contact.

Specification Checklist

| Requirement | Date Complete |
| --- | --- |
| Bias Voltage |  |
| 50 mV resolution | 14-Mar-01 |
| 50 mV stability |  |
| Temperature Control |  |
| 100 mK stability |  |
| Discriminated Signals |  |
| Variance of 128 signals |  |
| Trigger Simulation Complete |  |
| DFE Transmission Error Free | 14-Mar-01 |
| SEQ Transmission Error Free, < 8 us |  |
| 20 MHz Pulser Test * |  |
| Five day stability * |  |
| Analog Information |  |
| SEQ Transmission Error Free, < 8 us |  |
| Fit 50% channels, high gain |  |
| Fit 50% channels, low gain * |  |
| Mechanical |  |
| Properly Seated |  |
| No contact intracryostat |  |
| No contact intercryostat |  |
| * Not required |  |