WATER QUALITY MONITORING PROTOCOL

FOR KETTLE PONDS OF

CAPE COD NATIONAL SEASHORE

Prepared by:

J. Portnoy, K. Lee, J. Oset, E. Gwilliam & J. Budreski

April 2003

Table of Contents

Preface	4

Acknowledgements	4

Introduction and Objectives	5

Integration with Other CACO Monitoring	6

Background for selection of monitoring variables	7

Brief explanation of selected monitoring variables	9

Measurement methods	11

Seasonal Sampling Program and Collection Overview	12

Sample Location and Procedure	14

Field Logistics and Safety	15

Data management/Reporting/Staffing	15

Literature Cited	17

Appendix A. – YSI (temperature, conductivity, redox potential, dissolved oxygen, pH, depth)	18

Appendix B. – Transparency (Secchi and Li-cor Measurements)	24

Appendix C. – pH and Alkalinity titration (potentiometric, low-level)	27

Appendix D. – Dissolved Nutrients Analysis (Ammonium, Nitrate/Nitrite and Orthophosphate)	29

Appendix E. - Total Phosphorus and Nitrogen: Acid Persulfate Digestion	31

Appendix F – Cations (Ca, Mg, K, Na)	37

Appendix G – Anions (Cl- and SO4-)	38

Appendix H - Sampling Procedures for the Pond Monitoring Program	39

Appendix I – PondStage.	47

Appendix J - Data Forms.	48

Appendix K.  The Need for Annual Sampling.	52

Tables and Figures:

Table 1.	CACO Ponds Included in the Kettle Pond Monitoring Program.	4

Table 2.	Variables and Measurement Methods.	11

Table 3.	Changes in Secchi Transparency (m) Versus the Seasonal Range at CACO Kettle Ponds  as an Indicator of Necessary Sampling Frequency.	13

Table 4.	Summary of water quality variables by sampling period - Field Sampling Information	14

# Figure 1.     Map of CACO ponds………………………………………………………………5

# Figure 2.     Flow diagram of pond water sample collection………………………….…13

# Preface

This protocol describes the conceptual underpinnings and field and laboratory methods employed in the monitoring of 20 kettle ponds within Cape Cod National Seashore (CACO).  Selection of monitoring variables and methods derived from three prior reviews (Martin et al. 1993; Portnoy and Farris 1996; Portnoy and Lee 2000) by scientists and resource managers with expertise and experience in Cape Cod limnology, hydrology and biogeochemistry.

# Acknowledgements

CACO is indebted to the following individuals for periodic technical assistance and critical review: Elizabeth Colburn, Harvard Forest, Anne Giblin, Marine Biological Laboratory, Paul Joseph Godfrey, University of Massachusetts, Harry Hemond, Massachusetts Institute of Technology, National Park Service, Charles Roman, U.S. Geological Survey, University of Rhode Island, Marjorie Winkler, Center for Climatic Research, University of Wisconsin.

## CACO Ponds Included in the Kettle Pond Monitoring Program.  See Summer Biweekly Monitoring P.11.  GPS X/Y Coordinates (Easting/Northing) of deepest point.

| PRIMARY Ponds |  | Easting |  | Northing |  | Depth (ft) |  | SECONDARY Ponds |  | Depth (ft) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Gull |  | 416491 |  | 4644967 |  | 60 |  | Higgins |  | 20 |
| Duck |  | 417050 |  | 4642801 |  | 60 |  | Williams |  | 4 |
| Ryder |  | 414538 |  | 4646672 |  | 39 |  | Kinnacum |  | 12 |
| Spectacle |  | 417457 |  | 4644849 |  | 24 |  | Northeast |  | 11 |
| Great (Truro) |  | 414652 |  | 4647453 |  | 36 |  | Southeast |  | 12 |
| Great (Wellfleet) |  | 417193 |  | 4643575 |  | 52 |  | Turtle |  | 3 |
| Dyer |  | 416348 |  | 4643454 |  | 21 |  | Horseleech |  | 13 |
| Long |  | 416303 |  | 4644036 |  | 47 |  | Slough |  | 25 |
| Snow |  | 414867 |  | 4646937 |  | 23 |  | Round (west) |  | 31 |
| Herring |  | 415967 |  | 4646237 |  | 12 |  | Round (east) |  | 24 |

# Introduction and Objectives

Kettle pond management issues and monitoring objectives are detailed in Martin et al. (1993).  In brief, concerns include:

1) Increased human use and accompanying nutrient (phosphorus and nitrogen) loading from a variety of anthropogenic sources including:

a) Bank and soil erosion;

b) Shoreline wastewater disposal;

c) Visitors and their pets;

2) Atmospheric deposition of metals, sulfate, and nitrate and their effects on pH, alkalinity and sulfur and phosphorus cycling;

3) Hydraulic alterations, especially stream channelization, which may alter hydrology, trophic structure and native species diversity;

4) Fish stocking, for the enhancement of recreational fishing, and its possible effects on pond trophic dynamics.

To document current status and to monitor future change in pond water quality, this monitoring program is designed to:

1) Characterize the trophic status of the ponds;

2) Develop repeatable methods of monitoring trophic status and limnological processes to detect important changes;

3) Describe pond-specific and seasonal limnological and hydrological processes affecting water quality;

4) Identify and develop management actions to mitigate anthropogenic effects.

The selection of monitoring variables and temporal and spatial sampling frequencies is guided by expectations (i.e. hypotheses) regarding the biogeochemical and limnological responses of the ponds to added nutrients (nitrogen and phosphorus), acids (as from “acid rain”) or to manipulations of hydrography or of the resident fish community.  Monitoring the effects of "external" (extra-watershed) phosphorus loading, for example, on pond trophic status requires data on nutrient status, transparency, and dissolved oxygen.

# Integration with Other CACO Monitoring

Besides assessing the effects of nutrient and acid loading on the limnology and ecology of the ponds themselves, this monitoring program provides important information for a diverse array of research and management questions.  Pond water level data is critical to hydrologic monitoring which is relevant to understanding pond water quality changes that are linked to water level and/or flow direction.  Pond chemistry affects biota and often determines the species composition and abundance of aquatic macrophytes, benthic and pelagic invertebrates and fish, including rare species.  Pond chemistry has also been shown to reflect long-term and landscape-scale changes in the surrounding uplands, and thus to integrate watershed-wide changes in soils and vegetation.  Data from the National Atmospheric Deposition Program (NADP), including the local CACO-operated Truro site (MA01), have been employed for over two decades to help interpret pond chemistry and to assess the importance of atmospheric deposition in eutrophication and acidification.

# Background for selection of monitoring variables

##### Nutrient loading and eutrophication

Phosphorus (P) is the nutrient most likely limiting primary production and controlling water clarity in the clearest (oligotrophic) kettle ponds.  Local glacial outwash soils are low in P, probably explaining the natural clarity of most ponds.  [Nitrogen, the other major nutrient for plant and algal growth, is supplied by microbial fixation of atmospheric nitrogen gas in freshwater environments.]  It is hypothesized that kettle pond water quality deteriorates when humans import P into the pond basin.  P inputs from precipitation are minor compared to other sources, based upon data collected since 1981 at the NPS-operated NADP site in South Truro.  Wildlife (primarily bird) contributions of P are probably of importance for Gull Pond (NADP data, Portnoy and Soukup 1990).  With increasing shoreline development, cottage occupancy and pond use, humans provide additional P to pond nutrient budgets from septic wastes and through shore erosion.  Human wastes from shoreline septic systems are likely the major source of P to the kettle ponds.

Because P has no common gaseous form, once imported into a lake watershed it tends to remain there. Most P is adsorbed to iron minerals on or in the sediments, where it remains unless released mechanically or through a change in the chemical environment.  Chemical release of P generally occurs in summer as a result of thermal stratification and the loss of oxygen from bottom waters.  The process is as follows:  In spring and summer pond surface waters warm.  This warm water tends to remain “floating” at the surface because it is less dense than colder water; also less wind energy is available for mixing than in fall and winter.   Colder water accumulates at the bottom and, because of its higher density, is not readily brought to the surface by weak summer winds.  Bottom water thereby stagnates during this thermal stratification.  Because there is little or no light in the bottom waters of the deep (10-20 m) kettle ponds, oxygen production through photosynthesis is nil, while oxygen consumption by bacteria which decompose organic matter is high.  Oxygen depletion (anoxia) of the bottom waters results.  Once oxygen gas is no longer available to bacteria for the oxidation (decomposition) of organic matter, the microbes then use oxidized iron minerals.  This results in the dissolution of the iron minerals and, most importantly for pond water quality, the release of phosphate that is normally immobilized by the iron under more aerobic conditions. Fall brings surface water cooling so that the density difference between surface and bottom waters decreases and autumn winds can eventually mix the entire water column, i.e. fall overturn. With renewed mixing, PO4-enriched bottom water is carried to the well-lit surface where it is assimilated by phytoplankton, thereby stimulating algal blooms.

Pond eutrophication and accompanying algal blooms can cause earlier and more intense heating of the surface waters because the algal cells suspended in the water column absorb solar energy and re-radiate it as heat, unlike clear water which allows much of the energy to pass through to deeper water.  Thus, more eutrophic ponds thermally stratify sooner in the warm months, resulting in more prolonged and extensive deep-water anoxia and P release, which can feed back to increase algal growth.  The degree of temperature stratification and oxygen depletion within the water column can be monitored with submersible temperature, conductivity and oxygen probes.

Iron is also active in the lake's sulfur cycle.  Chemically reduced iron (Fe++), which is found in the sediment and in bottom waters during oxygen depletions, precipitates the sulfides produced by anaerobic sulfate-reducing bacteria in the sediment.  Sulfide precipitation may in fact limit PO4 binding sites on Fe complexes.  Thus, the presence of free sulfides, particularly the foul-smelling hydrogen sulfide (H2S) gas, can indicate that Fe is in short supply for sulfide binding, and more importantly, for phosphate complexation.  The odor of H2S in collected bottom water therefore indicates PO4 release from anaerobic sediment.  Acid deposition may increase the competition between PO4 and sulfide for Fe.  The Cape’s acidic precipitation is acidified primarily by sulfate contamination.  In this way  “acid rain” can contribute to pond eutrophication, as well as acidification (Caraco et al. 1989).

Nutrient additions manifest their effect in the water column; however, phytoplankton assimilation of a limiting nutrient is so rapid that little dissolved inorganic nitrogen or phosphorus may be detectable in pond water even during conditions of high loading.  Also, although P clearly limits algal growth in oligotrophic ponds, once P concentrations increase above a critical threshold (e.g. 30 ppb,) (Downing and  McCauley 1992 ), nitrogen can limit phytoplankton growth.  Thus, to assess total nutrient status, whole water samples must be chemically digested to measure their total phosphorus (TP) and nitrogen (TN) content.  TP and TN include all dissolved and particulate forms of these elements in the water column.

Increased algal productivity in kettle ponds should be evident as reduced clarity, as algal biomass increases.  Clarity can be measured directly using light attenuation and Secchi depth; chlorophyll concentration serves as a widely used surrogate for phytoplankton biomass.  Chlorophyll measurements were discontinued in 2002 due to poor data trends and major variability in sampling and analyses.

##### Acidification

The scarcity of natural inorganic buffers in native granitic soils makes the kettle ponds apparently sensitive to acid deposition.  All 20 kettle ponds are poorly buffered with alkalinities less than 4 mg/L and pH ranging from 4.3 to 6.5.   For the past two decades of record, precipitation pH has averaged about 4.6 with high sulfate loading of surface waters.  However, pond pH with few exceptions has remained stable both over the short-term (1981 – present) and apparently over the past 10 millennia (Winkler 1988).  Nevertheless, there is the potential for long-term effects on base cation leaching and pH due to continued acid deposition within pond watersheds.  Further, nitrate contamination from shoreline septic systems has the potential to acidify surface waters once the microbial community within a pond which normally removes nitrate is saturated (Kelly et al. 1990) .

# Brief explanation of selected monitoring variables

Pond Stage, especially relative to adjacent groundwater levels, affects both inflow and outflow of water and dissolved species.  It also affects the redox potential and pH of littoral sediments, in turn affecting the mobility of sedimentary Fe, P and S.

Temperature is measured along with electrical conductivity (specific conductance) to follow thermal stratification of the water column.  Water masses of different temperature and, thus, density can stratify with little mixing between layers during summer months leading to stagnant, and often oxygen-poor, conditions in bottom waters.  Temperature profiling describes the relative thermal resistance to mixing.  Anoxia induces the chemical reduction of iron and the release of sedimentary phosphorus and sulfides into the water column.  Phosphorus release can trigger algal blooms, if transported to well-lit depths, and sulfides are toxic to aquatic life.  Temperature profiles can also be useful for evaluating groundwater inputs and outputs and for computing heat budgets.

Electrical conductivity is a measure of the amount of dissolved constituents in water.  It is temperature-dependent (and readings must be temperature-compensated) and thus aids in identifying thermal stratification. More importantly, the loss of oxygen from deep pondwater in summer causes the dissolution of limnologically important substances (iron, phosphorus, and sulfides) that show up as increased conductivity. Lower conductivities are often associated with purer waterbodies.

Dissolved oxygen.  Besides being essential for aerobic respiration (and therefore critical for the survival of aquatic life), oxygen concentrations influence all major nutrient cycles, including the storage of phosphorus in lake sediments.  Lakes undergoing eutrophication typically show increases in the duration of deep-water anoxia and the thickness (i.e. vertical extent) of anoxic bottom waters.  By calculating the anoxic factor for our primary ponds, trends in pond health can be followed (Nurnberg, 1995).

Light penetration.  A light meter is used to measure incident light (visible range) above the lake surface and at specific depths, preferably simultaneously, to calculate the percent transmission through the intervening water column.  Photosynthesis and aquatic plant production are dependent upon available light; thus, photosynthetic oxygen production is often limited by light transmission to the bottom waters of turbid lakes.  This phenomenon can exacerbate deep-water oxygen depletions.  Light profiles can also show where plankton organisms are concentrated and thereby explain subsurface peaks in the oxygen profile.

Secchi depth.  A single measurement of the maximum depth at which the black and white 20-cm Secchi disk can be seen is widely used as an index of clarity. The Secchi disk measures the transparency of the water. The transparency can be affected by the color of the water, suspended sediments and algae. Transparency decreases as color, suspended sediments, or algal abundance increases.  Water may be stained yellow or brown by decaying plant matter; however, Cape Kettle Ponds have very low dissolved organic material (DOM), therefore, color is relatively unimportant in affecting water clarity.  Suspended sediments often come from sources such as re-suspension from the lake bottom, construction sites, agricultural fields, and urban storm runoff.   Algae abundance is related to the amount of plant nutrients, especially phosphorus and nitrogen. Transparency can therefore be affected by the amount of plant nutrients coming into the lake from sources such as sewage treatment plants, septic tanks, and lawn and agricultural fertilizer (Great North American Secchi Dip-In, 2002).

Transparency is an indicator of the impact of human activity on the land surrounding the lake. If transparency is measured through the season and from year to year, trends in transparency may be observed. Transparency can serve as an early warning that activities on the land are having an effect on a lake.

Hydrogen ion concentration is represented by pH, the negative logarithm of [H+], and measures the acid-base balance of water.  Pond water pH is determined by the combination of ions in solution. Ionic composition is affected by surrounding soils, atmospheric inputs and in-pond chemical and biological processes.  Conversely, pH affects both chemical reactions and the aquatic biotic community.

Alkalinity measures the acid-neutralizing (i.e. buffering) capacity of pondwater, and is therefore a measure of the ability to resist pH change while receiving acidic inputs, as from acid deposition.

Total Nitrogen and Total Phosphorus (TN and TP) are measures of the total concentration of these elements in the water column, including dissolved species, which are scarce in nutrient-limited kettle ponds, and the more abundant N and P contained in organic particles including plankton.  To measure TN and TP colorimetrically, water samples must be digested to convert all N and P to inorganic forms.  Since P is rapidly cycled by water column biota in P-limited pond systems, soluble reactive P is often undetectable in the water column and total P is a better measure of this element's availability.  Total P is measured by digestion of unfiltered samples.

Dissolved Nutrients, (orthophosphate, nitrate and ammonium) are included as potentially sensitive indicators of septic leachate loading. Because most ponds are likely P- rather than N-limited, dissolved inorganic N should persist in the water column to provide a tracer of contamination from shoreline wastewater disposal systems.

Anions  (SO4- and Cl-) are indicators of the atmospheric deposition and marine influences on the ponds.  SO4 is the principal acidifying anion in Cape Cod precipitation and, in addition to acidification, can contribute to eutrophication by enhancing P release from sediment.  Concurrent chloride monitoring allows SO4 from marine aerosols to be subtracted from the total, yielding a measure of non-sediment SO4.

Cations  (Mg, K, Na, Ca) are essential both physiologically and for maintaining a stable acid balance in pond water.  Cation depletion has been observed in freshwater and forest soils in areas of the Northeast United States receiving acid precipitation.

# Measurement methods

Field and laboratory measurements employ the following equipment and methods.  Staff adheres strictly to these procedures unless analytical problems or safety concerns arise.  Deviation from these protocols for any reason is reported to the project Technical Coordinator within 24 hours to ensure rapid resolution of the problem and continuance of the program.  Table 2 lists current methods used.

## Variables and Measurement Methods.

| Variable | Method |
| --- | --- |
| Temperature | YSI (see Appendix A,H) |
| Conductivity | YSI (see Appendix A,H) |
| Redox potential (Eh) | YSI (see Appendix A,H) |
| Dissolved oxygen | YSI (see Appendix A,H) |
| pH for Profiles | YSI (see Appendix A,H) |
| pH for Apr, Jul, Oct survey | Beckman meter with glass KCl-filled electrode, calomel junction (see Appendix C,H) |
| Transparency | Licor light meter with two underwater sensors (see Appendix B,H) |
|  | Secchi disk (20-cm diameter, black and white quadrants) (see Appendix B,H) |
| Alkalinity | Potentiometric titration using Beckman meter, Dosimat (see Appendix C,H) |
| Ammonium | Phenolate method, flow injection analysis (FIA-Lachat QuickChem Method 10-107-06-1-C; see Dissolved Nutrients Analysis Appendix D) |
| Nitrate | Cadmium reduction (FIA-Lachat QuickChem Method 31-107-04-1-C, see Dissolved Nutrients Analysis Appendix D) |
| Orthophosphate | Ammonium molybdate and antimony potassium tartrate -ascorbic acid reduction (FIA-Lachat QuickChem Method 31-115-01-1-G, see Dissolved Nutrients Analysis Appendix D) |
| Total phosphorus/ Total nitrogen | FIA- Simultaneous Determination of Total Phosphorus and Total Nitrogen in Natural Waters (alkaline persulfate digestion, see Appendix E) |
| Major cations Ca++, Mg++, K+, Na+ | Atomic absorption spectrometry (see Appendix F,H) |
| Major anions SO4 and Cl- | Turbidimetric method for sulfate and mercuric nitrate for chloride (see Appendix G,H) |
| Pond stage | Siphon well or staff gauge (see Appendix I) |

# Seasonal Sampling Program and Collection Overview

Monitoring comprises four sub-programs.  See the Field Sampling section for operational details and Table 4 for a complete summary of variables by sampling period.  Refer to Appendix A-K for further details.

##### 1) Sampling of pH and alkalinity three times per year (April, July and October) for all 20

##### ponds. This profiling tracks long-term trends in acidity and acid-neutralizing capacity, both of which can vary seasonally.  (See Appendix H and C for procedures)

Note:  The original monitoring program, begun in 1984 in cooperation with a statewide

survey (Massachusetts Acid Rain Monitoring Project -ARM), included January

sampling. This mid-winter sampling was discontinued after January 2000 because of safety

concerns and because pH and alkalinity data did not differ significantly  (p>0.05) between

January and April.  Other ARM sites are included with the sampling.

2) Spring Survey of all 20 ponds.  During the late March – early April well-mixed period, all 20 ponds are visited to characterize temperature, pH, Eh (redox potential), conductivity, dissolved oxygen, Secchi and light, dissolved nutrients (nitrate, ammonium and orthophosphate) TP, TN, major anions (SO4 and Chloride) and major cations (Ca, Mg, K, Na).  Field observers position a boat above the pond’s deepest point and collect depth profiles for field variables using a YSI.  Samples for laboratory analysis are collected from the same anchored position.  This spring visit can double as an opportunity to collect the spring pH/alkalinity samples (See Appendix H)

##### 3) Summer Biweekly (May through Fall overturn) sampling of the ten “primary” ponds for secchi transparency and for YSI profile measurements - temperature, conductivity, pH, redox potential, dissolved oxygen and depth.  Ten “primary” ponds were selected to represent trophic groups as follows: oligotrophic (Gull, Dyer, Great-Wellfleet, Great-Truro, Spectacle and Long), mesotrophic (Duck, Snow and Ryder) and eutrophic (Herring).  Past monitoring has shown that transparency (e.g. Secchi depth) must be measured at least biweekly to document major temporal change (Table 3).  The objective is to conduct field profiling, particularly of dissolved oxygen, to document the depth and duration of oxygen depletion in deep waters (i.e. hypolimnetic anoxia).  Thus it is essential that this sub-program completely encompass the period from onset of stratification in spring to fall overturn.

- August Survey for all 20 ponds.  For most ponds, stratification, hypolimnetic anoxia and sediment P mobilization are at an annual extreme during late summer; therefore, an annual characterization of the degree of eutrophication of all 20 ponds is performed during this worst-case period. With the exception of major anions, August variables are the same as the spring survey set.

5)  Monthly Pond Stage Pond water levels are monitored at all twenty ponds on a monthly basis to document major seasonal and annual hydrologic changes that could affect water chemistry and biota.  Siphon wells are read at Gull, Duck, Great-Wellfleet, Ryder, Dyer, Long, Herring, Snow and Great-Truro using an electric or chalked tape.  Elsewhere, an auto-level is used to determine pond surface elevation by differential leveling from local benchmarks. (see Appendix I)

## Changes in Secchi transparency (m) versus the seasonal range at CACO kettle ponds  as an indicator of necessary sampling frequency.

1994	                1995

Biweekly Change      		        Biweekly Change

Seasonal 				Seasonal

Pond	   Mean ± SE     Maximum	 Range	Mean ± SE	Maximum      Range

Ryder	1.2 ± 1.2	3.7	8.4	1.1 ± 1.0	3.0	4.5

Duck	0.8 ± 0.9	2.5	2.2	2.2 ± 1.5	4.6	5.7

Great(T)	1.5 ± 0.9	2.4	2.6	0.4 ± 0.5	1.0	1.0

Gull	0.5 ± 0.6	1.8	2.2	0.4 ± 0.2	0.5	1.0

Spectacle	1.7 ± 1.0	2.6	3.7	0.8 ± 0.9	2.3	3.5

## Summary of water quality variables by sampling period - Field Sampling Information

| Parameter to be Monitored | Sample Equipment, Container Sizes, Types and Required Prep | Number of samples or readings | Sample Location and Procedure (Always the deepest point except pH/ALK) | April, July, Oct (all ponds) ( | April Spring Survey (all ponds) X | Summer Biweekly (primary ponds) ( | August Survey (all ponds) ( | Monthly (all ponds) ( |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| pH and alkalinity | 250-ml HDPE bottle, cleaned and rinsed with distilled water | 1 per pond | wade beyond littoral weed beds, rinse the bottle with pond water and sample/cap about 0.5 m below the surface (elbow depth) | ( |  |  |  |  |
| temperature | YSI unit | vertical profile | take readings at 0.1 meters, 0.5 meters, 1 meter and every meter thereafter until 0.5m from bottom |  | X | ( | ( |  |
| pH | YSI unit (calibrate for pH monthly) | vertical profile | take readings at 0.1 meters, 0.5 meters, 1 meter and every meter thereafter until 0.5m from bottom |  | X | ( | ( |  |
| conductivity | YSI unit (calibrate for conductivity monthly) | vertical profile | take readings at 0.1 meters, 0.5 meters, 1 meter and every meter thereafter until 0.5m from bottom |  | X | ( | ( |  |
| redox potential (Eh) | YSI unit (calibrate for redox potential monthly) | vertical profile | take readings at 0.1 meters, 0.5 meters, 1 meter and every meter thereafter until 0.5m from bottom |  | X | ( | ( |  |
| dissolved oxygen | YSI unit (calibrate for DO daily) | vertical profile | take readings at 0.1 meters, 0.5 meters, 1 meter and every meter thereafter until 0.5m from bottom |  | X | ( | ( |  |
| depth | YSI unit (calibrate for depth daily) | vertical profile | take readings at 0.1 meters, 0.5 meters, 1 meter and every meter thereafter until 0.5m from bottom |  | X | ( | ( |  |
| transparency | Licor light meter with two submersible light sensors | vertical profile | take readings at 0.1 meters, 0.5 meters, 1 meter and every meter thereafter until the probe reads 1% |  | X | ( | ( |  |
| Secchi depth | Secchi disk; viewing scope | 1 per pond | take readings at the deepest point in the pond |  | X | ( | ( |  |
| Dissolved Nutrients (nitrates, ammonium, orthophosphate) & cations (Ca, Mg, K, Na) | 50mL clear centrifuge tubes pre-acidified with 0.2 ml of 2 N trace metal grade HCl; filter holders with 0.45 (m filter; peristaltic pump with tubing | 3 per pond | All three samples are pumped from about 0.5 meters through a 0.45(m filter; into clear pre-acidified 50-ml plastic centrifuge tubes (do not rinse bottles); store on ice |  | X |  | ( |  |
| major anions, (chloride & SO4) | 60-ml brown Nalgene bottles cleaned by rinsing 3 times with HCl rinse and then 3 times with distilled water; peristaltic pump with tubing | 3 per pond | All three samples are pumped from about 0.5 meters below surface through a 0.45(m filter into brown 60-ml bottles ; rinse with filtered water prior to filling; store on ice |  | X |  |  |  |
| total nitrogen and total phosphorus | 70-ml test tubes pre-washed with boiling phosphate-free soapy water, triple washed with dH20, triple washed with HCl rinse and triple washed again with dH20; 5mL repipet dispenser with persulfate oxidizing agent (made up of potassium persulfate, boric acid and NaOH); peristaltic pump with tubing | 3 per pond | All three samples are pumped from about 0.5 meters below surface (no filtration) into a pre-washed test tube (etched at the 40ml level); rinse with pond water before sampling; after sample is taken, add 5 ml of oxidizing agent and invert several times to mix; store in rack |  | X |  | ( |  |
| pond water levels | Use siphon well or staff gauge | 1 per pond | Use electric tape for siphon wells. For ponds w/o siphon, 2 people use autolevel. |  |  |  |  | ( |

Figure 1. Flow Diagram of Pond Water Sample Collection. For all lab-analyzed samples, triplicate samples are collected.

# Field Logistics and Safety

Fieldwork is conducted from a canoe or a 10-foot skiff transported to the ponds by truck.  CACO kettle ponds are small (most < 20 ha) and sheltered so that oars are sufficient propulsion.  All personnel are required to wear Personal Flotation Devices (PFDs), or survival suits when surface water temperature is < 18( C. Two technicians are required for equipment operation and safety; thus, no one goes out on the ponds alone.

At the beginning of each field season in late March, each pond’s deepest point is determined by fathometer and marked with an anchored buoy.  These “permanent” buoys must be removed, or their floats submerged, in the fall to avoid being displaced by ice. Technicians should always carry a park radio or cellular phone for emergency use, water and sunscreen.  If  inclement weather approaches, return to shore and seek shelter. (Cape Cod National Seashore SOP #S0014)

# Data management/Reporting/Staffing

Data entry

- Pond profile data (i.e. temperature, conductivity, pH, Eh and dissolved oxygen) are recorded by two different means while in the field: (1) profile data are saved in the YSI 650 Datalogger and (2) the same data are recorded manually on Form 2 (see Appendix J).  Secchi and Light data are recorded manually also on Form 2.

- pH and alkalinity data are recorded on Form 1 (see Appendix J).

Completed data forms are proofed by the field biotechnician and then by the Aquatic Ecologist.

Data for all parameters will be stored in two different places on the Y drive (CACO Server):

- As a working document in the folder titled labusers/Ponds Computer/Working Directory/CCNS Kettle Ponds/CACO_KettlePonds_97.mdb. This file is intended to be updated weekly as data becomes available.

- As archived data in the folder titled dataarchive/Kettle Ponds/PONDS ARCHIVE/Synthesis/CACO_KettlePonds_97.mdb. This file is intended to be updated annually by saving the same file from the directory indicated in part 1. above to the archive database annually, once all data has been entered and validated.

Hard copies of data for all parameters will be stored in the Ponds binder for that year (i.e. 2002)

Reporting and analysis

The Aquatic Ecologist is responsible for data reduction and annual reporting.  Annual reports include: 1) both paper and digital compendia of all collected data; 2) a summary of important trends relevant to management concerns and monitoring objectives; 3) method developments and changes to the protocol; and 4) recommendations for improved monitoring, research and management.

Budget and staffing

The pond field-monitoring program requires one full-time biological technician and a field assistant, typically an SCA or volunteer, from April through October.  Laboratory analysis is conducted by the above biotechnician plus the Laboratory Chemist.  Data analysis and reporting are the responsibility of the Aquatic Ecologist.

# Literature Cited

Caraco, N.F., J.J. Cole, AND G.E. Likens. 1989. Evidence for sulphate controlled phosphorus release for sediments of aquatic systems. Nature. 341: 316-318.

Downing, J.A., AND E. McCauley.  1992.  The nitrogen:phosphorus realtionship in lakes. Limnol. Oceanogr. 37: 936-945.

Great North American Secchi Dip-In, Department of Biological Sciences, Kent State University Kent OH 44242 website: http://dipin.kent.edu

Kelly, C.A., J.W.M. Rudd, AND D.W. Schindler. 1990. Acidification by nitric acid - future       considerations. Water Air Soil Pollut. 50: 49-61.

Martin, L., J.W. Portnoy, AND C. Roman. 1993. Water quality and research plans for Cape Cod National Seashore: Report of a workshop, March 2-3, 1992. NPS/NRWRD/NRTR-93/15.

Nurnberg, G.K.  1995.  Quantifying anoxia in lakes.  Limnol. Oceanogr. 40(6) 1100-1111.

Portnoy, J.W., AND C.N. Farris.  1996.  Revisions to protocol for kettle pond monitoring at Cape Cod National Seashore.

Portnoy, J.W., AND M.A. Soukup. 1990. Gull contributions of phosphorus and nitrogen to a Cape Cod kettle pond. Hydrobiol. 202: 61-69.

Schoenherr, A. 2000.  Seawater as an alkalinity standard.  Cape Cod National Seashore. 5 p.

Winkler, M.G. 1988. Paleolimnology of a Cape Cod kettle pond: Diatoms and reconstructed pH. Ecol. Monogr. 58: 197-214.

# Appendix A. – YSI (temperature, conductivity, redox potential, dissolved oxygen, pH, depth)

The laboratory owns a YSI 6820 Sonde (Serial # 0H932) and a YSI 650 Display )Serial # 02B0008A-B).  The 650 compensates for barometric pressure.  Dissolved oxygen and depth should be calibrated every day.  All other parameters (DO Membrane, pH, conductivity and redox) should be calibrated every two weeks. The sonde can be calibrated either right side up or upside down.  The amount of calibration solution necessary for the process varies between the two methods.  See section 2.6 in the YSI manual for more information regarding the calibration procedure.  YSI can provide annual instrument inspections and factory calibration / cable checks. Be sure to utilize the YSI calibration worksheet (Appendix J) and store this with all monthly calibrations in the laboratory.  Additional YSI information is available at YSI.com, YSI Massachusetts at (508) 748 0366, or the YSI Manual in the laboratory.

Calibration Procedure:

- Connect the probe, 650 and cable.

- Remove the probe guard (if performing pH, Eh, conductivity calibrations) and use the calibration cup for the calibration procedure.

- Turn on the YSI.

- Press “Escape” to get to the main menu.

- Select  “Calibration Mode” and press “Enter”

- A screen allowing you to choose which parameter to calibrate will appear.

- Fill a bucket with room temperature tap water to rinse the sonde between calibration solutions. Rinse the probe being calibrated with deionized (DI) water between calibrations.

- Utilize Calibration Worksheet (Appendix J.) during process

Conductivity

- Conductivity should be calibrated first.  Avoid contamination of conductivity standard.  Never use less than 1.0mS/cm Conductivity Standard.

- Rinse the probe and calibration cup with distilled water before starting the calibration.

- For freshwater pond sampling use the 1,000 uS/cm (microsiemens/cm) conductivity standard solution.  This is equal to 1 mS/cm (millisiemens/cm).  The Sonde requires input in millisiemens.

- Choose “Conductivity” and then “SpCond” in order to calibrate for specific conductivity.

- Add a small amount of calibration solution to the calibration cup to rinse the probe sensors and cup, then discard.

- Fill the calibration cup with the conductivity standard (~ about 200 ml) making sure that the standard is covering the entire conductivity bridge (ensure no air bubbles are present in conductivity hole) and screw the cup onto the probe.

- Enter the conductivity value of the calibrator when prompted.

- Allow the readings to stabilize for about 30 seconds and then press “Enter” to accept the calibration.

- The YSI will indicate if the calibration was successful.  Check the Conductivity Cell Constant (Advanced Menu) under Cal Constants.  The acceptable range is 5.0  +/-0.45.  Contamination is likely if it is out of this range.  If the sonde reports “Out Of Range”, investigate.  Never override a calibration without understanding  the cause.  A typical error is entering 1000 microsiemens instead of 1.0 millisiemens.

- Press any key as prompted, hit “Escape” twice to get back to the calibration menu.

pH

- Go to the sonde’s report menu and turn on the pH mv output.  This will allow the sonde to display the millivolts (probes raw output), in addition to the pH units during the calibration process.

- Recondition the probe if a slow response in the field has been reported.  The procedure can be found in the manual under the “Sonde Care and Maintenance Section”.

- First, rinse the calibration cup with DI water.

- Choose “ISE1 pH” from the menu and then “2 Point” calibration.

- Use pH 7 standard to calibrate for the first point.

- Ensure that the temperature probe is in solution with the standard, record pH millivolts at each calibration point.

- Add a small amount of calibration solution to the calibration cup and rinse the probe sensors and calibration cup.

- Discard the solution and refill the cup with the standard. You will need about 100 mL to completely immerse the pH sensor when the probe is held upright.

- Screw the cup onto the end of the probe.

- Input the calibration value, press “Enter”.

- Allow the readings to stabilize for about 30 seconds and then press “Enter” to accept the calibration.

- The YSI will indicate if the calibration was successful. Press any key as prompted.

- Press “Escape” to get to the second point of the pH calibration.

- Repeat the same process using pH 4 standard.

- After successfully calibrating for pH 4 standard, press any key as prompted.

The millivolts indicates the present status of the probe.  A good set of numbers to use are as follows.

BUFFER 4   =    +180   +/- 50mv

BUFFER 7   =    0         +/- 50mv

BUFFER 10 =  -180      +/- 50mv

The ideal numbers when a probe is new are close to the 0 and 180 numbers.  The numbers will shift to the higher side of the tolerance as the probe ages.

- After recording the pH millivolts for the calibration points, you must determine the slope of the sensor.  This is the difference between the two calibration points that were used.  For example, if you record +3mv for buffer 7 and –177 for buffer 10, then the slope is 180.  The acceptable slope range is 165 to 180.  Once the slope drops below span of 165, the sensor should be taken out of service.

- Press “Escape” twice to return to the calibration menu.

- Never override a calibration without understanding  the cause.  Proper storage of the sensor will greatly extend the probe life.

Redox (ORP or Eh)

- Rinse the probe and calibration cup with distilled water before starting calibration.  Dry the calibration cup and probe sensors as much as possible before adding the calibration standard.

- Use the YSI Zobell Solution (Toxic, Handle with caution! Use and discard per MSDS) to calibrate for redox.  Once reconstituted, the solution has only a six-month shelf life.

- Fill the calibration cup (~ 250 ml) with the solution and screw the cup onto the YSI.  The Redox (ORP or Eh) sensor shares the pH sensor reference so it must be in solution when the 6000 is being calibrated.

- To determine the redox value of the standard, use the Ag/AgCl formula given on the sheet included with the Zobell Solution, or see “Principles of Operation” section of the manual. Air temperature is needed to calculate the redox value.  Enter the number in millivolts that is closest to the measured solution temperature of the sonde.

- Choose “ISE2 Orp” from the calibration menu

- Enter the redox value of the standard when prompted and press “Enter.”

- Allow the readings to stabilize for about 30 seconds and then press “Enter” to accept the calibration.

- The YSI will indicate if the calibration was successful.

- Press any key as prompted; hit “Escape” once to get back to the calibration menu.

Dissolved Oxygen

- DAILY CALIBRATION:

- Rinse the probe and calibration cup with water.

- Use a kimwipe to remove all liquid from the dissolved oxygen membrane.

- Add enough water to cover the bottom of the calibration cup.  Make sure the probe membrane does not touch the water.

- Put the probe into the calibration cup, but do not screw the cup onto the probe.

- Allow the air within the cup to become fully saturated by letting the probe sit in the calibration cup for ten minutes.

- Select “Dissolved Oxy” from the calibration menu and then “DO%” from the next menu.

- Allow the readings to stabilize for about 30 seconds and then press “Enter” to accept the calibration.

- The YSI will indicate if the calibration was successful.

- When calibration is complete, go to the sondes “Advanced Menu” and then to the “Cal Constants” and record the “DO Gain”.  The Gain should be 1.0 with a range of –0.3 to +0.5.  The probe should now be successfully calibrated and ready to go.  Investigate any warning messages.

- Press any key as prompted, hit “Escape” twice to get back to the calibration menu.

- 3 WEEK CALIBRATION AND MEMBRANE CHANGE:

- Replace the DO Membrane using the 6035 Probe Reconditioning Kit and Directions.

- DO membranes will be slightly unstable for 3-6 hours after installation.  It is recommended that final calibration of DO sensor take place after replacing the membrane.  (The Sonde should be in a saturated evironment).

- Go to the sonde’s “Report” menu and enable the “DO Charge”

- Now go to the “Run” menu and start the sonde in “Discrete Run” mode at a 4 second rate.  Allow the sonde to run for 10 minutes.

- Record the DO Charge after 5 minutes.  It should be 50  +/- 25.

- Once again, start the probe in “Discrete Run” mode at a 4 second rate and record the first 10 DO% numbers on paper.  The numbers should start high and drop with each 4 second reading (eg) 110, 105, 102, 101.5, 101.1, 101.0…etc..It does not matter that exactly 100% is not reached, only that the high to low trend exists.

Note: Initial power-up can corrupt the first two DO % samples; disregard.  Now the probe can be calibrated as explained above for daily calibration.

Depth

- Select “Pressure-Abs” from the calibration menu.

- Enter “0.0” at the prompt and then press “Enter”.

- Allow the readings to stabilize for about 30 seconds and then press “Enter” to accept the calibration.

- The YSI will indicate if the calibration was successful.

- Hit “escape” until the prompt has returned to the main menu.

YSI Calibration information adapted from M. Lizotte, YSI MA, May 2002.

# Appendix B. – Transparency (Secchi and Li-cor Measurements)

- Secchi Measurements

The laboratory owns numerous secchi disks.  It is vital to examine the ponds secchi disk to ensure the line is correctly marked in meter increments.  Secchi depth (SD) should be measured first so as to ascertain the depth for subsequent YSI measurements and minimizing the chance of the YSI sonde disturbing the bottom sediments which will adversely affect readings.

- Li-cor Measurements for Light Attenuation

The laboratory owns two Li-cor underwater sensors : (serial number – UWQ4539 and serial number UWQ6339).  These sensors must be calibrated every two years.  In the past the light sensors have been calibrated every other spring before the April sampling season has begun.  It will take at least a full month for the calibration process to be completed, so be sure to plan ahead.  If there are any questions, consult the aquatic ecologist. Also, all calibration information should be given to the aquatic ecologist.  The calibration needs to be scheduled with Li-cor.  Call the company at (402) 467-3576 to receive an RMA number.  Then box up the sensors and send them to the company.

The datalogger does not need to be calibrated.  However, before the season begins it is a good idea to check that the Datalogger has been programmed with the correct multipliers.   Each light sensor has a unique multiplier that is stored with Li-cor information in the lab.

Li-cor setup:

- Connect a sensor to an underwater cable by pushing the two-pronged connector of the sensor itself into the cable, aligning the marks on the cable and sensor. If the marks are opposite you will get negative (but accurate) readings. The cable can then be connected to the logger at one of the bayonet mount terminals on the top of the logger. The terminal to the left is for "Channel 1" or “surface” and the one on the right is for "Channel 2" or deep profiling sensor.

- Turn on the LI-1000 by pushing the "FCT ON" button firmly. Instantaneous mode is the correct mode for CACO operations.  Before taking readings channels MUST be configured. Once started you must complete the configuration procedure before doing anything else; the logger will not let you do otherwise.

- Configurations are done by channel, starting with Channel 1. Push the CFG button and the display will read "MODE IS INST", meaning that instantaneous mode has been selected.  Press ENTER to accept (or up or down arrow to toggle to LOG). A message saying "ch1 is OFF" (or LIGHT or GEN or THERM) appears.  Using the up or down arrows to toggle to LIGHT and press ENTER. The next display says "range = A". This means the machine's sensitivity range will be automatically adjusted for different light intensities. Press ENTER to accept, and don't be concerned with other choices. The next display is "mult = xxx.x" where the number is a calibration coefficient. Each sensor has a small plastic card listing its factory-calibrated coefficient. This number differs for measurements of light flux in air or water. Choose the proper one and enter the number, including the minus sign first, from the keypad, then press ENTER. The next display is "label=XX" where XX is simply 2 characters chosen to denote the particular channel or type of measurement in progress. The 2 characters you choose will be displayed in data readouts. Note that letters must be selected by using the up or down arrow then pressing the appropriate button, e.g., up arrow-9 gets you "Q". Press ENTER to accept the characters  chosen. The next display is "ave=X" where X is some number. Any number can be entered; this denotes the period in seconds over which the logger will take a moving average of light flux. This feature allows the LI-1000 to average short scale fluctuations in light intensity, such as from moving clouds, but requires patience to get an accurate reading. Ten seems to work well. Press ENTER to accept the number chosen. This completes configuration of channel 1. This process may be repeated for channel 2 if two sensors are in use (and likewise for additional sensors hard-wired in on the pin connectors), or simply toggle each channel OFF and press ENTER until past channel channels 8 and M (the latter used for math functions). The next display will be 1X and a decimal number, which gives readings in microeinsteins (=micromoles of photons) m-2 sec-1. Configuration is now complete. The logger can be turned off by pushing OFF; it also will turn itself off after a 15 minutes with no key activity to save battery life.

- Measurements of light flux can now be taken by reading the display after a suitable time has elapsed (bearing in mind the averaging time) and/or by hitting ENTER whenever storing a reading in memory.

- Use of the memory to store data requires that time and date be set before beginning (also important for use of log mode). To do this, push TIME. The display will read a preset date. If it is incorrect, simply enter two numbers for year (YY), two for month (MM) and two for day (DD), then press ENTER. If a mistake is made, push the back arrow to delete the wrong character and enter the right one. Push ENTER to accept the date. The next display will give military time in hours and minutes. If this is not right, enter the correct numbers as above, then push ENTER, which will return to the main display.

# Notes & Troubleshooting

- On sunny, windy days the light readings will fluctuate. There are a few things you can do to reduce the amount of fluctuation.  First, adjust the range of the submerged probe as it’s moved through the water column.  Pages 4-1 and 4-2 in the LI-1000 Datalogger Instruction Manual explain the range settings in the instantaneous mode (the mode which should be used to collect data).  The autorange (A) setting covers the widest input signal range but tends to fluctuate rapidly when underwater radiation sensors are used.  Therefore, adjusting the range as the probe is lowered through the water column can decrease the amount of fluctuation.

- To change the range settings:

- Press the “CFG” button.

- Use the “ENTER” button to scroll down to the channel, typically channel 2 (the submerged probe), which will be changed.

- Press any button to scroll through the range values.

- Choose a range value by hitting “CFG”.

- Another way to stabilize readings is to adjust the average value; the time interval over which values are averaged (“ave =”).

Adapted from:

http://people.uncw.edu/cahoon/LiCor.htm

# Appendix C. – pH and Alkalinity titration (potentiometric, low-level)

Surface samples (0.5 m depth) are collected April, July and October by wading out beyond weed beds, rinsing a 250 mL HDPE bottle with pond water, and filling the bottle to the top, i.e. no air in head space.  Samples are kept cold in a cooler, or refrigerated, and titrated within 24 hours of collection.

Stock seawater standard:

Collect and filter (0.45 (m) seawater from Coast Guard Beach, Truro.

Dilute standard:

Pipet exactly 10.0 ml of the above stock into a 1-L volumetric flask and fill to the mark with distilled water.  The alkalinity of this solution should be about 0.0232 mEq/L = 1.16 mg CaCO3/L (Schoenherr 2000).

pH & titration procedure (All solutions at room temperature):

1. Calibrate pH meter and electrode pair in pH 7 and pH 4 buffers with stirring.(medium speed)

2. Measure 100 ml of the standard into a 150-ml beaker.  Add a small stir bar.

3. Turn stirrer speed to medium.

4. Place pH and reference electrodes in the standard and record pH after 3 min or sooner if pH is stable, i.e. no change > 0.05 pH units for 30 sec.

5. Submerse the tip of the Dosimat dispenser, containing 0.1N HCl (certified 1N  diluted 1:10), into the solution.  Set mode to DIS C and delta volume to 0.005 ml.

6. Add titrant slowly to reduce pH to 4.50 (see notes below).

7. Record titrant volume to pH 4.5 (quantity A).

8. Again add titrant until pH reads exactly 4.20.

9. Record total titrant volume (quantity B).

10. Alkalinity (mg CaCO3 L-1) is equal to:

(2A-B) X 0.1 X 50,000/100

11. Compare your results with previous calibration data for the standard.  If they differ more than

0.5 pH units and/or more than 0.3 alkalinity units, probe (or, less likely, meter) maintenance is

indicated.

Notes:

If you overshoot the 4.50 pH by a few hundredths (no more!) don't abandon the titration.  Simply record the pH (e.g. 4.48) and the titrant volume used and then titrate exactly 0.3 pH units lower (e.g. 4.18) and record total titrant volume.  The formula still works.

For ponds with pH below 4.50, record the ambient pH as usual and then titrate to exactly 0.3 pH units lower.  Alkalinity is computed as usual although quantity A is zero.

# Appendix D. – Dissolved Nutrients Analysis (Ammonium, Nitrate/Nitrite and Orthophosphate)

Sample Handling for Dissolved Nutrients

Samples taken for dissolved nutrient analysis are also used for cation analysis.

For cation and dissolved nutrient samples, direct the GeoPump tube discharge through a 47-mm filter holder with 0.45-(m membrane filter (Millipore HA).  Allow the first few milliliters to discharge to waste and then direct the flow into a pre-acidified (0.2 ml of 2N trace metal grade HCl) 50-ml clear freestanding HDPE centrifuge tube.  Labeled samples are stored on ice or refrigerated.

- Ammonium (Phenolate method)

(Adapted from Lachat QuickChem Method 10-107-06-1-C)

Note: EDTA is added to the sample in-line in order to prevent precipitation of calcium and magnesium ions.

Principle

This method is based on the Berthelot reaction. Ammonia reacts with alkaline phenol, then with sodium hypochlorite to form indophenol blue. Sodium nitroprusside is added to enhance sensitivity. The absorbance of the reaction product is measured at 630 nm and is directly proportional to the original ammonia concentration.

See Lachat QuickChem Method 10-107-06-1-C for preparation of reagents, standards, timing specifications, manifold diagram, and system notes.

- Nitrate/Nitrite (Cadmium reduction method)

(Adapted from Lachat QuickChem method #31-107-04-1-C)

Note: Samples containing high concentrations of iron, copper or other metals may give low results. EDTA is added to the buffer to reduce this interference.

Principle

Nitrate is quantitatively reduced to nitrite by passage of the sample through a copperized cadmium column. The nitrite (reduced nitrate plus original nitrite) is then determined by diazotizing with sulfanilamide followed by coupling with N-(1-naphthyl)ethylenediamine dihydrochloride. The resulting solution has a magenta color, which is detectable at 520 nm.

See Lachat QuickChem Method 31-107-04-1-C for preparation of standards, reagents, packing of cadmium column, timing specifications, manifold diagram, and system notes.

- OrthoPhosphates

(Adapted from Lachat QuickChem method #31-115-01-1-G)

Principle

Ammonium molybdate and antimony potassium tartrate react in an acid medium with phosphate to form an antimony-phospho-molybdate complex.  This complex is reduced to an intensely blue-colored complex by ascorbic acid.  The color produced is proportional to the phosphate concentration in the sample.  Though there is a density difference between seawater and reagent water the bias is less than 2%.

Though the method is written for seawater and brackish water it is also applicable to non-saline sample matrixes.  The method is calibrated using standards prepared in deionized water.  Once calibrated, samples of varying salinities (0 to 35 ppt) may be analyzed.  The determination of background absorbance is necessary only for samples, which have color absorbing at 880 nm.

See Lachat QuickChem Method 31-115-01-1-G for preparation of standards, reagents, timing specifications, manifold diagram, and system notes.

##### Special Apparatus

Heater Module (175 cm of 0.8 mm i.d. Teflon tubing).

# Appendix E. - Total Phosphorus and Nitrogen: Acid Persulfate Digestion

(Adapted by K. Lee from Lachat QuickChem method #’s 31-115-01-1-F and 31-107-04-1-B or C)

Note: This method is extremely sensitive to pH in the sample. The higher the acidity, the less sensitive the method.

Principle

All forms of phosphorus are converted to orthophosphate. The orthophosphate ion (PO4-) reacts with ammonium molybdate and antimony potassium tartrate under acidic conditions to form a blue complex which absorbs light at 880nm.

##### Sample Handling

Wash all sample test tubes (70 mL etched glass tubes) with HOT phosphate free soapy water and triplicate flush with dH2O. Acid wash the tubes and triplicate rinse with DI H2O. After drying, store sample tubes capped until day of sample collection. Collect triplicate 40 mL samples and immediately treat with 5ml of persulfate oxidizing reagent (50 g Aldrich low nitrogen Potassium Persulfate, recrystallized; 30g Fisher Boric Acid and 350 ml of Fisher1M NaOH [40g/L]).  Cap tubes securely.  Invert twice to mix.  Record reagent date on each tube and run a blank.

Sample Digestion

Place samples in pressure cooker and process at 15 psi for approximately 50 minutes on the day of collection. Allow sample tubes to cool and store upright at room temperature until analyzed. Holding time for samples should be limited to one month (28 days) if at all possible. (Ameel, 1993 and Lachat Instruments method 10-115-01-1-F)

##### Method

Introduction

The method described here is an adaptation of the Total Dissolved and Total Particulate Nitrogen and Phosphorus method by Barbara L. Nowicki, published in the Manual of Biological and Geochemical Techniques in Coastal Areas, MERL Series, Report No. 1, 3rd edition.  This method is performed on a whole water sample utilizing an alkaline persulfate digestion in order to oxidize both organic nitrogen and ammonia to nitrate. Total phosphorus can also be determined from the same digestion.  The samples are analyzed on a multi-channel Quikchem FIA + 8000 Series autoanalyzer (Lachat Instruments, Inc., Milwaukee, WI).

For simultaneous oxidation of total nitrogen (TN) and total phosphorus (TP) it is necessary to utilize an alkaline medium for nitrate production and an acidic medium for the oxidation of phosphorus.  This simultaneous oxidation method starts at a pH of approximately 10 and ends at a pH of approximately 3-4.  These conditions may be obtained from a sodium hydroxide-boric acid system.  The digestion time is a function of the autodecomposition of peroxodisulfate and temperatures must reach at least 110(C in the pressure cooker to achieve this decomposition (Valderrama, 1981).

Equipment

- Digestion tubes are 70 mL Pyrex #9825 borosilicate glass screw-top tubes with autoclavable rubber-lined screw caps (Fisher Scientific, catalog #15-327). Sample caps and tubes are washed thoroughly with HOT soapy water (Citranox; phosphate free acidic detergent), distilled water rinse to remove all soapy residue triplicate acid rinse with 1:4 HCl solution triplicate distilled water rinse, air dried and capped until used.

- Hot plate

- Pressure cooker (maximum pressure ~20psi)

- Orion pH meter

- Quikchem FIA + 8000 Series Lachat autoanalyzer with nitrate/nitrite and orthophosphate channels

- Eppendorf pipettes and tips

Reagent Preparation

Note: Use deionized water for all solutions.

Reagent 1: Stock Ammonium Molybdate Solution: Dissolve 40.0 g. ammonium molybdate tetrahydrate in approximately 800 ml of dH2O.  Dilute to the mark in a 1L volumetric flask and invert to mix. Store in plastic and refrigerate.

Reagent 2: Stock Antimony Potassium Tartrate Solution: Dissolve 3.0 g. antimony potassium tartrate hemihydrate (K(SbO)C4H4O6·1/2H2O) in approximately 800 ml of dH2O. Dilute to the mark in a 1L volumetric flask and invert to mix. Store in a dark bottle and refrigerate.

Reagent 3: Molybdate Color Reagent: Slowly add 35.0 ml concentrated TMG sulfuric acid to about 500 ml dH2O. Mix and allow to cool. Add 72.0 ml of Reagent 2 and 213.0 ml of Reagent 1. Dilute to the mark in a one-liter volumetric flask and invert to mix.

Reagent 4: Ascorbic Acid Reducing Solution: Dissolve 60.0 g. ascorbic acid in about 700 ml dH2O. Dilute to the mark in a 1L volumetric flask and invert to mix. Prepare this reagent weekly.

Reagent 5: Sodium Hydroxide-EDTA Rinse: Dissolve 65 g. sodium hydroxide and 6 g. tetrasodium ethylene diamine tetraacetic acid in 1L of dH2O.

Reagent 6: Carrier Solution: The carrier solution is distilled water.

Standard Preparation

The linear range of standards found to be appropriate for analysis of natural water samples by this method are 1-100(M NO3 and 0-10(M PO4.   Inorganic standards (the same standards used in the dissolved nutrients method) are used for the standard calibration. An organic standard (B) is prepared as follows and digested exactly as sample/standard procedure outlined in the next section and is analyzed as an unknown.

NOTE:  Seashore vernal pool samples will likely exceed the range for NO3 and should be diluted accordingly.  All other Cape Cod National Seashore natural water samples are expected to fall within these standard ranges for TP and TN.

Preparation of Inorganic Standards:

The standard used for Total Phosphorus/Total Nitrogen is the same standard used in the Dissolved Nutrients Analysis.  It consists of one standard (Standard A). The Lachat Total Phosphorus/Total Nitrogen calibration method should be programmed to dilute this Standard A for determining the standard curve. This program is set so that the first standard uses no dilution and the last standard is the blank.  Make sure at least 4 test tubes in the first column, (first positions of Tray 4) are ready for dilution.  The total points for the standard curve will be 6.

To make Standard A, a stock standard solution is prepared for each nutrient and a given amount of each of the three stock solutions are combined in distilled water for the desired resulting concentration.  Below is the procedure for preparing each of the stock solutions along with Standard A.

Stock Standard Solution for NH3-H+ 			Stock Solution, 10mM

In a 1 L volumetric flask dissolve 0.535 g of ammonium chloride (NH4Cl) that has been

dried for two hours at 110(C in about 800 mL DI water. Dilute to the mark with DI water

and invert to mix. Store refrigerated in a properly labeled bottle for up to six months

.

Stock Standard Solution for PO4				Stock Solution, 10 mM

In a 1 L volumetric flask dissolve 1.361g primary standard grade anhydrous potassium dihydrogen phosphate (KH2PO4) that has been dried for two hours at 110(C in about 800 mL DI water.  Dilute to the mark with DI water and invert to mix. Store refrigerated in a properly labeled bottle for up to six months.

Stock Standard Solution for NO3 				Stock Solution, 10 mM

In a 1 L volumetric flask dissolve 0.850g of sodium nitrate (NaNO3) that has been dried for one hour at 60(C in about 800 mL water. Dilute to the mark with DI water and invert to mix.   Store refrigerated in a properly labeled bottle for up to six months.

STANDARD A

In a 100 mL volumetric flask, add the following to approximately 80 mL of distilled water:

| To 80 mL of DI H20 in a volumetric flask add: | Resulting Concentration |
| --- | --- |
| 1000 μL NH4 Stock Standard | 1400 ppb |
| 100 μL PO4 Stock Standard | 310 ppb |
| 500 μL NO3 Stock Standard | 700 ppb |

Organic Standard for analyses as an unknown:

Nitrogen Organic Standard:	Urea (NH2CONH2)	M.W. = 60.06	(Fisher certified A.C.S.)

10mM stock solution:	0.060g/100mL dH2O

Standard	(M		mL of 10mM stock/100mL	mg/L N

B		 40			0.4			0.56

Phosphorus Organic Standard: D-fructose 1,6-diphosphate (C6H11O12P2Na3)

M.W. = 406.1		(Sigma grade)

1mM stock solution:	0.0413 g/100mL dH2O

Standard	(M			mL of 1mM stock/100mL	mg/L P

B		 5			0.5				0.155

Percent recovery of organic standards is checked against the recovery of inorganic nitrate and phosphate standards.

Sample/Standard Procedure and Storage

Recrystallization of Potassium Persulfate:

Acid clean all glassware and dry in oven. Dissolve 48 g. of potassium persulfate (Aldrich, low nitrogen) in 300 mL of distilled water in a 1000 mL Erlenmeyer flask.  Heat the solution with stirring and swirling to 65(C at which point the persulfate dissolves.  Continue heating to 75(C.  Place immediately in ice-water bath for approximately 45 minutes while crystals form.  Pour through acid-cleaned, Distilled water-rinsed Buchner funnel with No. 40 Whatman filter paper.  Apply 5 psi vacuum. Rinse and scrape remaining crystals from flask into funnel.  Dry and fluff crystals in funnel for 5 minutes.  Put in petri dish and dry in dessicator.

Persulfate Oxidizing Reagent:

50g Potassium persulfate (low nitrogen) (Aldrich)

30g Boric Acid (Fisher)

350mL  1M NaOH (40g/L)  (Fisher)

Dissolve above reagents in a 1-liter volumetric flask (acid and dH2O rinsed!) and bring to volume.  Transfer to amber 5mL-repipet dispenser.  Holding time of oxidizing reagent not to exceed one week, although it is recommended that new reagent be made daily.

Sample/Standard Digestion:

Forty milliliters (40mL) of organic standard B or whole water sample is collected directly into the 70 mL glass tubes. Five milliliters (5mL) of persulfate oxidizing reagent is added immediately following sample collection to each tube and tubes are capped. Sample/standard tubes are cooked for approximately 50 minutes at 15psi in pressure cooker.  Tubes are allowed to cool to room temperature and stored upright until analysis is performed.  Samples should be analyzed as soon as possible, with a maximum holding time not longer than 28 days following digestion (Ameel, 1993 and Lachat Instruments method 10-115-01-1-F).

Sample analysis

Simultaneous analysis of TP and TN is performed on a Quickchem 8000+ Series autoanalyzer (Lachat Instruments, Inc., Milwaukee, WI).  TN is determined by cadmium reduction using Lachat Instruments method 31-107-04-1-B or C, based on U.S. EPA Method 353.2.  TP is measured as orthophosphate by automated ascorbic acid method on the Quickchem IV autoanalyzer (Lachat Instruments method 31-115-01-1-F, based on U.S. EPA Method 365.3).

NOTE:  Carrier solution should be prepared to match sample/standard matrix (i.e. add an appropriate amount of TMG H2SO4 to match pH of samples).

REFERENCES

Nowicki, B.L.  1994.  Total Dissolved and Total Particulate Nitrogen and Phosphorus.  MERL Series, Report No. 1, 3rd Edition, pages  88-91, 1994.

Ameel, J.J., R.P Axler, and C.J. Owen,  1993.  Persulfate digestion for determination of total nitrogen and phosphorus in low-nutrient waters.  American Environmental Laboratory.  1993.

Valderrama, J.C.  1981.  The simultaneous analysis of total nitrogen and total phosphorus in natural waters.  Marine Chemistry, 10:109-122.

American Public Health Association.  Standard methods for the examination of water and wastewater, 19th edition.  Eaton, AD, ed. Washington DC, 1995.

# Appendix F – Cations (Ca, Mg, K, Na)

Sample Handling for Cations

Samples taken for cation analysis are also used for dissolved nutrient analysis.

For cation and dissolved nutrient samples, direct the tube discharge through a 47-mm filter holder with 0.45-(m membrane filter (Millipore HA).  Allow the first few milliliters to discharge to waste and then direct the flow into a pre-acidified (0.2 ml of 2N trace metal grade HCl) 50-ml clear HDPE centrifuge tube.  Samples are stored on ice or refrigerated.

Direct Air-Acetylene Flame Atomic Absorption Method

Principle

In flame atomic absorption spectroscopy a liquid sample is aspirated and mixed as an aerosol with combustible gases (acetylene and air). The mixture is ignited in a flame of temperature ranging from 2100 c to 2800 c (depending on the fuel gas used.) During combustion, atoms of the element of interest in the sample are reduced to the atomic state. A light beam from a lamp whose cathode is made of the element of interest is passed through the flame into a monochronometer and detector. Free, unexcited ground state atoms of the element absorb light at characteristic wavelengths; this reduction of the light energy at the analytical wavelength is a measure of the amount of the element in the sample.

(Follow procedures outlined in the Cape Cod National Seashore, North Atlantic Coastal Lab Method A-001 )

# Appendix G – Anions (Cl- and SO4-)

Sample Handling for Anions

Filter samples (Millipore type HA 0.45-(M filters) directly into brown, 60-ml polyethylene, distilled water-rinsed bottles.  Refrigerate samples until analyzed.

Principle
Sulfate

Sulfate ion (SO4-2) is precipitated in an acetic acid medium with barium chloride so as to form barium sulfate crystals of uniform size. Light absorbance of the barium sulfate suspension is measured by a photometer and the SO4-2 concentration is determined by comparison of the reading with a standard curve.

Refer to method E Turbidimetric Method 4500-SO4-2. in Standard Methods for the Examination of Water and Wastewater, 20th Edition

Chloride

Chloride can be titrated with mercuric nitrate because of the formation of soluble, slightly dissociated mercuric chloride.  In the pH range 2.3 to 2.8, diphenylcarbazone indicates the titration end point by formation of a purple complex with the excess mercuric ions.  Xylene cyanol FF serves as a pH indicator and end-point enhancer.  Increasing the strength of the titrant and modifying the indicator mixtures extend the range of measurable chloride concentrations.

Refer to method C. Mercuric Nitrate Method 4500-Cl- in Standard Methods for the Examination of Water and Wastewater, 20th Edition

.

# Appendix H - Sampling Procedures for the Pond Monitoring Program

- Biweekly Pond Sampling Procedure

Equipment List for Biweekly Sampling

Check here

| Truck |  |
| --- | --- |
| Boat |  |
| Oars/Paddles |  |
| Radio |  |
| 2 Anchors |  |
| Fathometer |  |
| Clipboard |  |
| Pencils |  |
| Pond Field sheets copied onto “Write in the Rain” paper (Profile) |  |
| YSI (Four parts – 650, cable, probe and sensor guard) |  |
| Secchi disk |  |
| Viewing Scope |  |
| Light meter (Three parts – Datalogger, long cable with attached light sensor for variable depth readings and short cable with attached sensor for 10 cm readings) |  |

Before Sampling

- Before sampling, the YSI should be calibrated.  DO and depth must be calibrated

each day.  See Appendix A. for more information.

- Make sure the YSI is properly charged.  Once fully charged the battery will run for 6-8

hours.  When recharging make sure that the charger is on the positive setting.

- Copy enough pond field sheets onto “write in the rain paper”.

(   Put the YSI together in the laboratory before going out into the field.  Replace the     calibration cup with the probe guard.  Be careful when doing this because the DO

membrane is not protected and can easily be nicked.  To keep the DO membrane wet during transport, wrap a wet towel around the probe and place the entire probe upright into an empty bucket.

- Load equipment and boat into the truck.  Check twice to make sure nothing has been forgotten.

- Some sampling sites require you to walk across private property. Be aware of the private property owners and their personalities especially during July and August.  Shoreline residents at Duck, Ryder, Herring and Southeast, Great (W) ponds support NPS monitoring but are especially sensitive to trespass.

- Paddle out to the marker buoy or find the deepest point in the pond.  Use the anchor in conjunction with the buoy line to situate the canoe over the deep point.  If windy, do not tie the boat to the buoy.  Instead paddle up-wind of the buoy about 20 m, drop the bow anchor (the heaviest anchor) and allow the boat to float back to the buoy.  Cleat off the front anchor and drop the back anchor.  This should keep the boat from dragging anchor or drifting with the wind.  In high winds let out more anchor line (i.e. increase scope) to hold the bottom.

Data Collection – Transparency

Note: Data collection should take place as close to the deepest point as possible.

Secchi Disk

- Lower the Secchi disk over the sunny side of the boat.  Using the viewing scope, record the depth at which the disk cannot be seen by the naked eye.  The disk usually has to be raised and lowered a couple of times close to the point where the disk disappears in order to determine the Secchi depth.  Remove sunglasses when testing the Secchi depth. Have helper repeat this procedure and record the average of the two readings.

Light Meter

- Set up the light meter. The cords connecting the cable to the datalogger are color-coded.  The blue cord is the surface sensor. Take the protective red covers off of the two light sensors (cells).  Try to keep the datalogger out of direct sunlight.

- Turn the light meter on (press the “FCT ON” button) and switch to “MI” by pressing the “Chan” button twice.  “MI” is the math channel.  It calculates the ratio between values received from the two light cells.  Channel 1 displays the light readings from the water surface, while the second channel 2 displays the values from the profiling cell.

- Position the two light probes next to each other in the boat and record this measurement, so that their relative sensitivity can be assessed.

- Next, lower the surface cell and the profiling cell to 0.1m and take a reading.  Avoid shadows (especially from the boat as the probe is lowered through the water column) which will result in incorrect light measurements.

- Take readings at 0.5 meters, 1 meter and every meter thereafter until the probe reads one percent.

Data Collection –YSI (temperature, pH, conductivity, redox potential, dissolved oxygen, depth)

- Record DO/Temperature in saturated towel conditions before deploying sonde.

- Unwrap the YSI probe from the towel and turn it on.  Try to keep the 650 Display out of direct sunlight.

- Place the probe in the water and allow it to equilibrate to the pond conditions.  This usually takes 5 mins.

- Take readings while out of the water and then all profile data at 0.1 meters, 0.5 meters, 1 meter and every meter thereafter until about 0.5 m off of the sediment avoiding contact with the sediment if possible.

- Before taking each reading, make sure the values have stabilized.  Dissolved oxygen is usually the one parameter that takes the longest to stabilize, especially when passing through the thermocline. Allow 60 seconds at each depth interval for meter stabilization.

- Record readings in the YSI datalogger and on the field sheets (see Appendix I, Form 2)

After Sampling

- Upon returning to the lab, properly store all equipment.

- Disconnect the YSI Sonde from the cable and 650 datalogger.  Fill the calibration cup with a small amount of water and replace the cup onto the end of the probe.  Replace the plastic protectors on the cable ends and the top of the Sonde.

(	Download the YSI 650 data file to the ECOwin program at the ponds computer.

- Enter the data collected that day into the working directory. CACO_KettlePonds_97.mdb file.

Division of labor

CACO found the most efficient division of labor was to have one person working with the YSI/ Light Meter in combination.  The second person was in charge of recording the YSI profile values as they are read aloud by the first person.

- April and August Sampling Procedures

NOTE: Major Anions (Cl- and SO4-) samples are only collected in April.

#### Equipment List for April and August Sampling

#### Check here

| Truck |  |
| --- | --- |
| Boat |  |
| Oars/Paddles |  |
| Survival suits; if water 40 c |  |
| 2 Anchors |  |
| Fathometer |  |
| Clipboard |  |
| Pencils |  |
| Marker |  |
| Labeling tape |  |
| Pond Field sheets copied onto “Write in the Rain” paper (Profile) |  |
| Cooler |  |
| Icepacks |  |
| Peristaltic pump with tubing |  |
| YSI (Four parts – 650 datalogger, cable, probe and sensor guard) |  |
| Secchi disk |  |
| Viewing Scope |  |
| Light meter (Three parts – datalogger, long cable with attached light sensor for variable depth readings and short cable with attached sensor for 10 cm readings) |  |
| 60-mL plastic syringe |  |
| Clean beaker |  |
| Filter holders |  |
| Luer-lock/threaded stainless steel adapter |  |
| Millipore filters (type HA 0.45-(m) |  |
| 50-ml pre-acidified clear freestanding HDPE centrifuge tubes (3 per pond) |  |
| 60-ml amber HDPE bottles (3 per pond) without acid (FOR APRIL SAMPLING ONLY) |  |
| 70-ml glass test tubes (3 per pond) |  |
| Potassium persulfate oxidizing reagent (optional – this may be added immediately upon returning to lab) |  |
| Bottle of distilled water |  |
| 250-mL plastic bottles (1 per pond), if doing ARM pH and alkalinity on same visit |  |

Laboratory Preparation Before Sampling

- Collect all equipment

- Use a short length (2 m) of tubing on the pump.  Attach a float to the tubing in order to suspend the submerged end of the tubing 0.5 meters under the water’s surface.  Clean out tubing by pumping 50 ml 1N HCl and then rinsing with distilled water until pH returns to about 5.6.

- Clean 70-ml glass test tubes.

Pre-rinse tubes with tap water. Fill the Branson ultrasonic cleaner with tap water.   Add a dash of Citranox.  Heat water in ultrasonic unit to maximum heat and then turn on ultrasound for minimum 90 minutes.  Check all caps to make sure the seals are still tight. Once cleaned in the ultrasonic bath, flush the tubes with DI water to eliminate soapy residue.  Rinse three times or soak over night in HCl rinse (1:4 HCl:water).  Caution: When making this acid rinse, remember to add the acid to water.  Rinse three times with distilled water.  Dry tubes inverted in test tube rack and once dry, cap test tubes to keep clean.

- Clean the 50-ml centrifuge bottles, scintillation vials and 250-mL plastic bottles by rinsing three times with HCl rinse (1:4 HCl to water) and then rinse three times with distilled water.

- Pre-acidify 50-ml clear, HDPE centrifuge bottles by adding 0.2 ml of 2 N trace metal grade HCl to each bottle.

- Label all bottles and glassware according to the pond name, date and method (if applicable).

- Calibrate the YSI  (see Appendix A).

- Copy field sheets onto “Write on the Rain” paper.

- Make reagents:

Potassium persulfate oxidizing reagent (for Total Nitrogen/Total Phosphorus)

50 g Potassium persulfate (low nitrogen– see page 33 for procedure)

30 g Boric Acid

350 ml 1M NaOH (40 g/L)

Dissolve above reagents in a 1-L volumetric flask (acid and DI water rinsed) and bring to volume.  Transfer to amber 5-mL repipet dispenser.  Holding time of oxidizing reagent is one week although it is recommended that new reagent be made daily.

Data Collection - Profile and Transparency

- Follow the directions for biweekly pond sampling.

Sample Collection - Alkalinity and pH

- Rinse the 250-ml plastic bottles with pond water.

- Fill the bottle with pond water.  Eliminate all air pockets from the top of the bottle by holding the bottle and bottle top under water while screwing on the cover.

- Store in the cooler.

Sample Collection and Preservation for Dissolved Nutrients (orthophosphate, NH4, NO3 and major cations (Ca, Mg, K, Na) (in triplicate)

Note: Water samples for nutrients, metals, cations and anions are collected in triplicate using the Geotech peristaltic pump.

- Connect the battery to the pump.

- Lower the intake tubing to 0.5 meters below the pond surface.

- Rinse and condition the tubing by pumping pond water though the hose.  To make sure the tubing has been thoroughly rinsed, pump about 250 ml of waste before sampling.

- Attach a 47mm Millipore filter holder with a 0.45 millipore filter to the discharge end of the tubing.

- Turn on the pump. Allow the first few milliliters to discharge to waste and then direct the flow into a pre-acidified (0.2 ml of 2N trace metal grade HCl) 50-ml clear HDPE centrifuge tube. Cap each sample, shake and place in the cooler. Samples are stored on ice or refrigerated.  Replace filter when pump begins laboring

- See Appendix D for Dissolved Nutrients laboratory methods and Appendix F for cations laboratory methods..

Sample Collection - Anions (Cl- and SO4-) (collected in triplicate only during the April survey)

- This procedure is most efficiently done before or after the Dissolved Nutrient procedure since the equipment needs for both are basically the same.

- You may use the same filter holder and filter (change the Millipore filter when it becomes clogged and the pump starts to labor).

- Rinse each vial with filtered pond water before filling.

- Pond water is passed from the pump tube through a 0.45-(m membrane filter and into an acid-cleaned, distilled-water rinsed 60-ml brown polyethylene bottle. Note that this bottle must be thoroughly rinsed as any acid residue will contaminate samples with acid anions.   Samples are stored on ice or refrigerated.

See Appendix G for anions laboratory methods.

Sample Collection - Total Phosphorus and Total Nitrogen (triplicates)

- Pump 40 mL unfiltered pond water directly into the clean borosilicate test tube.

- Add 5 mL of persulfate oxidizing reagent. Alternatively, cap and store samples in cooler and add the persulfate oxidizing reagent immediately upon returning to the laboratory.

- Cap and invert numerous times to mix.

- Samples are digested the same day of collection by pressure cooking at approximately15 psi for 50 min and then stored at room temperature for no more than 28 days.

See Appendix E for the total phosphorus and total nitrogen laboratory method.

After Sampling

- Refrigerate the dissolved nutrients and anion samples (record information on the log-in sheet on the refrigerator).

###### If TP/TN samples were not treated with oxidizing reagent in the field, add 5 mL of persulfate oxidizing reagent to the TP/TN samples.   Place the samples in a pressure cooker and cook at 15 psi for 50 minutes.  Label the reagent date on tubes.

- Clean all used filter holders with distilled water and fill with new filters.

- Recharge pump battery if needed.

- Put ice packs back into the freezer.

- Discard used Millipore filters.

- Put the equipment away.  Disconnect the YSI Sonde from the cable and 650 datalogger.  Fill the calibration cup with a small amount of water and replace the cup onto the end of the probe.  Replace the plastic protectors on the cable ends and the top of the Sonde.

- Allow the alkalinity and pH samples to warm up to room temperature.  These samples must be analyzed the day they are collected.  The procedure for this analysis is described in Appendix C.

# Appendix I – Pond Stage.

Siphon Measurements

Ryder, Duck, Great Wellfleet, Great Truro, Herring, Snow, Dyer and Long Pond all have siphon wells.

Using an electric well tape, lower the sensor until an audible beep occurs.  Record the position where the tape intersects the well casing and record to the nearest 0.01ft. The casing will be marked with ink at the measurement point.  Refer to the Long Term Hydrologic Monitoring Protocol for Coastal Ecosystems, 2003.

# Appendix J - Data Forms.

Form 1 - pH and Alkalinity Form

| Alkalinity Form |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| (P=Provincetown T=Truro W=Wellfleet) |  |  |  |  |  |  |
|  |  | date | date |  | ml titrant to: |  |  |
| POND | ID# | Collection | Analysis | pH | pH4.5 | pH4.2 | Alkalinity |  |
| Seawater |  |  |  |  |  |  |  |  |
| DUCK (P) | 96866 |  |  |  |  |  | 0 |  |
| DUCK (W) | 96868 |  |  |  |  |  | 0 |  |
| DYER (W) | 96070 |  |  |  |  |  | 0 |  |
| FRESHBROOK (W) | 9661575 |  |  |  |  |  | 0 |  |
| GREAT (T) | 96114 |  |  |  |  |  | 0 |  |
| GREAT (W) | 96117 |  |  |  |  |  | 0 |  |
| GULL (W) | 96123 |  |  |  |  |  | 0 |  |
| HERRING POND (W) | 96134 |  |  |  |  |  | 0 |  |
| HERRING RIVER (T) | 9661650 |  |  |  |  |  | 0 |  |
| HIGGINS (W) | 96137 |  |  |  |  |  | 0 |  |
| HORSELEECH (W) | 96144 |  |  |  |  |  | 0 |  |
| KINNACUM (W) | 96163 |  |  |  |  |  | 0 |  |
| LIT.BENNET (P) | 96347 |  |  |  |  |  | 0 |  |
| LONG1 (W) | 96179 |  |  |  |  |  | 0 |  |
| LONG2 (W) |  |  |  |  |  |  | 0 |  |
| NORTHEAST(W) | 96226 |  |  |  |  |  | 0 |  |
| PAMET RIVER (T) | 9661725 |  |  |  |  |  | 0 |  |
| ROUND EAST (T) | 96261 |  |  |  |  |  | 0 |  |
| ROUND WEST (T) | 96260 |  |  |  |  |  | 0 |  |
| RYDER (T) | 96268 |  |  |  |  |  | 0 |  |
| SLOUGH (T) | 96298 |  |  |  |  |  | 0 |  |
| SNOW (T) | 96303 |  |  |  |  |  | 0 |  |
| SOUTHEAST (W) | 96305 |  |  |  |  |  | 0 |  |
| SPECTACLE (W) | 96306 |  |  |  |  |  | 0 |  |
| TURTLE (W) |  |  |  |  |  |  | 0 |  |
| WILLIAMS (W) | 96341 |  |  |  |  |  | 0 |  |

Note:  The non-kettle ponds listed are part of the A.R.M. study.

Form 2. Field form for kettle pond water quality monitoring.

| Kettle Pond Field Form | Fathometer: |  | (M or ft) |  |  |
| --- | --- | --- | --- | --- | --- |
| Pond |  |  | Date |  |  | Pre-profile %DO: |  |
| Weather |  |  | Secchi |  |  | Post-profile %DO: |  |
| Depth M | F/CTemp | Conduct | %DO | DO ppm | pH | Eh | %Trans |
| Air |  |  |  |  |  |  |  |
| 0.1 |  |  |  |  |  |  |  |
| 0.5 |  |  |  |  |  |  |  |
| 1 |  |  |  |  |  |  |  |
| 2 |  |  |  |  |  |  |  |
| 3 |  |  |  |  |  |  |  |
| 4 |  |  |  |  |  |  |  |
| 5 |  |  |  |  |  |  |  |
| 6 |  |  |  |  |  |  |  |
| 7 |  |  |  |  |  |  |  |
| 8 |  |  |  |  |  |  |  |
| 9 |  |  |  |  |  |  |  |
| 10 |  |  |  |  |  |  |  |
| 11 |  |  |  |  |  |  |  |
| 12 |  |  |  |  |  |  |  |
| 13 |  |  |  |  |  |  |  |
| 14 |  |  |  |  |  |  |  |
| 15 |  |  |  |  |  |  |  |
| 16 |  |  |  |  |  |  |  |
| 17 |  |  |  |  |  |  |  |
| 18 |  |  |  |  |  |  |  |
| 19 |  |  |  |  |  |  |  |
| 20 |  |  |  |  |  |  |  |

# Appendix K.  The Need for Annual Sampling.

Analysis of transparency data collected through year 2000 indicates that water quality monitoring should continue on an annual basis.  Reasons are as follows:

The most direct measure of clarity is Secchi depth.  We compare inter-annual coefficients of variation (CV) for Secchi given 1, 2, 3, 4 or 5 yr. monitoring intervals for the six ponds with the most data (Fig.1).  Inter-annual variation is very high even with annual sampling.  Low CVs for the 4- or 5-yr interval are attributed to chance given the small sample size, in most cases only 7-15 years of fairly continuous annual data (Fig 2).

We also compare CV in Secchi within years to inter-annual CV at annual, biennial, etc. sampling intervals for representative Duck (Fig. 3) and Ryder (Fig. 4) ponds.  Given the large annual CVs, there is simply not enough data, i.e. years of observation, to say that less frequent sampling can discern trends.  We have already determined that biweekly observations are necessary to characterize Secchi depth, because month to month changes can equal the full range of Secchi over the entire April-October season of stratification (see Table 3, above).

Other monitoring variables (TN, TP, DO and light profiles) do not directly indicate the water’s clarity, but are essential in understanding the reasons for water quality changes, though they do take time and considerable expertise to do well.  A recent review of data from throughout the world in “Limnology & Oceanography” (Nurnberg, 1995, 40:1100-1111) showed that at least monthly profiles of DO are needed to track changes in hypolimnetic anoxia, a variable that reviewers of this protocol have consistently stressed as critical to assess eutrophication.

# Text boxes

Membrane 0.45-(m filter in 47-mm holder

Peristaltic pump

Store cold in DI-rinsed 60-mL HDPE brown bottles for Anions (Cl- and SO4-)

40-ml sample taken in a 70-mL test tube to which 5 mL of persulfate oxidizing reagent is added
for Total P and Total N

Store in pre-acidified freestanding 50-ml HDPE clear centrifuge tubes for Dissolved Nutrients and cations (Ca, Mg, K, Na)

Figure 1.

Figure 2.

Figure 4.

Figure 3.