Proposed Medical Implant Rule Changes

401-406 MHz

During 1996 and 1997 a manufacturer of medical devices petitioned the FCC to establish rules allowing the operation of data telemetry links (Medical Implant Communications or MICS) in the band 401-406 MHz for communication with implants within the human body.  At the time the most prevalent planned use was pacemakers, where the pacemaker could be programmed in a hospital or clinic environment.  The band 401-406 MHz offered many advantages due to propagation characteristics of flesh and a relatively quiet operating environment.  The links were only intended to be operated at a range of less than 10 to 15 feet.  Studies were conducted within the U.S. and internationally that led to the development of International Telecommunications Union- Radiocommunication Sector (ITU-R) Recommendation SA. 1346.  The studies showed the potential for interference to radiosonde systems was very low, but without appropriate mitigation techniques, the MICS link could receive interference from radiosondes.  Interference to an implant during a programming session could potentially harm a patient.  ITU-R SA.1346 recommended mitigation techniques to ensure radiosondes would not cause harm to patients through interference caused to the MICS operating on a non-interference basis.  The primary concern from the meteorological community was ensuring patient harm could not occur, even though MICS was operating on a non-interference basis (NIB).  Radiosonde operations in the band clearly have priority over the MICS, but concern existed over the potential perception and emotional response if radiosonde transmissions were to cause harm.

The current FCC rules allowing the operation of MICS within the band 401-406 MHz require implementation of the mitigation techniques defined in ITU-R SA.1346.  Over the last several years the FCC has received two waiver requests asking for relief from some of the mitigation techniques intended to protect patients from interference.  The medical implant devices for which the waivers were requested were of a nature that interference would cause delay in transmission of data from the implant to a body-worn receiver, but no patient harm could result from the delay.  The FCC, with NTIA concurrence, has granted both devices waivers based on the conclusion that patient harm could not occur, if interference was to occur from radiosondes.   Each one of these waiver approvals required careful analysis to ensure patient harm would not occur.  Based on these two waivers, a medical implant manufacturer has petitioned the FCC for creation of new set of rules that do not require the mitigation techniques for patient protection.  The new rules would apply to Medical Data Services (MEDS), devices the manufacturers determine are not performing a function where interference could cause harm or loss of life.

While other Federal agencies are the primary users of radiosondes in the 401-406 MHz band, NTIA has been looking to DOC/NOAA for advice on this issue.  As with MICS, the potential for interference from MEDS to radiosondes is small, though greater than MICs due to a higher intended proliferation of the MEDS devices and the fact they will transmit outside the hospital/clinic environment.   The primary concern is allowing the manufacturers to determine whether a radiosonde interference event could cause a life-threatening issue.  In many cases the manufacturers do not understand radiosonde operations or how those operations could impact a MEDS user.  The agencies using the 401-406 MHz band for radiosonde operations should evaluate the proposed rules and work with their frequency management offices to raise issues, concerns or agreement to the proposed new rules.  Attention should given to the potential for radiosondes to cause interference to MEDS, and the resulting impact if it was shown radiosonde transmissions were harming MEDS patients.  Sharing of information on agency views would also be of value to all WG-FM members.