##### QADATA Summary – February 2008

The Inorganic Blind Sample Project (IBSP) of the Branch of Quality Systems reviews about 90 inorganic analytical lines at the National Water Quality Laboratory (NWQL) using double-blind, quality-control samples.  A summary of this review, the QADATA Summary, is a short report commenting on only those analytical lines for which the quality-control charts indicate or have indicated a bias and/or higher than expected variability.  Overall, the vast majority of the inorganic methods at the NWQL demonstrate acceptable variability and little or no bias.

The data reviewed in this summary is for the time period 07/01/07 through 01/31/08; however, longer-range charts may be used when they add clarity to the problem.  The attached quality-control charts contain data from samples that have been subject to the same laboratory handling procedures as the samples submitted by the USGS district offices, and the data have the same level of quality-control review as the data released to the districts.  Therefore, any factors that affect the data displayed on the IBSP quality-control charts should also affect the districts’ data in a similar manner.  Data from samples known to have been affected by a problem independent of the NWQL are removed from the charts.

In this report, the “Problem” is determined and written by IBSP and then sent to the NWQL.  The NWQL then gets two weeks to respond to the issue by writing an “NWQL Comment” and a “Corrective Action.”  Comments from the lab are incorporated into the QADATA Summary as received.  The most recent comments and responses are in bold followed by comments and responses from previous QADATA Summaries.  Approximately one month after the QADATA Summary is distributed, the NWQL and IBSP personnel usually meet to follow-up on corrective actions and on how the corrective actions are affecting subsequent data.

Current charts (updated twice weekly) may be accessed at  http://bqs.usgs.gov/bsp/Controlcharts.htm.  Laboratory and district personnel may have their own custom quality-assurance data retrieval to parallel their objectives by contacting Ted Struzeski (303.236.1872, struzesk@usgs.gov).  Charts pertaining specifically to this report can be accessed electronically by following the underlined hyperlinks embedded within the report and associated with each analysis discussed.

Historical QADATA Summaries may be accessed at http://bqs.usgs.gov/bsp/QAData.htm.

Tedmund M. Struzeski

struzesk@usgs.gov

Table of Contents

Report Highlights 	Page

One-Page Summary	3

Detailed Description by Analyte

Trace Metals

Copper, Unfiltered,  cICP-MS, Labcode 3129, Parmcode 01042PLM10	4

Manganese, Unfiltered,  ICP, Labcode 2363, Parmcode 01055PLA15	4

Mercury, Filtered,  CVAF, Labcode 2707, Parmcode 71890CV014	5

Mercury, Unfiltered,  CVAF, Labcode 2708, Parmcode 71900CV018	6

Vanadium, Unfiltered,  cICP-MS, Labcode 3135, Parmcode 01087PLM10	6

Zinc, Filtered,  cICP-MS, Labcode 3138, Parmcode 01090PLM10	7

Majors

Fluoride, Filtered,  ISE, Labcode 0031, Parmcode 00950ISE05	9

Magnesium, Unfiltered,  ICP, Labcode 2362, Parmcode 00927PLA15	9

Sulfate, Filtered,  IC, Labcode 1572, Parmcode 00945IC022	10

Report Highlights

Note:  Click the underlined text to access the control chart associated with the comment.

- Copper, unfiltered, cICP-MS, parmcode 01042PLM10– High variability starting mid October 2007.

- Fluoride, filtered, ISE, parmcode 0950ISE05– Positive bias starting mid December 2007.  Current bias is 16%

- Mercury, filtered, CVAF, parmcode 71890CV014 – Negative bias starting May 2007.  Most recent data show improvement.

- Mercury, unfiltered, CVAF, parmcode 71900CV018 – Negative bias starting mid May 2007.  Most recent data show improvement.

- Magnesium, unfiltered, ICP, parmcode 00927PLA15– Negative bias starting mid August 2007.  Current bias is -6.6%.

- Vanadium, unfiltered, cICP-MS, parmcode 01087PLM10– High variability and positive bias (for blind-blanks only) starting November 2007.

- Zinc, filtered, cICP-MS, parmcode 01090PLM10 – Positive bias starting January 2007.  Current bias is 21%.

NWQL PERFORMANCE ON IBSP SAMPLES

Note:  Click the underlined text in the left-hand column to access the control chart for the associated comment.

| Trace Metal Determinations | Problem | NWQL Comment | Corrective Action |
| --- | --- | --- | --- |
| Copper, Unfiltered cICP-MS labcode 3129 parmcode 1042PLM10 | High variability is demonstrated starting mid October 2007. This variability is not shown for filtered copper by cICP-MS (pc 1040PLM10). | We have looked at a number of possible causes, but none of the attempted corrections has been proven to correct the issue yet. The most recent points are well within limits and one of our attempted corrections may have fixed the problem. | We will continue to investigate possible causes and request a review of the process by the Quality Assurance Section of the NWQL. The NWQL will continue to monitor the newest data closely. |
| Manganese, Unfiltered ICP labcode 2363 parmcode 01055PLA15 | Data from samples analyzed since the last assessment show acceptable bias and variability. No corrective action necessary. |  | The NWQL will continue to monitor the instrument for bias and variability. |
| Previous comments for manganese, unfiltered, labcode 2363, parmcode 01055PLA15 |
| From QADATA Summary December 2007 | Data for samples analyzed as of the last assessment show an improvement in bias. As of then results show a negative bias of less than 5% on average. Analysis should be closely monitored until next QADATA Summary is distributed. | Internal NWQL charts indicate a similar trend. | The NWQL will continue to monitor the instrument for bias and variability. |
| From QADATA Summary October 2007 | The data show a slight negative bias with all results in the last six months below the target value. Since the last assessment results are on average 6.4% below the target value. | NWQL online charts also show a negative bias near the concentrations on the BQS charts. | The NWQL had the instrument manufacturer come out (the week of October 15th, 2007) and do a complete check of the instrument. He optimized the instrument and initial runs indicate reduced variability. NWQL will continue to monitor the instrument. |
| Mercury, Filtered CVAF labcode 2707 parmcode 71890CV014 | Samples analyzed as of the last assessment show an improvement in the negative bias. Analysis should be closely monitored until the next QADATA Summary is distributed. |  | The NWQL will continue to monitor the instrument for bias. |
| Previous comments for mercury, filtered, labcode 2707, parmcode 71890CV014 |
| From QADATA Summary December 2007 | Results as of the last assessment show an improvement in the negative bias. As of then, results show an average negative bias of 7.8% below the target value. | New calibration standards were prepared in early December; however, due to instrument problems that have recently been fixed, not enough internal QC points have been generated in order to know if the bias has been corrected. | The NWQL will continue to monitor the instrument for bias as more data becomes available. |
| From QADATA Summary October 2007 | The negative bias has increased in magnitude. As of the last assessment, results are on average 13% below the target value, and all but one result is less than the target value. | New calibration standards were prepared in September and current QC charts indicate that bias has been fixed. | The NWQL will continue to monitor the instrument for bias. |
| From QADATA Summary August 2007 | A negative bias is still shown. As of May 2007, results are on average 12% below the target value, and 23 out of the last 25 results are less than the target value. | There was improvement at first after new standards were made in early July, but the last points show a downward shift. On-line QC charts indicate the bias is at the higher values, but not near the low end of the curve. | We will continue to monitor mercury closely and check reagents and the instrument for problems. |
| From QADATA Summary June 2007 | The data demonstrate a downward shift starting May 2007. Since then, results are on average 12% below the target value. | NWQL QC charts indicate a similar trend. | We will make new standards and will closely monitor the analysis for bias. |
| Mercury, Unfiltered CVAF labcode 2708 parmcode 71900CV018 | Samples analyzed as of the last assessment show an improvement in the negative bias. Analysis should be closely monitored until the next QADATA Summary is distributed. |  | The NWQL will continue to monitor the instrument for bias. |
| Previous comments for mercury, unfiltered, labcode 2708, parmcode 71900CV018 |
| From QADATA Summary December 2007 | A negative bias is still shown. As of the last assessment, results are on average 12% below the target value. | New calibration standards were prepared in early December; however, due to instrument problems that have recently been fixed, not enough internal QC points have been generated in order to know if the bias has been corrected. | The NWQL will continue to monitor the instrument for bias as more data becomes available. |
| From QADATA Summary October 2007 | The negative bias has increased in magnitude. As of the last assessment, results are on average 11% below the target value, and most results are less than the target value. | New calibration standards were prepared in September and current QC charts indicate that bias has been fixed. | The NWQL will continue to monitor the instrument for bias. |
| From QADATA Summary August 2007 | The data exhibit a negative bias starting mid May 2007. Since then, results are on average 6.2% below the target value. | The most recent points show improvement. | We will continue to monitor mercury closely. |
| Vanadium, Unfiltered cICP-MS labcode 3135 parmcode 01087PLM10 | High variability and a positive bias are shown for blind blanks. Since November 2007, 63% of the blind-blank results are greater than the IRL. | We are seeing this with our blanks as well. We believe that the variability is instrumental and have contacted the manufacturer to have maintenance performed. If the variability is related to the instrument aging, we may need to raise the IRL for this element. | We will monitor the instrument closely after the maintenance and take appropriate steps based on the information gathered at that time. |
| Zinc, Filtered cICP-MS labcode 3138 parmcode 01090PLM10 | The positive bias continues to increase. Since the last assessment, almost half of the results are outside of control limits and the overall bias is 21%. Results for unfiltered zinc by cICP-MS (pc 1092PLM10) are not exhibiting this bias or variability. | We have looked at a number of possible causes, but none of the attempted corrections has been proven to correct the issue yet. The most recent modification to the calibration procedure is showing promise, but we want to see more data to be sure. | It will take 2-3 more weeks of data to ensure that this solution is working. We will continue to monitor this line closely and if the line is still having problems we will request input from the Quality Assurance Section of the NWQL on March 11, 2008. |
| Previous comments for zinc, filtered, labcode 3138, parmcode 01090PLM10 |
| From QADATA Summary December 2007 | The positive bias displayed since January 2007 has increased even more since the last assessment. The most recent data show a positive bias greater than 13%, and one out of every four results is outside of control limits. Results for unfiltered zinc by cICP-MS are not exhibiting this problem. | On-line QC charts indicate the bias does not appear to be related to standards. The bias also does not seem to be related to carry-over or contamination. The analyst is investigating the possibility of interference, however, the instrument is currently undergoing repair. More research needs to be conducted as soon as the instrument is operational again. | The NWQL will continue to monitor the instrument for bias. |
| From QADATA Summary October 2007 | This analytical line has shown a positive bias since January 2007. As of then, the results are on average 13% above the target value, and 89 of the 95 results are greater than the target value. Results for unfiltered zinc by cICP-MS (pc 1092PLM10) come from the same reference materials but do not exhibit this bias. | On-line QC was showing no bias so the analyst purchased another standard from a different source. Initial analyses show a bias similar to the one on the BQS charts. | The NWQL will make new calibration and check standards from new stock solutions. |
| From QADATA Summary August 2007 | Although there is some improvement, a positive bias is still present. Results received since the last assessment are on average 14% above the target value with most results greater than the target value. A positive bias is not evident for unfiltered zinc by standard ICP-MS (pc 1092PLM48) or by cICP-MS (pc 1092PLM10). | On-line QC charts indicate an equal number of points above and below the MPV at 10 and 100 ug/L. At 10 ug/L there is a difference between the 2 instruments with one trending high and one trending low. | We will monitor the line to see if new results continue to improve the charts. |
| From QADATA Summary June 2007 | The positive bias is increasing. Results received since the last assessment are on average 18% above the target value with three out of 22 results (14%) outside of +/-2 f-pseudosigmas. A positive bias is not evident for unfiltered zinc by cICP-MS (pc 1092E). |  | We are making new standards and we will continue to closely monitor the analysis for bias. |
| From QADATA Summary April 2007 | The data demonstrate a positive bias. As of January 2007, results are on average 11% above the target value, and most results are greater than the target value. | NWQL QC charts indicated a similar bias and new calibration standards were prepared in the end of March 2007. We are still gathering data to see if the bias has improved. | We will continue to closely monitor the analysis for bias. |

| Major Determinations | Problem | NWQL Comment | Corrective Action |
| --- | --- | --- | --- |
| Fluoride, Filtered ISE labcode 0031 parmcode 0950ISE05 | A positive bias is shown for the most recent data. As of mid December 2007, the bias is about 16%, on average, above the target value. Note: The concentration of the blind samples submitted cover a fairly limited portion of the method’s analytical range. | The electrodes (ISE and reference) on the system typically need to be replaced on a 5-6 month schedule. The analyst continued to use the old electrodes while analyzing some difficult samples that arrived in December in order to not potentially ruin new electrodes that are very expensive. | New electrodes are being conditioned this week and will be used on the line when they are ready for use. This should lower the observed bias. |
| Magnesium, Unfiltered ICP labcode 2362 parmcode 00927PLA15 | Although there is some improvement in the negative bias, a bias is still present. Results received since the last assessment are on average 6.6% below the target value. | The new standards from previously purchased stock standards had some improvement, but we still show a negative bias. | We will purchase new stock solutions and make new standards. New standards will be on-line one week after the new solutions arrive. |
| Previous comments for magnesium, unfiltered, labcode 2362, parmcode 00927PLA15 |
| From QADATA Summary December 2007 | The magnitude of the negative bias is increasing. Results received since the last assessment are on average 10% below the target value with two out of the six results (33%) outside of -2 f-pseudosigmas. | On-line QC charts indicate similar bias. It is probably due to calibration standards. New calibration standards were prepared the second week of December. | The NWQL will continue to monitor the instrument for bias. |
| From QADATA Summary October 2007 | The data demonstrate a negative bias. As of mid August 2007, results are on average 7.4% below the target value. | Online charts show a slight high bias for this analysis at concentrations above and below those submitted by BQS. | The NWQL had the instrument manufacturer come out (the week of October 15th, 2007) and do a complete check of the instrument. He optimized the instrument and initial runs indicate reduced variability. NWQL will continue to monitor the instrument. |
| Sulfate, Filtered IC labcode 1572 parmcode 00945IC022 | A positive bias is shown starting late October 2007 and going through mid December for blind standard reference materials with a concentration close to 14 mg/L. As of mid December, the bias is no longer apparent. No corrective action necessary. | Changes to the instrument calibration calculation resulted in a positive bias for samples in the 10-15 mg/L concentration range. | Samples with results affected by the change in the calibration calculation have had results recalculated using a new calibration calculation. Results showing a positive bias will be updated to their corresponding lower values and re-released. |
| Previous comments for sulfate, filtered, labcode 1572, parmcode 00945IC022 |
| From QADATA Summary December 2007 | Data exhibit increased variability starting late October 2007. Analysis should be closely monitored until next QADATA Summary is distributed. |  |  |