LBL Trip Summary and Task List

Thursday, February 21, 2008

On the week of February 18, 2008 Peter Denes, Dionisio Doering and John Weizeorick worked to test the FCCD’s Clock Module.  The major accomplishment of this week is seeing first light out of the 480 x 480 CCD using external voltage bias and the FCCD Clock Module.  A more detail list of tasks completed are listed below.

- Tested the small 20 x 480 CCD with the Leach controller.  Needed to modify the clock pin out to match the FCCD clock module.  Observed the CCD output on a scope while modifying the time an LED is turned on during exposure.

- Tested the small CCD with the Leach controller providing the bias voltages and the FCCD detector providing the clocks.  This was done by using a 10 inch adapter cable from the FCCD Clock Module to the DB37 pin connector on the Frame Board.  The CCD output behaved the same with the faster clocks.

- Tested out how shorting the vertical and horizontal right and left signals affected the clock signals.  The waveforms appeared to become cleaner.

- Tested the 480 x 480 CCD with the Leach controller providing all of the bias voltages except the VDDOUT, which required -22V at 140mA, and used the FCCD’s Clock Module to provide the clock signals.  Observed and plotted how the longer light pulses during the exposure increased the CCD output signal.

- Tried different voltage bias, clock rail voltages, and clock timing and observed the effects on the CCD output.

- Improved the linearity of the CCD output by slowing down the horizontal clocks.

In preparation for the next meeting the following task need to be completed by Argonne.

- Do not need to fix -25V error on Interface Module.  Voltage is never used.

- Need to modify VDDOUT circuit to provide 0 to -25V at 300 mA.

- Need to change the source of  VDDA_AS (regulated 3.3V at ? Amps) VDDA_ADC (regulated 3.3V) These voltages are currently coming form the Data Module.  These voltages will tied together and come from the Clock Module.

- Final Clock Module will also need a connector change.  We will be making it a larger connector and make the pin out symmetrical.

- The Data Module will need a connector soldered onto its edge connector.  This connector will plug into an edge connector coming out of the vacuum chamber.

- Move the single point ground into the Acopian power supply.  Argonne will do the power supply at Argonne and email a marked up schematic so that LBNL can make the same changes.

- Fix Data Module’s FPGA program so that the fCRIC power up reliably.

- Add more features to GUI to allow greater control over the clock signals.

- Fix state machine problem which is generating fewer clocks on the first line.

- Remove small caps on clock driver outputs.

The tasks that LBNL will need to complete before the next meeting are.

- Make small adapter boards.  One between Data Module and vacuum edge connector.  One for between Clock Module and vacuum edge connector.  The adaptor boards are meant to provide a method of connecting the original FCCD Clock and Data Modules with the final vacuum edge connectors.

- Do glue test on silicon wafer and PCB board.

- Assemble vacuum housing with vacuum feed through boards epoxy in place.

- Finnish Assembly and install into vacuum housing.

- Let Argonne know what the power up voltages should be for NGD, NCON, and GUARD.  I believe that they currently power up at +10Volts and have a range of -10 Volts to +10 Volts.

- Set up temperature controller for final system.

- Simulate vertical clocks and determine the optimum RC Network.