## Ca
Statistical summary of reported data for standard reference sample P-42 (low ionic strength constituents)
+3 Fs
+1.5 Fs
MPV
-1.5 Fs
-3 Fs
SUMMARY |  |  | Methods | [8 empty] | Statistics |  |  |  |  | SCALE
|  |  | 1 | 2 | 4 | 5 | 6 | 7 | 20 | Method Codes | [6 empty] | min | 0.43384
|  | n = | 5 | 1 | 18 | 1 | 2 | 3 | 2 |  | 01 Atomic absorption: direct, air |  | MPV = | 0.495 | mg/L |  | max | 0.55616
|  | Minimum = | 0.442 | 0.48 | 0.244 | 0.496 | 0.45 | 0.479 | 0.16 |  | 02 Atomic absorption: direct, nitrous oxide |  | F-pseudosigma = | 0.020385 |  |  | major | 0.030578
|  | Maximum = | 0.625 |  | 0.64 |  | 0.48 | 0.5808 | 0.168 |  | 04 Inductively coupled plasma |  | Rating criterion = | 0.02475 |  |  | minor | 0.030578
|  | Median = | 0.5 |  | 0.4975 | [5 empty] | 05 Direct current plasma |  | n = | 32
|  | F-pseudosigma = | 0.044477 |  | 0.015567 | [5 empty] | 06 Inductively coupled plasma / mass spectrometry |  | Uh = | 0.502
[11 empty] | 07 Ion chromatography |  | Lh = | 0.4745 |  |  | Order | 1 | 2 | 4 | 5 | 6 | 7 | 20 | Lab | RV | Rating
[11 empty] | 20 Titration: colorimetric | [5 empty] | 1 | 0.442 | [6 empty] | 1 | 0.502 | 4
[17 empty] | 2 | 0.47 | [6 empty] | 2 | 0.499 | 4
|  |  | Method Codes | [13 empty] | 3 | 0.5 | [6 empty] | 5 | 0.47 | 2
Lab | Rating | Z-value | 1 | 2 | 4 | 5 | 6 | 7 | 20 | [7 empty] | 4 | 0.53 | [6 empty] | 8 | 0.49 | 4
1 | 4 | 0.28283 | -- | -- | 0.502 | -- | -- | -- | -- | [7 empty] | 5 | 0.625 | [6 empty] | 23 | 0.47 | 2
2 | 4 | 0.16162 | -- | -- | -- | -- | -- | 0.499 | -- | [7 empty] | 6 |  | 0.48 | [5 empty] | 25 | 0.47 | 2
5 | 2 | -1.0101 | -- | -- | 0.47 | -- | -- | -- | -- | [7 empty] | 7 |  |  | 0.244 |  |  |  |  | 33 | 0.496 | 4
8 | 4 | -0.20202 | -- | -- | 0.49 | -- | -- | -- | -- | [7 empty] | 8 |  |  | 0.47 |  |  |  |  | 38 | 0.48 | 3
23 | 2 | -1.0101 | 0.47 | -- | -- | -- | -- | -- | -- | [7 empty] | 9 |  |  | 0.47 |  |  |  |  | 45 | 0.48 | 3
25 | 2 | -1.0101 | -- | -- | 0.47 | -- | -- | -- | -- | [7 empty] | 10 |  |  | 0.48 |  |  |  |  | 59 | 0.479 | 3
33 | 4 | 0.040404 | -- | -- | -- | 0.496 | -- | -- | -- | [7 empty] | 11 |  |  | 0.481 |  |  |  |  | 64 | 0.5 | 4
38 | 3 | -0.60606 | -- | 0.48 | -- | -- | -- | -- | -- | [7 empty] | 12 |  |  | 0.49 |  |  |  |  | 105 | 0.495 | 4
45 | 3 | -0.60606 | -- | -- | -- | -- | 0.48 | -- | -- | [7 empty] | 13 |  |  | 0.494 |  |  |  |  | 110 | 0.502 | 4
59 | 3 | -0.64646 | -- | -- | -- | -- | -- | 0.479 | -- | [7 empty] | 14 |  |  | 0.495 |  |  |  |  | 134 | 0.495 | 4
64 | 4 | 0.20202 | 0.5 | -- | -- | -- | -- | -- | -- | [7 empty] | 15 |  |  | 0.495 |  |  |  |  | 180 | 0.481 | 3
105 | 4 | 0 | -- | -- | 0.495 | -- | -- | -- | -- | [7 empty] | 16 |  |  | 0.5 |  |  |  |  | 190 | 0.5 | 4
110 | 4 | 0.28283 | -- | -- | 0.502 | -- | -- | -- | -- | [7 empty] | 17 |  |  | 0.5 |  |  |  |  | 193 | 0.45 | 1
134 | 4 | 0 | -- | -- | 0.495 | -- | -- | -- | -- | [7 empty] | 18 |  |  | 0.5 |  |  |  |  | 209 | 0.554 | 0
180 | 3 | -0.56566 | -- | -- | 0.481 | -- | -- | -- | -- | [7 empty] | 19 |  |  | 0.502 |  |  |  |  | 265 | 0.5 | 4
190 | 4 | 0.20202 | -- | -- | 0.5 | -- | -- | -- | -- | [7 empty] | 20 |  |  | 0.502 |  |  |  |  | 268 | 0.625 | 0
193 | 1 | -1.8182 | -- | -- | -- | -- | 0.45 | -- | -- | [7 empty] | 21 |  |  | 0.52 |  |  |  |  | 273 | 0.244 | 0
209 | 0 | 2.3838 | -- | -- | 0.554 | -- | -- | -- | -- | [7 empty] | 22 |  |  | 0.554 |  |  |  |  | 274 | 0.16 | 0
265 | 4 | 0.20202 | -- | -- | 0.5 | -- | -- | -- | -- | [7 empty] | 23 |  |  | 0.555 |  |  |  |  | 279 | 0.442 | 0
268 | 0 | 5.2525 | 0.625 | -- | -- | -- | -- | -- | -- | [7 empty] | 24 |  |  | 0.64 |  |  |  |  | 284 | 0.555 | 0
273 | 0 | -10.1414 | -- | -- | 0.244 | -- | -- | -- | -- | [7 empty] | 25 |  |  |  | 0.496 |  |  |  | 321 | 0.53 | 2
274 | 0 | -13.5354 | -- | -- | -- | -- | -- | -- | 0.16 | [7 empty] | 26 |  |  |  |  | 0.45 |  |  | 323 | 0.48 | 3
279 | 0 | -2.1414 | 0.442 | -- | -- | -- | -- | -- | -- | [7 empty] | 27 |  |  |  |  | 0.48 |  |  | 327 | 0.52 | 2
284 | 0 | 2.4242 | -- | -- | 0.555 | -- | -- | -- | -- | [7 empty] | 28 | [5 empty] | 0.479 |  | 328 | 0.64 | 0
321 | 2 | 1.4141 | 0.53 | -- | -- | -- | -- | -- | -- | [7 empty] | 29 | [5 empty] | 0.499 |  | 333 | 0.5 | 4
323 | 3 | -0.60606 | -- | -- | 0.48 | -- | -- | -- | -- | [7 empty] | 30 | [5 empty] | 0.5808 |  | 336 | 0.168 | 0
327 | 2 | 1.0101 | -- | -- | 0.52 | -- | -- | -- | -- | [7 empty] | 31 | [6 empty] | 0.16 | 372 | 0.494 | 4
328 | 0 | 5.8586 | -- | -- | 0.64 | -- | -- | -- | -- | [7 empty] | 32 | [6 empty] | 0.168 | 393 | 0.5808 | 0
333 | 4 | 0.20202 | -- | -- | 0.5 | -- | -- | -- | --
336 | 0 | -13.2121 | -- | -- | -- | -- | -- | -- | 0.168
372 | 4 | -0.040404 | -- | -- | 0.494 | -- | -- | -- | --
393 | 0 | 3.4667 | -- | -- | -- | -- | -- | 0.5808 | --

## Cl
Statistical summary of reported data for standard reference sample P-42 (low ionic strength constituents)
+3 Fs
+1.5 Fs
MPV
-1.5 Fs
-3 Fs
SUMMARY |  |  | Methods | [5 empty] | Statistics |  |  |  |  | SCALE
|  |  | 6 | 7 | 20 | 22 | 40 | Method Codes | [5 empty] | min | 4.9858
|  | n = | 1 | 27 | 4 | 1 | 1 | 06 Inductively coupled plasma / mass spectrometry |  | MPV = | 6.12 | mg/L |  | max | 7.2542
|  | Minimum = | 6.3 | 3.14 | 5.396 | 7 | 4.8 | 07 Ion chromatography |  | F-pseudosigma = | 0.37806 |  |  | major | 0.56709
|  | Maximum = |  | 10.88 | 18.33 |  |  | 20 Titration: colorimetric |  | n = | 34 |  |  | minor | 0.56709
|  | Median = |  | 6.14 |  |  |  | 22 Colorimetric |  | Uh = | 6.3
|  | F-pseudosigma = |  | 0.31653 |  |  |  | 40 Ion selective electrode |  | Lh = | 5.79
[14 empty] | Order | 6 | 7 | 20 | 22 | 40 | Lab | RV | Rating
[14 empty] | 1 | 6.3 |  |  |  |  | 1 | 6.34 | 3
[14 empty] | 2 |  | 3.14 |  |  |  | 2 | 5.821 | 3
|  |  | Method Codes | [10 empty] | 3 |  | 5.63 |  |  |  | 5 | 6 | 4
Lab | Rating | Z-value | 6 | 7 | 20 | 22 | 40 | [6 empty] | 4 |  | 5.73 |  |  |  | 8 | 6.08 | 4
1 | 3 | 0.58192 | -- | 6.34 | -- | -- | -- | [6 empty] | 5 |  | 5.78 |  |  |  | 23 | 6.07 | 4
2 | 3 | -0.79088 | -- | 5.821 | -- | -- | -- | [6 empty] | 6 |  | 5.786 |  |  |  | 25 | 5.9 | 3
5 | 4 | -0.31741 | -- | 6 | -- | -- | -- | [6 empty] | 7 |  | 5.79 |  |  |  | 30 | 6.2 | 4
8 | 4 | -0.1058 | -- | 6.08 | -- | -- | -- | [6 empty] | 8 |  | 5.821 |  |  |  | 33 | 6.51 | 2
23 | 4 | -0.13225 | -- | 6.07 | -- | -- | -- | [6 empty] | 9 |  | 5.9 |  |  |  | 38 | 6.6 | 2
25 | 3 | -0.58192 | -- | 5.9 | -- | -- | -- | [6 empty] | 10 |  | 6 |  |  |  | 45 | 5.79 | 3
30 | 4 | 0.21161 | -- | 6.2 | -- | -- | -- | [6 empty] | 11 |  | 6 |  |  |  | 59 | 10.88 | 0
33 | 2 | 1.0316 | -- | 6.51 | -- | -- | -- | [6 empty] | 12 |  | 6.07 |  |  |  | 64 | 6.3 | 4
38 | 2 | 1.2696 | -- | 6.6 | -- | -- | -- | [6 empty] | 13 |  | 6.08 |  |  |  | 89 | 6.38 | 3
45 | 3 | -0.87288 | -- | 5.79 | -- | -- | -- | [6 empty] | 14 |  | 6.1 |  |  |  | 105 | 6 | 4
59 | 0 | 12.5907 | -- | 10.88 | -- | -- | -- | [6 empty] | 15 |  | 6.14 |  |  |  | 110 | 3.14 | 0
64 | 4 | 0.47612 | 6.3 | -- | -- | -- | -- | [6 empty] | 16 |  | 6.1519 |  |  |  | 113 | 6.1 | 4
89 | 3 | 0.68773 | -- | 6.38 | -- | -- | -- | [6 empty] | 17 |  | 6.2 |  |  |  | 134 | 6.14 | 4
105 | 4 | -0.31741 | -- | -- | 6 | -- | -- | [6 empty] | 18 |  | 6.2 |  |  |  | 180 | 5.78 | 3
110 | 0 | -7.8824 | -- | 3.14 | -- | -- | -- | [6 empty] | 19 |  | 6.26 |  |  |  | 190 | 6.27 | 4
113 | 4 | -0.052902 | -- | 6.1 | -- | -- | -- | [6 empty] | 20 |  | 6.27 |  |  |  | 208 | 6.26 | 4
134 | 4 | 0.052902 | -- | 6.14 | -- | -- | -- | [6 empty] | 21 |  | 6.275 |  |  |  | 209 | 6.275 | 4
180 | 3 | -0.89933 | -- | 5.78 | -- | -- | -- | [6 empty] | 22 |  | 6.3 |  |  |  | 265 | 6 | 4
190 | 4 | 0.39676 | -- | 6.27 | -- | -- | -- | [6 empty] | 23 |  | 6.34 |  |  |  | 268 | 5.786 | 3
208 | 4 | 0.37031 | -- | 6.26 | -- | -- | -- | [6 empty] | 24 |  | 6.38 |  |  |  | 273 | 4.8 | 0
209 | 4 | 0.40999 | -- | 6.275 | -- | -- | -- | [6 empty] | 25 |  | 6.51 |  |  |  | 274 | 18.33 | 0
265 | 4 | -0.31741 | -- | 6 | -- | -- | -- | [6 empty] | 26 |  | 6.6 |  |  |  | 284 | 7 | 0
268 | 3 | -0.88346 | -- | 5.786 | -- | -- | -- | [6 empty] | 27 |  | 7 |  |  |  | 321 | 6.3 | 4
273 | 0 | -3.4915 | -- | -- | -- | -- | 4.8 | [6 empty] | 28 |  | 10.88 |  |  |  | 323 | 6.2 | 4
274 | 0 | 32.2966 | -- | -- | 18.33 | -- | -- | [6 empty] | 29 |  |  | 5.396 |  |  | 327 | 5.7 | 2
284 | 0 | 2.3277 | -- | -- | -- | 7 | -- | [6 empty] | 30 |  |  | 5.7 |  |  | 328 | 7 | 0
321 | 4 | 0.47612 | -- | 6.3 | -- | -- | -- | [6 empty] | 31 |  |  | 6 |  |  | 333 | 5.73 | 2
323 | 4 | 0.21161 | -- | 6.2 | -- | -- | -- | [6 empty] | 32 |  |  | 18.33 |  |  | 336 | 5.396 | 1
327 | 2 | -1.1109 | -- | -- | 5.7 | -- | -- | [6 empty] | 33 |  |  |  | 7 |  | 372 | 5.63 | 2
328 | 0 | 2.3277 | -- | 7 | -- | -- | -- | [6 empty] | 34 |  |  |  |  | 4.8 | 393 | 6.1519 | 4
333 | 2 | -1.0316 | -- | 5.73 | -- | -- | --
336 | 1 | -1.9151 | -- | -- | 5.396 | -- | --
372 | 2 | -1.2961 | -- | 5.63 | -- | -- | --
393 | 4 | 0.084379 | -- | 6.1519 | -- | -- | --

## F
Statistical summary of reported data for standard reference sample P-42 (low ionic strength constituents)
+3 Fs
+1.5 Fs
MPV
-1.5 Fs
-3 Fs
SUMMARY |  |  | Methods | [5 empty] | Statistics |  |  |  |  | SCALE
|  |  | 7 | 21 | 22 | 40 | Method Codes | [6 empty] | min | 0.039948
|  | n = | 11 | 1 | 3 | 5 |  | 07 Ion chromatography |  | MPV = | 0.11 | mg/L |  | max | 0.18005
|  | Minimum = | 0.08 | 0.1 | 0.085 | 0.098 |  | 21 Titration: electrometric |  | F-pseudosigma = | 0.023351 |  |  | major | 0.035026
|  | Maximum = | 0.48 |  | 0.131 | 0.13 |  | 22 Colorimetric |  | n = | 20 |  |  | minor | 0.035026
|  | Median = | 0.114 |  |  | 0.11 |  | 40 Ion selective electrode |  | Uh = | 0.1305
|  | F-pseudosigma = | 0.029577 |  |  | 0.0074129 |  |  |  | Lh = | 0.099
[14 empty] | Order | 7 | 21 | 22 | 40 | Lab | RV | Rating
[14 empty] | 1 | 0.08 |  |  |  | 1 | 0.093 | 3
[14 empty] | 2 | 0.093 |  |  |  | 2 | 0.112 | 4
|  |  | Method Codes | [10 empty] | 3 | 0.1 |  |  |  | 5 | 0.15 | 1
Lab | Rating | Z-value | 7 | 21 | 22 | 40 | [7 empty] | 4 | 0.105 |  |  |  | 8 | <0.5 | NR
1 | 3 | -0.72803 | 0.093 | -- | -- | -- | [7 empty] | 5 | 0.112 |  |  |  | 23 | 0.125 | 3
2 | 4 | 0.085651 | 0.112 | -- | -- | -- | [7 empty] | 6 | 0.114 |  |  |  | 25 | <0.1 | NR
5 | 1 | 1.713 | 0.15 | -- | -- | -- | [7 empty] | 7 | 0.125 |  |  |  | 33 | 0.179 | 0
8 | NR | -- | <0.5 | -- | -- | -- | [7 empty] | 8 | 0.1348 |  |  |  | 45 | 0.105 | 4
23 | 3 | 0.64238 | 0.125 | -- | -- | -- | [7 empty] | 9 | 0.15 |  |  |  | 59 | 0.11 | 4
25 | NR | -- | <0.1 | -- | -- | -- | [7 empty] | 10 | 0.179 |  |  |  | 105 | <0.20 | NR
33 | 0 | 2.955 | 0.179 | -- | -- | -- | [7 empty] | 11 | 0.48 |  |  |  | 113 | 0.08 | 2
45 | 4 | -0.21413 | 0.105 | -- | -- | -- | [7 empty] | 12 | <0.1 |  |  |  | 134 | 0.1 | 4
59 | 4 | 0 | -- | -- | -- | 0.11 | [7 empty] | 13 | <0.20 |  |  |  | 180 | 0.114 | 4
105 | NR | -- | <0.20 | -- | -- | -- | [7 empty] | 14 | <0.5 |  |  |  | 190 | 0.098 | 3
113 | 2 | -1.2848 | 0.08 | -- | -- | -- | [7 empty] | 15 | <5 |  |  |  | 273 | 0.13 | 3
134 | 4 | -0.42825 | -- | -- | -- | 0.1 | [7 empty] | 16 |  | 0.1 |  |  | 274 | 0.131 | 3
180 | 4 | 0.1713 | 0.114 | -- | -- | -- | [7 empty] | 17 |  |  | 0.085 |  | 277 | 0.1 | 4
190 | 3 | -0.5139 | -- | -- | -- | 0.098 | [7 empty] | 18 |  |  | 0.09 |  | 279 | 0.085 | 2
273 | 3 | 0.85651 | -- | -- | -- | 0.13 | [7 empty] | 19 |  |  | 0.131 |  | 284 | 0.1 | 4
274 | 3 | 0.89933 | -- | -- | 0.131 | -- | [7 empty] | 20 |  |  |  | 0.098 | 323 | 0.09 | 3
277 | 4 | -0.42825 | 0.1 | -- | -- | -- | [7 empty] | 21 |  |  |  | 0.1 | 327 | <5 | NR
279 | 2 | -1.0706 | -- | -- | 0.085 | -- | [7 empty] | 22 |  |  |  | 0.11 | 328 | 0.11 | 4
284 | 4 | -0.42825 | -- | 0.1 | -- | -- | [7 empty] | 23 |  |  |  | 0.11 | 372 | 0.48 | 0
323 | 3 | -0.85651 | -- | -- | 0.09 | -- | [7 empty] | 24 |  |  |  | 0.13 | 393 | 0.1348 | 2
327 | NR | -- | <5 | -- | -- | --
328 | 4 | 0 | -- | -- | -- | 0.11
372 | 0 | 15.8454 | 0.48 | -- | -- | --
393 | 2 | 1.0621 | 0.1348 | -- | -- | --

## K
Statistical summary of reported data for standard reference sample P-42 (low ionic strength constituents)
+3 Fs
+1.5 Fs
MPV
-1.5 Fs
-3 Fs
SUMMARY |  |  | Methods | [8 empty] | Statistics |  |  |  |  | SCALE
|  |  | 1 | 3 | 4 | 5 | 6 | 7 | 12 | Method Codes | [6 empty] | min | 0.22715
|  | n = | 9 | 1 | 11 | 1 | 2 | 3 | 2 |  | 01 Atomic absorption: direct, air |  | MPV = | 0.325 | mg/L |  | max | 0.42285
|  | Minimum = | 0.264 | 0.33 | 0.188 | 0.289 | 0.315 | 0.308 | 0.3 |  | 03 Atomic absorption: graphite furnace |  | F-pseudosigma = | 0.032617 |  |  | major | 0.048925
|  | Maximum = | 0.56 |  | 1.36 |  | 0.33 | 0.3427 | 0.315 |  | 04 Inductively coupled plasma |  | n = | 29 |  |  | minor | 0.048925
|  | Median = | 0.326 |  | 0.32 | [5 empty] | 05 Direct current plasma |  | Uh = | 0.34
|  | F-pseudosigma = | 0.031875 |  | 0.055337 | [5 empty] | 06 Inductively coupled plasma / mass spectrometry |  | Lh = | 0.296
[11 empty] | 07 Ion chromatography | [5 empty] | Order | 1 | 3 | 4 | 5 | 6 | 7 | 12 | Lab | RV | Rating
[11 empty] | 12 Flame emission | [5 empty] | 1 | 0.264 | [6 empty] | 1 | 0.34 | 4
[17 empty] | 2 | 0.29 | [6 empty] | 2 | 0.331 | 4
|  |  | Method Codes | [13 empty] | 3 | 0.297 | [6 empty] | 5 | <1.00 | NR
Lab | Rating | Z-value | 1 | 3 | 4 | 5 | 6 | 7 | 12 | [7 empty] | 4 | 0.325 | [6 empty] | 8 | 0.32 | 4
1 | 4 | 0.45989 | -- | -- | 0.34 | -- | -- | -- | -- | [7 empty] | 5 | 0.326 | [6 empty] | 23 | 0.3535 | 3
2 | 4 | 0.18395 | -- | -- | -- | -- | -- | 0.331 | -- | [7 empty] | 6 | 0.33 | [6 empty] | 25 | 0.2707 | 1
5 | NR | -- | -- | -- | <1.00 | -- | -- | -- | -- | [7 empty] | 7 | 0.34 | [6 empty] | 33 | 0.289 | 2
8 | 4 | -0.1533 | -- | -- | 0.32 | -- | -- | -- | -- | [7 empty] | 8 | 0.3535 | [6 empty] | 38 | 0.33 | 4
23 | 3 | 0.87378 | 0.3535 | -- | -- | -- | -- | -- | -- | [7 empty] | 9 | 0.56 | [6 empty] | 45 | 0.315 | 4
25 | 1 | -1.6648 | -- | -- | 0.2707 | -- | -- | -- | -- | [7 empty] | 10 |  | 0.33 | [5 empty] | 59 | 0.308 | 3
33 | 2 | -1.1037 | -- | -- | -- | 0.289 | -- | -- | -- | [7 empty] | 11 |  |  | 0.188 |  |  |  |  | 64 | 0.34 | 4
38 | 4 | 0.1533 | 0.33 | -- | -- | -- | -- | -- | -- | [7 empty] | 12 |  |  | 0.258 |  |  |  |  | 105 | <1.00 | NR
45 | 4 | -0.30659 | -- | -- | -- | -- | 0.315 | -- | -- | [7 empty] | 13 |  |  | 0.27 |  |  |  |  | 110 | 0.325 | 4
59 | 3 | -0.5212 | -- | -- | -- | -- | -- | 0.308 | -- | [7 empty] | 14 |  |  | 0.2707 |  |  |  |  | 134 | 0.326 | 4
64 | 4 | 0.45989 | 0.34 | -- | -- | -- | -- | -- | -- | [7 empty] | 15 |  |  | 0.296 |  |  |  |  | 180 | <1.67 | NR
105 | NR | -- | -- | -- | <1.00 | -- | -- | -- | -- | [7 empty] | 16 |  |  | 0.32 |  |  |  |  | 190 | 0.296 | 3
110 | 4 | 0 | 0.325 | -- | -- | -- | -- | -- | -- | [7 empty] | 17 |  |  | 0.34 |  |  |  |  | 193 | 0.33 | 4
134 | 4 | 0.030659 | 0.326 | -- | -- | -- | -- | -- | -- | [7 empty] | 18 |  |  | 0.34 |  |  |  |  | 209 | 0.376 | 1
180 | NR | -- | -- | -- | <1.67 | -- | -- | -- | -- | [7 empty] | 19 |  |  | 0.35 |  |  |  |  | 264 | 0.56 | 0
190 | 3 | -0.88911 | -- | -- | 0.296 | -- | -- | -- | -- | [7 empty] | 20 |  |  | 0.376 |  |  |  |  | 265 | 0.35 | 3
193 | 4 | 0.1533 | -- | -- | -- | -- | 0.33 | -- | -- | [7 empty] | 21 |  |  | 1.36 |  |  |  |  | 268 | 0.264 | 1
209 | 1 | 1.5636 | -- | -- | 0.376 | -- | -- | -- | -- | [7 empty] | 22 |  |  | <0.50 |  |  |  |  | 273 | 0.258 | 0
264 | 0 | 7.2049 | 0.56 | -- | -- | -- | -- | -- | -- | [7 empty] | 23 |  |  | <1 |  |  |  |  | 274 | 0.3 | 3
265 | 3 | 0.76648 | -- | -- | 0.35 | -- | -- | -- | -- | [7 empty] | 24 |  |  | <1.00 |  |  |  |  | 279 | 0.297 | 3
268 | 1 | -1.8702 | 0.264 | -- | -- | -- | -- | -- | -- | [7 empty] | 25 |  |  | <1.00 |  |  |  |  | 284 | <1 | NR
273 | 0 | -2.0542 | -- | -- | 0.258 | -- | -- | -- | -- | [7 empty] | 26 |  |  | <1.67 |  |  |  |  | 321 | 0.29 | 2
274 | 3 | -0.76648 | -- | -- | -- | -- | -- | -- | 0.3 | [7 empty] | 27 |  |  |  | 0.289 |  |  |  | 323 | <0.50 | NR
279 | 3 | -0.85845 | 0.297 | -- | -- | -- | -- | -- | -- | [7 empty] | 28 |  |  |  |  | 0.315 |  |  | 326 | 0.33 | 4
284 | NR | -- | -- | -- | <1 | -- | -- | -- | -- | [7 empty] | 29 |  |  |  |  | 0.33 |  |  | 327 | 0.27 | 1
321 | 2 | -1.0731 | 0.29 | -- | -- | -- | -- | -- | -- | [7 empty] | 30 | [5 empty] | 0.308 |  | 328 | 1.36 | 0
323 | NR | -- | -- | -- | <0.50 | -- | -- | -- | -- | [7 empty] | 31 | [5 empty] | 0.331 |  | 333 | 0.34 | 4
326 | 4 | 0.1533 | -- | 0.33 | -- | -- | -- | -- | -- | [7 empty] | 32 | [5 empty] | 0.3427 |  | 336 | 0.315 | 4
327 | 1 | -1.6863 | -- | -- | 0.27 | -- | -- | -- | -- | [7 empty] | 33 | [6 empty] | 0.3 | 372 | 0.188 | 0
328 | 0 | 31.7322 | -- | -- | 1.36 | -- | -- | -- | -- | [7 empty] | 34 | [6 empty] | 0.315 | 393 | 0.3427 | 3
333 | 4 | 0.45989 | -- | -- | 0.34 | -- | -- | -- | --
336 | 4 | -0.30659 | -- | -- | -- | -- | -- | -- | 0.315
372 | 0 | -4.2003 | -- | -- | 0.188 | -- | -- | -- | --
393 | 3 | 0.54267 | -- | -- | -- | -- | -- | 0.3427 | --

## Mg
Statistical summary of reported data for standard reference sample P-42 (low ionic strength constituents)
+3 Fs
+1.5 Fs
MPV
-1.5 Fs
-3 Fs
SUMMARY |  |  | Methods | [8 empty] | Statistics |  |  |  |  | SCALE
|  |  | 0 | 1 | 3 | 4 | 5 | 6 | 7 | 20 | Method Codes | [5 empty] | min | 0.039105
|  | n = | 1 | 5 | 1 | 13 | 1 | 2 | 3 | 1 | 00 Other |  | MPV = | 0.048 | mg/L |  | max | 0.056895
|  | Minimum = | 0.29 | 0.05 | 0.05 | 0.037 | 0.039 | 0.046 | 0.031 | 0.1 | 01 Atomic absorption: direct, air |  | F-pseudosigma = | 0.0029652 |  |  | major | 0.0044477
|  | Maximum = | Methods 1: 0.077; Methods 4: 0.11; Methods 6: 0.0475; Methods 7: 0.048 | 03 Atomic absorption: graphite furnace |  | n = | 27 |  |  | minor | 0.0044477
|  | Median = | Methods 1: 0.05; Methods 4: 0.047 | 04 Inductively coupled plasma |  | Uh = | 0.05
|  | F-pseudosigma = | Methods 1: 0.0081542; Methods 4: 0.0022239 | 05 Direct current plasma |  | Lh = | 0.046
[11 empty] | 06 Inductively coupled plasma / mass spectrometry | [5 empty] | Order | 0 | 1 | 3 | 4 | 5 | 6 | 7 | 20 | Lab | RV | Rating
[11 empty] | 07 Ion chromatography | [5 empty] | 1 | 0: 0.29 | 1 | 0.047 | 4
[11 empty] | 20 Titration: colorimetric | [5 empty] | 2 | 1: 0.05 | 2 | 0.048 | 4
|  |  | Method Codes | [13 empty] | 3 | 1: 0.05 | 5 | 0.05 | 3
Lab | Rating | Z-value | 0 | 1 | 3 | 4 | 5 | 6 | 7 | 20 | [6 empty] | 4 |  | 0.05 | [6 empty] | 8 | 0.049 | 4
1 | 4 | -0.33725 | -- | -- | -- | 0.047 | -- | -- | -- | -- | [6 empty] | 5 |  | 0.061 | [6 empty] | 23 | <0.05 | NR
2 | 4 | 0 | -- | -- | -- | -- | -- | -- | 0.048 | -- | [6 empty] | 6 |  | 0.077 | [6 empty] | 25 | 0.0468 | 4
5 | 3 | 0.6745 | -- | -- | -- | 0.05 | -- | -- | -- | -- | [6 empty] | 7 |  | <0.05 | [6 empty] | 33 | 0.039 | 0
8 | 4 | 0.33725 | -- | -- | -- | 0.049 | -- | -- | -- | -- | [6 empty] | 8 |  | <0.5 | [6 empty] | 38 | 0.061 | 0
23 | NR | -- | -- | <0.05 | -- | -- | -- | -- | -- | -- | [6 empty] | 9 |  |  | 0.05 | [5 empty] | 45 | 0.0475 | 4
25 | 4 | -0.4047 | -- | -- | -- | 0.0468 | -- | -- | -- | -- | [6 empty] | 10 |  |  |  | 0.037 |  |  |  |  | 59 | 0.031 | 0
33 | 0 | -3.0352 | -- | -- | -- | -- | 0.039 | -- | -- | -- | [6 empty] | 11 |  |  |  | 0.04 |  |  |  |  | 64 | 0.05 | 3
38 | 0 | 4.3842 | -- | 0.061 | -- | -- | -- | -- | -- | -- | [6 empty] | 12 |  |  |  | 0.044 |  |  |  |  | 105 | 0.048 | 4
45 | 4 | -0.16862 | -- | -- | -- | -- | -- | 0.0475 | -- | -- | [6 empty] | 13 |  |  |  | 0.046 |  |  |  |  | 110 | 0.046 | 3
59 | 0 | -5.7332 | -- | -- | -- | -- | -- | -- | 0.031 | -- | [6 empty] | 14 |  |  |  | 0.0465 |  |  |  |  | 134 | 0.0465 | 4
64 | 3 | 0.6745 | -- | 0.05 | -- | -- | -- | -- | -- | -- | [6 empty] | 15 |  |  |  | 0.0468 |  |  |  |  | 180 | <0.111 | NR
105 | 4 | 0 | -- | -- | -- | 0.048 | -- | -- | -- | -- | [6 empty] | 16 |  |  |  | 0.047 |  |  |  |  | 190 | 0.04 | 0
110 | 3 | -0.6745 | -- | -- | -- | 0.046 | -- | -- | -- | -- | [6 empty] | 17 |  |  |  | 0.048 |  |  |  |  | 193 | 0.046 | 3
134 | 4 | -0.50588 | -- | -- | -- | 0.0465 | -- | -- | -- | -- | [6 empty] | 18 |  |  |  | 0.048 |  |  |  |  | 209 | 0.054 | 0
180 | NR | -- | -- | -- | -- | <0.111 | -- | -- | -- | -- | [6 empty] | 19 |  |  |  | 0.049 |  |  |  |  | 264 | <0.5 | NR
190 | 0 | -2.698 | -- | -- | -- | 0.04 | -- | -- | -- | -- | [6 empty] | 20 |  |  |  | 0.05 |  |  |  |  | 265 | 0.048 | 4
193 | 3 | -0.6745 | -- | -- | -- | -- | -- | 0.046 | -- | -- | [6 empty] | 21 |  |  |  | 0.054 |  |  |  |  | 268 | 0.077 | 0
209 | 0 | 2.0235 | -- | -- | -- | 0.054 | -- | -- | -- | -- | [6 empty] | 22 |  |  |  | 0.11 |  |  |  |  | 274 | 0.1 | 0
264 | NR | -- | -- | <0.5 | -- | -- | -- | -- | -- | -- | [6 empty] | 23 |  |  |  | <0.10 |  |  |  |  | 279 | 0.05 | 3
265 | 4 | 0 | -- | -- | -- | 0.048 | -- | -- | -- | -- | [6 empty] | 24 |  |  |  | <0.10 |  |  |  |  | 284 | <1 | NR
268 | 0 | 9.7802 | -- | 0.077 | -- | -- | -- | -- | -- | -- | [6 empty] | 25 |  |  |  | <0.111 |  |  |  |  | 321 | 0.05 | 3
274 | 0 | 17.537 | -- | -- | -- | -- | -- | -- | -- | 0.1 | [6 empty] | 26 |  |  |  | <1 |  |  |  |  | 323 | <0.10 | NR
279 | 3 | 0.6745 | -- | 0.05 | -- | -- | -- | -- | -- | -- | [6 empty] | 27 |  |  |  |  | 0.039 |  |  |  | 326 | 0.05 | 3
284 | NR | -- | [14 empty] | 28 | [5 empty] | 0.046 |  |  | 327 | <0.10 | NR | bar: Method Codes 0 to Method Codes 20
321 | 3 | 0.6745 | -- | 0.05 | -- | -- | -- | -- | -- | -- | [6 empty] | 29 | [5 empty] | 0.0475 |  |  | 328 | 0.11 | 0
323 | NR | -- | -- | -- | -- | <0.10 | -- | -- | -- | -- | [6 empty] | 30 | [6 empty] | 0.031 |  | 333 | 0.044 | 2
326 | 3 | 0.6745 | -- | -- | 0.05 | -- | -- | -- | -- | -- | [6 empty] | 31 | [6 empty] | 0.0434 |  | 336 | 0.29 | 0
327 | NR | -- | -- | -- | -- | <0.10 | -- | -- | -- | -- | [6 empty] | 32 | [6 empty] | 0.048 |  | 372 | 0.037 | 0
328 | 0 | 20.9095 | -- | -- | -- | 0.11 | -- | -- | -- | -- | [6 empty] | 33 | [7 empty] | 0.1 | 393 | 0.0434 | 1
333 | 2 | -1.349 | -- | -- | -- | 0.044 | -- | -- | -- | --
336 | 0 | 81.6145 | 0.29 | -- | -- | -- | -- | -- | -- | --
372 | 0 | -3.7097 | -- | -- | -- | 0.037 | -- | -- | -- | --
393 | 1 | -1.5513 | -- | -- | -- | -- | -- | -- | 0.0434 | --

## Na
Statistical summary of reported data for standard reference sample P-42 (low ionic strength constituents)
+3 Fs
+1.5 Fs
MPV
-1.5 Fs
-3 Fs
SUMMARY |  |  | Methods | [8 empty] | Statistics |  |  |  |  | SCALE
|  |  | 1 | 3 | 4 | 5 | 6 | 7 | 12 | Method Codes | [6 empty] | min | 0.16834
|  | n = | 9 | 1 | 14 | 1 | 2 | 3 | 2 |  | 01 Atomic absorption: direct, air |  | MPV = | 0.304 | mg/L |  | max | 0.43966
|  | Minimum = | 0.3 | 0.3 | 0.164 | 0.258 | 0.283 | 0.313 | 0.25 |  | 03 Atomic absorption: graphite furnace |  | F-pseudosigma = | 0.045219 |  |  | major | 0.067828
|  | Maximum = | 0.6 |  | 0.42 |  | 0.3 | 1.445 | 0.701 |  | 04 Inductively coupled plasma |  | n = | 32 |  |  | minor | 0.067828
|  | Median = | 0.335 |  | 0.298 | [5 empty] | 05 Direct current plasma |  | Uh = | 0.3375
|  | F-pseudosigma = | 0.042995 |  | 0.035582 | [5 empty] | 06 Inductively coupled plasma / mass spectrometry |  | Lh = | 0.2765
[11 empty] | 07 Ion chromatography | [5 empty] | Order | 1 | 3 | 4 | 5 | 6 | 7 | 12 | Lab | RV | Rating
[11 empty] | 12 Flame emission | [5 empty] | 1 | 0.3 | [6 empty] | 1 | 0.369 | 2
[17 empty] | 2 | 0.301 | [6 empty] | 2 | 0.313 | 4
|  |  | Method Codes | [13 empty] | 3 | 0.31 | [6 empty] | 5 | 0.27 | 3
Lab | Rating | Z-value | 1 | 3 | 4 | 5 | 6 | 7 | 12 | [7 empty] | 4 | 0.31 | [6 empty] | 8 | 0.296 | 4
1 | 2 | 1.4375 | -- | -- | 0.369 | -- | -- | -- | -- | [7 empty] | 5 | 0.335 | [6 empty] | 23 | 0.335 | 3
2 | 4 | 0.19903 | -- | -- | -- | -- | -- | 0.313 | -- | [7 empty] | 6 | 0.34 | [6 empty] | 25 | 0.2554 | 2
5 | 3 | -0.7519 | -- | -- | 0.27 | -- | -- | -- | -- | [7 empty] | 7 | 0.368 | [6 empty] | 33 | 0.258 | 2
8 | 4 | -0.17692 | -- | -- | 0.296 | -- | -- | -- | -- | [7 empty] | 8 | 0.404 | [6 empty] | 38 | 0.31 | 4
23 | 3 | 0.68556 | 0.335 | -- | -- | -- | -- | -- | -- | [7 empty] | 9 | 0.6 | [6 empty] | 45 | 0.283 | 4
25 | 2 | -1.0748 | -- | -- | 0.2554 | -- | -- | -- | -- | [7 empty] | 10 |  | 0.3 | [5 empty] | 59 | 1.445 | 0
33 | 2 | -1.0173 | -- | -- | -- | 0.258 | -- | -- | -- | [7 empty] | 11 |  |  | 0.164 |  |  |  |  | 64 | 0.31 | 4
38 | 4 | 0.13269 | 0.31 | -- | -- | -- | -- | -- | -- | [7 empty] | 12 |  |  | 0.224 |  |  |  |  | 105 | 0.313 | 4
45 | 4 | -0.46441 | -- | -- | -- | -- | 0.283 | -- | -- | [7 empty] | 13 |  |  | 0.2554 |  |  |  |  | 110 | 0.3 | 4
59 | 0 | 25.2329 | -- | -- | -- | -- | -- | 1.445 | -- | [7 empty] | 14 |  |  | 0.26 |  |  |  |  | 134 | 0.301 | 4
64 | 4 | 0.13269 | 0.31 | -- | -- | -- | -- | -- | -- | [7 empty] | 15 |  |  | 0.269 |  |  |  |  | 180 | 0.307 | 4
105 | 4 | 0.19903 | -- | -- | 0.313 | -- | -- | -- | -- | [7 empty] | 16 |  |  | 0.27 |  |  |  |  | 190 | 0.269 | 3
110 | 4 | -0.088459 | 0.3 | -- | -- | -- | -- | -- | -- | [7 empty] | 17 |  |  | 0.296 |  |  |  |  | 193 | 0.3 | 4
134 | 4 | -0.066344 | 0.301 | -- | -- | -- | -- | -- | -- | [7 empty] | 18 |  |  | 0.3 |  |  |  |  | 209 | 0.308 | 4
180 | 4 | 0.066344 | -- | -- | 0.307 | -- | -- | -- | -- | [7 empty] | 19 |  |  | 0.3 |  |  |  |  | 264 | 0.6 | 0
190 | 3 | -0.77402 | -- | -- | 0.269 | -- | -- | -- | -- | [7 empty] | 20 |  |  | 0.307 |  |  |  |  | 265 | 0.3 | 4
193 | 4 | -0.088459 | -- | -- | -- | -- | 0.3 | -- | -- | [7 empty] | 21 |  |  | 0.308 |  |  |  |  | 268 | 0.404 | 0
209 | 4 | 0.088459 | -- | -- | 0.308 | -- | -- | -- | -- | [7 empty] | 22 |  |  | 0.313 |  |  |  |  | 273 | 0.164 | 0
264 | 0 | 6.546 | 0.6 | -- | -- | -- | -- | -- | -- | [7 empty] | 23 |  |  | 0.369 |  |  |  |  | 274 | 0.25 | 2
265 | 4 | -0.088459 | -- | -- | 0.3 | -- | -- | -- | -- | [7 empty] | 24 |  |  | 0.42 |  |  |  |  | 279 | 0.368 | 2
268 | 0 | 2.2115 | 0.404 | -- | -- | -- | -- | -- | -- | [7 empty] | 25 |  |  | <0.30 |  |  |  |  | 284 | <1 | NR
273 | 0 | -3.0961 | -- | -- | 0.164 | -- | -- | -- | -- | [7 empty] | 26 |  |  | <1 |  |  |  |  | 321 | 0.34 | 3
274 | 2 | -1.1942 | -- | -- | -- | -- | -- | -- | 0.25 | [7 empty] | 27 |  |  |  | 0.258 |  |  |  | 323 | <0.30 | NR
279 | 2 | 1.4153 | 0.368 | -- | -- | -- | -- | -- | -- | [7 empty] | 28 |  |  |  |  | 0.283 |  |  | 326 | 0.3 | 4
284 | NR | -- | -- | -- | <1 | -- | -- | -- | -- | [7 empty] | 29 |  |  |  |  | 0.3 |  |  | 327 | 0.26 | 3
321 | 3 | 0.79613 | 0.34 | -- | -- | -- | -- | -- | -- | [7 empty] | 30 | [5 empty] | 0.313 |  | 328 | 0.42 | 0
323 | NR | -- | -- | -- | <0.30 | -- | -- | -- | -- | [7 empty] | 31 | [5 empty] | 0.3262 |  | 333 | 0.3 | 4
326 | 4 | -0.088459 | -- | 0.3 | -- | -- | -- | -- | -- | [7 empty] | 32 | [5 empty] | 1.445 |  | 336 | 0.701 | 0
327 | 3 | -0.97305 | -- | -- | 0.26 | -- | -- | -- | -- | [7 empty] | 33 | [6 empty] | 0.25 | 372 | 0.224 | 1
328 | 0 | 2.5653 | -- | -- | 0.42 | -- | -- | -- | -- | [7 empty] | 34 | [6 empty] | 0.701 | 393 | 0.3262 | 4
333 | 4 | -0.088459 | -- | -- | 0.3 | -- | -- | -- | --
336 | 0 | 8.7796 | -- | -- | -- | -- | -- | -- | 0.701
372 | 1 | -1.7692 | -- | -- | 0.224 | -- | -- | -- | --
393 | 4 | 0.49095 | -- | -- | -- | -- | -- | 0.3262 | --

## pH
Statistical summary of reported data for standard reference sample P-42 (low ionic strength constituents)
+3 Fs
+1.5 Fs
MPV
-1.5 Fs
-3 Fs
SUMMARY |  |  | Methods |  |  |  | Statistics |  |  |  | SCALE
|  |  | 40 | 41 | Method Codes | [5 empty] | min | 3.5442
|  | n = | 3 | 34 |  | 40 Ion selective electrode |  | MPV = | 3.9 |  | max | 4.2558
|  | Minimum = | 3.38 | 3.54 |  | 41 Electrometric |  | F-pseudosigma = | 0.11861 |  | major | 0.17791
|  | Maximum = | 4.04 | 4.3 |  |  |  | Rating criterion = | 0.195 |  | minor | 0.17791
|  | Median = |  | 3.9 |  |  |  | n = | 37
|  | F-pseudosigma = |  | 0.119 |  |  |  | Uh = | 3.96
[8 empty] | Lh = | 3.8 |  | Order | 40 | 41 | Lab | RV | Rating
[11 empty] | 1 | 3.38 |  | 1 | 3.96 | 4
[11 empty] | 2 | 4.02 |  | 2 | 3.932 | 4
|  |  | Method Codes | [7 empty] | 3 | 4.04 |  | 5 | 3.59 | 1
Lab | Rating | Z-value | 40 | 41 | [6 empty] | 4 |  | 3.54 | 8 | 3.89 | 4
1 | 4 | 0.30769 | -- | 3.96 | [6 empty] | 5 |  | 3.56 | 23 | 3.96 | 4
2 | 4 | 0.1641 | -- | 3.932 | [6 empty] | 6 |  | 3.58 | 25 | 3.83 | 4
5 | 1 | -1.5897 | -- | 3.59 | [6 empty] | 7 |  | 3.59 | 30 | 4 | 3
8 | 4 | -0.051282 | -- | 3.89 | [6 empty] | 8 |  | 3.6 | 33 | 3.88 | 4
23 | 4 | 0.30769 | -- | 3.96 | [6 empty] | 9 |  | 3.63 | 38 | 3.9 | 4
25 | 4 | -0.35897 | -- | 3.83 | [6 empty] | 10 |  | 3.74 | 45 | 3.78 | 3
30 | 3 | 0.51282 | -- | 4 | [6 empty] | 11 |  | 3.78 | 59 | 3.82 | 4
33 | 4 | -0.10256 | -- | 3.88 | [6 empty] | 12 |  | 3.8 | 64 | 3.83 | 4
38 | 4 | 0 | -- | 3.9 | [6 empty] | 13 |  | 3.82 | 89 | 4.04 | 3
45 | 3 | -0.61538 | -- | 3.78 | [6 empty] | 14 |  | 3.83 | 105 | 4 | 3
59 | 4 | -0.41026 | -- | 3.82 | [6 empty] | 15 |  | 3.83 | 110 | 3.96 | 4
64 | 4 | -0.35897 | -- | 3.83 | [6 empty] | 16 |  | 3.84 | 113 | 3.9 | 4
89 | 3 | 0.71795 | 4.04 | -- | [6 empty] | 17 |  | 3.87 | 134 | 3.92 | 4
105 | 3 | 0.51282 | -- | 4 | [6 empty] | 18 |  | 3.88 | 180 | 4.02 | 3
110 | 4 | 0.30769 | -- | 3.96 | [6 empty] | 19 |  | 3.89 | 190 | 3.6 | 1
113 | 4 | 0 | -- | 3.9 | [6 empty] | 20 |  | 3.9 | 193 | 3.93 | 4
134 | 4 | 0.10256 | -- | 3.92 | [6 empty] | 21 |  | 3.9 | 209 | 4.02 | 3
180 | 3 | 0.61538 | 4.02 | -- | [6 empty] | 22 |  | 3.9 | 264 | 4.3 | 0
190 | 1 | -1.5385 | -- | 3.6 | [6 empty] | 23 |  | 3.91 | 268 | 3.58 | 1
193 | 4 | 0.15385 | -- | 3.93 | [6 empty] | 24 |  | 3.92 | 273 | 3.63 | 2
209 | 3 | 0.61538 | -- | 4.02 | [6 empty] | 25 |  | 3.93 | 274 | 3.74 | 3
264 | 0 | 2.0513 | -- | 4.3 | [6 empty] | 26 |  | 3.93 | 277 | 4 | 3
268 | 1 | -1.641 | -- | 3.58 | [6 empty] | 27 |  | 3.932 | 279 | 3.38 | 0
273 | 2 | -1.3846 | -- | 3.63 | [6 empty] | 28 |  | 3.934 | 284 | 4 | 3
274 | 3 | -0.82051 | -- | 3.74 | [6 empty] | 29 |  | 3.96 | 321 | 3.87 | 4
277 | 3 | 0.51282 | -- | 4 | [6 empty] | 30 |  | 3.96 | 323 | 3.9 | 4
279 | 0 | -2.6667 | 3.38 | -- | [6 empty] | 31 |  | 3.96 | 326 | 3.91 | 4
284 | 3 | 0.51282 | -- | 4 | [6 empty] | 32 |  | 4 | 327 | 3.8 | 3
321 | 4 | -0.15385 | -- | 3.87 | [6 empty] | 33 |  | 4 | 328 | 3.56 | 1
323 | 4 | 0 | -- | 3.9 | [6 empty] | 34 |  | 4 | 333 | 3.84 | 4
326 | 4 | 0.051282 | -- | 3.91 | [6 empty] | 35 |  | 4 | 336 | 3.54 | 1
327 | 3 | -0.51282 | -- | 3.8 | [6 empty] | 36 |  | 4.02 | 372 | 3.93 | 4
328 | 1 | -1.7436 | -- | 3.56 | [6 empty] | 37 |  | 4.3 | 394 | 3.934 | 4
333 | 4 | -0.30769 | -- | 3.84
336 | 1 | -1.8462 | -- | 3.54
372 | 4 | 0.15385 | -- | 3.93
394 | 4 | 0.17436 | -- | 3.934

## PO4
Statistical summary of reported data for standard reference sample P-42 (low ionic strength constituents)
+3 Fs
+1.5 Fs
MPV
-1.5 Fs
-3 Fs
SUMMARY |  |  | Methods |  |  |  |  | Statistics |  |  |  |  | SCALE
|  |  | 0 | 7 | 22 | Method Codes | [6 empty] | min | 0.0058249
|  | n = | 1 | 3 | 18 |  | 00 Other |  | MPV = | 0.05475 | mg/L |  | max | 0.10368
|  | Minimum = | 0.12 | 0.055 | 0.04 |  | 07 Ion chromatography |  | F-pseudosigma = | 0.016308 |  |  | major | 0.024463
|  | Maximum = |  | 0.328 | 0.609 |  | 22 Colorimetric |  | n = | 22 |  |  | minor | 0.024463
|  | Median = |  |  | 0.053 |  |  |  | Uh = | 0.072
|  | F-pseudosigma = |  |  | 0.0064492 |  |  |  | Lh = | 0.05
[13 empty] | Order | 0 | 7 | 22 | Lab | RV | Rating
[13 empty] | 1 | 0.12 |  |  | 5 | 0.057 | 4
[13 empty] | 2 |  | 0.055 |  | 8 | <0.2 | NR
|  |  | Method Codes | [9 empty] | 3 |  | 0.061 |  | 23 | 0.04 | 3
Lab | Rating | Z-value | 0 | 7 | 22 | [7 empty] | 4 |  | 0.328 |  | 25 | 0.048 | 4
5 | 4 | 0.13797 | -- | -- | 0.057 | [7 empty] | 5 |  | <0.2 |  | 33 | 0.079 | 2
8 | NR | -- | -- | <0.2 | -- | [7 empty] | 6 |  |  | 0.04 | 38 | 0.054 | 4
23 | 3 | -0.90444 | -- | -- | 0.04 | [7 empty] | 7 |  |  | 0.041 | 45 | 0.061 | 4
25 | 4 | -0.4139 | -- | -- | 0.048 | [7 empty] | 8 |  |  | 0.044 | 59 | 0.328 | 0
33 | 2 | 1.487 | -- | -- | 0.079 | [7 empty] | 9 |  |  | 0.048 | 64 | 0.169 | 0
38 | 4 | -0.045989 | -- | -- | 0.054 | [7 empty] | 10 |  |  | 0.0483 | 89 | 0.0483 | 4
45 | 4 | 0.38324 | -- | 0.061 | -- | [7 empty] | 11 |  |  | 0.05 | 105 | 0.041 | 3
59 | 0 | 16.7552 | -- | 0.328 | -- | [7 empty] | 12 |  |  | 0.05 | 113 | 0.056 | 4
64 | 0 | 7.0056 | -- | -- | 0.169 | [7 empty] | 13 |  |  | 0.052 | 134 | 0.0545 | 4
89 | 4 | -0.3955 | -- | -- | 0.0483 | [7 empty] | 14 |  |  | 0.052 | 180 | 0.044 | 3
105 | 3 | -0.84313 | -- | -- | 0.041 | [7 empty] | 15 |  |  | 0.054 | 190 | 0.052 | 4
113 | 4 | 0.076648 | -- | -- | 0.056 | [7 empty] | 16 |  |  | 0.0545 | 264 | 0.052 | 4
134 | 4 | -0.01533 | -- | -- | 0.0545 | [7 empty] | 17 |  |  | 0.056 | 274 | 0.072 | 2
180 | 3 | -0.65917 | -- | -- | 0.044 | [7 empty] | 18 |  |  | 0.057 | 284 | <0.1 | NR
190 | 4 | -0.16863 | -- | -- | 0.052 | [7 empty] | 19 |  |  | 0.057 | 321 | 0.609 | 0
264 | 4 | -0.16863 | -- | -- | 0.052 | [7 empty] | 20 |  |  | 0.072 | 323 | 0.057 | 4
274 | 2 | 1.0577 | -- | -- | 0.072 | [7 empty] | 21 |  |  | 0.079 | 327 | 0.05 | 4
284 | NR | -- | -- | -- | <0.1 | [7 empty] | 22 |  |  | 0.169 | 328 | 0.12 | 0
321 | 0 | 33.9856 | -- | -- | 0.609 | [7 empty] | 23 |  |  | 0.609 | 333 | 0.055 | 4
323 | 4 | 0.13797 | -- | -- | 0.057 | [7 empty] | 24 |  |  | <0.1 | 372 | 0.05 | 4
327 | 4 | -0.29126 | -- | -- | 0.05
328 | 0 | 4.001 | 0.12 | -- | --
333 | 4 | 0.01533 | -- | 0.055 | --
372 | 4 | -0.29126 | -- | -- | 0.05

## SO4
Statistical summary of reported data for standard reference sample P-42 (low ionic strength constituents)
+3 Fs
+1.5 Fs
MPV
-1.5 Fs
-3 Fs
SUMMARY |  |  | Methods |  |  |  |  | Statistics |  |  |  |  | SCALE
|  |  | 7 | 21 | 51 | Method Codes | [6 empty] | min | -0.1964
|  | n = | 22 | 1 | 3 |  | 07 Ion chromatography |  | MPV = | 0.344 | mg/L |  | max | 0.8844
|  | Minimum = | 0.09 | 1.51 | 0.138 |  | 21 Titration: electrometric |  | F-pseudosigma = | 0.18013 |  |  | major | 0.2702
|  | Maximum = | 1.78 |  | 5 |  | 51 Turbidimetric |  | n = | 26 |  |  | minor | 0.2702
|  | Median = | 0.344 | [5 empty] | Uh = | 0.523
|  | F-pseudosigma = | 0.11119 | [5 empty] | Lh = | 0.28
[13 empty] | Order | 7 | 21 | 51 | Lab | RV | Rating
[13 empty] | 1 | 0.09 |  |  | 1 | 0.25 | 3
[13 empty] | 2 | 0.25 |  |  | 2 | 0.283 | 4
|  |  | Method Codes | [9 empty] | 3 | 0.2661 |  |  | 5 | 0.58 | 2
Lab | Rating | Z-value | 7 | 21 | 51 | [7 empty] | 4 | 0.27 |  |  | 8 | 0.95 | 0
1 | 3 | -0.52184 | 0.25 | -- | -- | [7 empty] | 5 | 0.28 |  |  | 25 | <5.0 | NR
2 | 4 | -0.33864 | 0.283 | -- | -- | [7 empty] | 6 | 0.28 |  |  | 33 | 0.314 | 4
5 | 2 | 1.3101 | 0.58 | -- | -- | [7 empty] | 7 | 0.283 |  |  | 38 | 0.09 | 2
8 | 0 | 3.3642 | 0.95 | -- | -- | [7 empty] | 8 | 0.3 |  |  | 45 | 0.687 | 1
25 | NR | -- | <5.0 | -- | -- | [7 empty] | 9 | 0.314 |  |  | 59 | 0.367 | 4
33 | 4 | -0.16654 | 0.314 | -- | -- | [7 empty] | 10 | 0.32 |  |  | 64 | 0.28 | 4
38 | 2 | -1.4101 | 0.09 | -- | -- | [7 empty] | 11 | 0.321 |  |  | 89 | 0.523 | 3
45 | 1 | 1.9041 | 0.687 | -- | -- | [7 empty] | 12 | 0.367 |  |  | 105 | <1.00 | NR
59 | 4 | 0.12768 | 0.367 | -- | -- | [7 empty] | 13 | 0.37 |  |  | 110 | 0.3 | 4
64 | 4 | -0.35529 | 0.28 | -- | -- | [7 empty] | 14 | 0.394 |  |  | 134 | 0.28 | 4
89 | 3 | 0.99371 | 0.523 | -- | -- | [7 empty] | 15 | 0.404 |  |  | 180 | 0.37 | 4
105 | NR | -- | <1.00 | -- | -- | [7 empty] | 16 | 0.42 |  |  | 190 | 0.42 | 4
110 | 4 | -0.24426 | 0.3 | -- | -- | [7 empty] | 17 | 0.43 |  |  | 193 | <0.5 | NR
134 | 4 | -0.35529 | 0.28 | -- | -- | [7 empty] | 18 | 0.523 |  |  | 208 | <3 | NR
180 | 4 | 0.14434 | 0.37 | -- | -- | [7 empty] | 19 | 0.58 |  |  | 209 | 0.404 | 4
190 | 4 | 0.42191 | 0.42 | -- | -- | [7 empty] | 20 | 0.687 |  |  | 264 | 1.51 | 0
193 | NR | -- | <0.5 | -- | -- | [7 empty] | 21 | 0.95 |  |  | 265 | 0.32 | 4
208 | NR | -- | <3 | -- | -- | [7 empty] | 22 | 1.78 |  |  | 268 | 0.394 | 4
209 | 4 | 0.33309 | 0.404 | -- | -- | [7 empty] | 23 | <0.5 |  |  | 274 | 0.25 | 3
264 | 0 | 6.473 | -- | 1.51 | -- | [7 empty] | 24 | <1.00 |  |  | 284 | 5 | 0
265 | 4 | -0.13323 | 0.32 | -- | -- | [7 empty] | 25 | <3 |  |  | 321 | 0.43 | 4
268 | 4 | 0.27757 | 0.394 | -- | -- | [7 empty] | 26 | <5.0 |  |  | 323 | 0.321 | 4
274 | 3 | -0.52184 | -- | -- | 0.25 | [7 empty] | 27 |  | 1.51 |  | 327 | <5 | NR
284 | 0 | 25.8475 | -- | -- | 5 | [7 empty] | 28 |  |  | 0.138 | 333 | 0.27 | 4
321 | 4 | 0.47742 | 0.43 | -- | -- | [7 empty] | 29 |  |  | 0.25 | 336 | 0.138 | 2
323 | 4 | -0.12768 | 0.321 | -- | -- | [7 empty] | 30 |  |  | 5 | 372 | 1.78 | 0
327 | NR | -- | -- | -- | <5 | [7 empty] | 31 |  |  | <5 | 393 | 0.2661 | 4
333 | 4 | -0.41081 | 0.27 | -- | --
336 | 2 | -1.1436 | -- | -- | 0.138
372 | 0 | 7.9719 | 1.78 | -- | --
393 | 4 | -0.43246 | 0.2661 | -- | --

## SpCond
Statistical summary of reported data for standard reference sample P-42 (low ionic strength constituents)
+3 Fs
+1.5 Fs
MPV
-1.5 Fs
-3 Fs
SUMMARY |  |  | Methods |  |  |  |  | Statistics |  |  |  |  | SCALE
|  |  | 40 | 41 | 51 | Method Codes | [6 empty] | min | 52.253
|  | n = | 2 | 29 | 1 |  | 40 Ion selective electrode |  | MPV = | 63.15 | mS/cm |  | max | 74.047
|  | Minimum = | 64.9 | 54.56 | 101.7 |  | 41 Electrometric |  | F-pseudosigma = | 3.6323 |  |  | major | 5.4485
|  | Maximum = | 65 | 70 |  |  | 51 Turbidimetric |  | n = | 32 |  |  | minor | 5.4485
|  | Median = |  | 63 |  |  |  |  | Uh = | 65.2
|  | F-pseudosigma = |  | 3.6323 |  |  |  |  | Lh = | 60.3
[13 empty] | Order | 40 | 41 | 51 | Lab | RV | Rating
[13 empty] | 1 | 64.9 |  |  | 1 | 65.3 | 3
[13 empty] | 2 | 65 |  |  | 2 | 65.7 | 3
|  |  | Method Codes | [9 empty] | 3 |  | 54.56 |  | 5 | 63 | 4
Lab | Rating | Z-value | 40 | 41 | 51 | [7 empty] | 4 |  | 56 |  | 8 | 63.7 | 4
1 | 3 | 0.59191 | -- | 65.3 | -- | [7 empty] | 5 |  | 58 |  | 23 | 101.7 | 0
2 | 3 | 0.70203 | -- | 65.7 | -- | [7 empty] | 6 |  | 58.3 |  | 25 | 64 | 4
5 | 4 | -0.041296 | -- | 63 | -- | [7 empty] | 7 |  | 58.7 |  | 33 | 54.56 | 0
8 | 4 | 0.15142 | -- | 63.7 | -- | [7 empty] | 8 |  | 60 |  | 38 | 64.7 | 4
23 | 0 | 10.6131 | -- | -- | 101.7 | [7 empty] | 9 |  | 60 |  | 45 | 62 | 4
25 | 4 | 0.23401 | -- | 64 | -- | [7 empty] | 10 |  | 60.2 |  | 59 | 62.9 | 4
33 | 0 | -2.3649 | -- | 54.56 | -- | [7 empty] | 11 |  | 60.4 |  | 64 | 68.9 | 1
38 | 4 | 0.42672 | -- | 64.7 | -- | [7 empty] | 12 |  | 60.4 |  | 89 | 65 | 4
45 | 4 | -0.3166 | -- | 62 | -- | [7 empty] | 13 |  | 61.5 |  | 105 | 60.4 | 3
59 | 4 | -0.068827 | -- | 62.9 | -- | [7 empty] | 14 |  | 62 |  | 110 | 60.4 | 3
64 | 1 | 1.583 | -- | 68.9 | -- | [7 empty] | 15 |  | 62.9 |  | 113 | 65.1 | 3
89 | 4 | 0.50932 | 65 | -- | -- | [7 empty] | 16 |  | 63 |  | 134 | 64.9 | 4
105 | 3 | -0.75709 | -- | 60.4 | -- | [7 empty] | 17 |  | 63 |  | 180 | 63 | 4
110 | 3 | -0.75709 | -- | 60.4 | -- | [7 empty] | 18 |  | 63 |  | 190 | 60.2 | 3
113 | 3 | 0.53685 | -- | 65.1 | -- | [7 empty] | 19 |  | 63.3 |  | 193 | 60 | 3
134 | 4 | 0.48179 | 64.9 | -- | -- | [7 empty] | 20 |  | 63.7 |  | 264 | 56 | 1
180 | 4 | -0.041296 | -- | 63 | -- | [7 empty] | 21 |  | 64 |  | 268 | 58.3 | 2
190 | 3 | -0.81215 | -- | 60.2 | -- | [7 empty] | 22 |  | 64.7 |  | 273 | 65.5 | 3
193 | 3 | -0.86721 | -- | 60 | -- | [7 empty] | 23 |  | 65 |  | 274 | 63.3 | 4
264 | 1 | -1.9684 | -- | 56 | -- | [7 empty] | 24 |  | 65.1 |  | 284 | 63 | 4
268 | 2 | -1.3352 | -- | 58.3 | -- | [7 empty] | 25 |  | 65.3 |  | 321 | 58.7 | 2
273 | 3 | 0.64697 | -- | 65.5 | -- | [7 empty] | 26 |  | 65.5 |  | 323 | 67 | 2
274 | 4 | 0.041296 | -- | 63.3 | -- | [7 empty] | 27 |  | 65.7 |  | 326 | 60 | 3
284 | 4 | -0.041296 | -- | 63 | -- | [7 empty] | 28 |  | 67 |  | 327 | 58 | 2
321 | 2 | -1.2251 | -- | 58.7 | -- | [7 empty] | 29 |  | 68 |  | 328 | 70 | 1
323 | 2 | 1.0599 | -- | 67 | -- | [7 empty] | 30 |  | 68.9 |  | 333 | 65 | 4
326 | 3 | -0.86721 | -- | 60 | -- | [7 empty] | 31 |  | 70 |  | 372 | 68 | 2
327 | 2 | -1.4178 | -- | 58 | -- | [7 empty] | 32 |  |  | 101.7 | 394 | 61.5 | 4
328 | 1 | 1.8858 | -- | 70 | --
333 | 4 | 0.50932 | -- | 65 | --
372 | 2 | 1.3352 | -- | 68 | --
394 | 4 | -0.45426 | -- | 61.5 | --