Building 510C

P.O. Box 5000

Upton, NY 11973-5000

Phone 516 344-6102

Fax 516 344-3253

sivertz@bnl.gov

managed by Brookhaven Science Associates
for the U.S. Department of Energy

Memo

date:		July 27, 2001

to:		Yousef Makdisi

from:		Michael Sivertz

subject:	PHENIX Drift Chamber gas piping modifications

This note gives the plan for adding a pair of gas pipes to the existing DC/PC gas system to allow for separate controls of the DC Gas System in the Gas Mixing House (GMH).

Motivation:

Currently the DC/PC Gas System uses a single pipe to bring the Argon/Ethane 50/50 gas mixture from the GMH to the IR where the pipes to the East Carriage and West Carriage are separated.  The final division between DC and PC is accomplished in the Carriage Racks.

The DC group would like to allow for bubbling the DC operating gas through alcohol or water, while the PC group does not want to have alcohol or water in its operating gas.  In order to allow for different gas mixtures, the DC and PC gas pipes must be separated in the GMH.

A pair of flexible stainless steel pipes was run from the GMH to the IR.  One pipe would be used for each DC on the East or West Carriage.  The pipes are 3/4" in diameter.  These would replace the 1.5" diameter pipe now being used for both the DC and PC.  The existing pipe would be used solely by the PC.

Future gas system modifications involving the addition of a bubbler, new pressure regulators and pressure transducers will be addressed in another review at a future date.

Safety Precautions:

All High Voltage to the DC and PC shall be turned off before beginning this work.

All subsequent work in the IR and GMH will be monitored with the use of a flammable gas sensor to ensure that there is no flammable gas present during the work.  The PHENIX flammable gas safety system will be operational in the GMH and IR throughout the duration of this work.

Inerting the gas pipes:

In order to connect the new pipes, the existing pipes must be purged of flammable gas before cutting into them.  The volume of the pipes is less than 2.5 cubic feet (CF).  Flowing at a rate of 40 cubic feet per hour gives a complete exchange of gas every six minutes, or eight exchanges in 50 minutes.  The inert gas would be allowed to flow into the DC and PC, effectively purging the pipes into the chambers.  After eight exchanges the gas in the pipes will be non-flammable.  Solenoid valves in the Carriage Racks will be closed sealing off the East and West DC and PC.  Solenoid valves in the GMH will be closed stopping the gas flow.  At this point it is safe to work on the gas pipes.

Modifications in the GMH:

We will cut into the gas supply piping in the DC/PC Gas Rack in the GMH.  Existing connections will be cut using pipe cutters.  A junction will be added allowing the new DC pipes to be connected.  A single pipe will bring the gas from the existing DC/PC Rack in the GMH to a new DC Bubbler Rack, also in the GMH.  There the gas pipe will be divided again to allow for separate control of the East and West DC gas supplies.  The gas pressure will be regulated separately as well with the addition of three new regulators.  A four-way valve in each of the East and West DC gas pipes will allow for the addition of a pair of bubblers in the future.  All new connections will be made with Swagelok fittings making cleaning of the pipes unnecessary.

Modifications in the IR:

In the IR, there are separate flow meters in the Carriage Racks for the DC and PC.  The existing pipe supplying these flow meters will be replaced by a pair of pipes; one coming from the new flexible stainless steel pipe, and one from the end of the existing pipe.  All connections will be made with 3/4" Swagelok fittings.

Restarting flammable gas flow:

Before starting to flow flammable gas, it is important to flash the new pipes with inert gas to assure that no oxygen will be introduced into the DC or PC.  We will begin flowing inert gas to the new and old pipes by opening up the solenoid valves in the GMH.  The new pipe connections to the Carriage Gas Racks will be left loose so that the inert gas leaks out at the carriages.  The air in the pipes will be flushed into the IR in this process.  Flowing at a rate of 40 SCFH, we will have purged the pipes of oxygen in 50 minutes.  At that time the connections at the Carriage Gas Racks will be made gas tight.  Then the solenoid valves to the DC and PC detectors will be opened and we will begin flowing operating gas again.

All new gas connections will be checked with a flammable gas detector to assure that there are no leaks in the new or old piping.