May 9, 2008

Division of Dockets Management (HFA-305)

Food and Drug Administration

5630 Fishers lane, Rm 1061

Rockville, MD 20852

Comments on Draft Guidance for Industry: Validation of Growth-Based Rapid Microbiological Methods for Sterility Testing of Cellular and Gene Therapy Products

- The Draft Guidance, under C. What Elements Should Be Considered to Validate an RMM? pg. 5, states “We recommend that validation include comparison studies between the RMM and the method described in 21 CFR 610.12 to demonstrate equivalence as required under 21 CFR 610.9.  For this purpose, equivalence may be demonstrated by evidence that the performance of the RMM assay provides assurances equal to or greater than assurances provided by the method in 21 CFR 610.12.  Because validation studies of microbiological methods encompass a large degree of variability, the concept of the RMM being “equal to” or “greater than” the traditional method is demonstrated in parallel comparisons to the traditional method using as many identical parameters as possible to claim equivalence.”

# Comment & Justification

The development of cellular and gene therapy based products often involves the use of materials that are in limited supply. In the development of a new product, appropriate test methods for assuring sterility are developed.  A RMM used during the development process and validated for use throughout the product development cycle, following the growth promotion and bacteriostasis and fungistasis guidelines in 21 CFR 610.12 or USP <71> sterility test methods, should be considered to provide assurances of sterility equal to 21 CFR 610.12.  For the development of a new product, requiring a parallel validation with 21 CFR 610.12 to establish assurances of sterility that are equal to or greater may create a burden on resources.

- The Draft Guidance, under D. How Do I Select a Panel of Appropriate Challenge Microorganisms for Validating an RMM? pg 7, states “We recommend that such studies include the initial challenge test panel of organisms indicated in Table 2, cultured with the indicated media, or a recognized equivalent media, at the prescribed temperatures.”

# Comments & Justification

The adjective ‘recognized’ used to describe ‘equivalent media’ is confusing.  It implies acceptance by a regulatory entity.  Further description of ‘recognized’ is needed for clarification or removal of the word should be considered.

This section clearly states the recommendations for the conditions under which validation should be performed.  However, for clarification, specifically pertaining to ‘prescribed temperatures’, although it is accepted that incubations at 20-25 oC and 30-35 oC will promote growth of the widest variety of microorganisms, incubation at a single temperature within the range of 20-35 oC, should be allowed when validated as suitable for recovery of challenge, bioburden and environmental isolates representing contaminants associated with the product being tested.

In Table 2, in the row associated with Organisms that grow well under aerobic conditions, Propionibacterium acnes is listed under Microorganism.  Since this is an anaerobic, non-sporeforming species, it should be moved and listed under Organisms that grow well under anaerobic conditions.

On page 7 under the recommendations of categories of microorganisms to include in the challenge panel, it states “fastidious microorganisms for the RMM”.  We propose this statement is expanded to state “fastidious microorganisms for the RMM that are relevant to the product and process”.  Rarely is a RMM sold specifically for a clearly defined application.  Manufacturers of RMMs list general limitations associated with certain microorganisms in order to inform customers for whom the limitation may have significance.  Although this guidance contains general recommendations, often users interpret them as required.  Expanding this statement will clarify the recommendation thereby avoiding any confusion in establishing the appropriateness of the RMM for specific applications.

- Comment & Justification

Routinely, the FDA Guideline on General Principles of Process Validation is referenced by companies and professional organizations when describing approaches to process validation (IQ, OQ, PQ).  This document is not listed in the Draft Guidance references.  Although it is not listed, does it have relevance to the validation of a RMM?

Sincerely,

Elizabeth J. Landon

Staff Regulatory Affairs Specialist

919-620-2329

elizabeth.landon@na.biomerieux.com