Spice Model

Model of 8 GeV Style BPM includes 50 Ohm Terminating Resistor and 500 ft. of 3/8 ‘’ Heliax cable . The cable is terminated in 50 Ohms. The BPM model includes the plate-plate coupling capacitance and the plate capacitance. The current source is shunted with 1 MOhm resistor for modeling purposes.

Cable length is 600 ft RG-8/U attn = 1.3 dB/100 ft @ 50 MHz. Beam Intensity is 3.94e12/batch. The signal at the BPM is  4Vpp /0.407 =  9.83 Vpp. For 1e13 p/batch the signal at the detector would be ~ 25 Vpp. If the termination is 25 Ohms the amplitude would be 12.5 Vpp.

A 12.5 Vpp signal sent through 3/8 Heliax (atten=.848 dB/100 ft)   767 ft of cable = 5.9 Vpp and for 175 ft = 10.53 Vpp. At receiver end.

The 4 Vpp signal of Q846A was connected to a 53 MHz Bandpass filter and terminated in 50 Ohms. The signal level changed from 4 Vpp to 1.8 Vpp  a –6.93 dB of filter insertion loss.

A 12.5 Vpp signal sent through 3/8 Heliax (atten=.848 dB/100 ft)   767 ft of cable = 5.9 Vpp and for 175 ft = 10.53 Vpp. At receiver end. These signals connected directly to a 53 MHz bandpass filter would have 2.66 Vpp and 4.73 Vpp of 53 MHz component at the input of the digital down converter board.

System Requirements:

| NuMI spill (pbar operation) | 5 batches x 84 bunches = 8.2 (sec |
| --- | --- |
| NuMI spill (no pbar operation) | 6 batches x 84 bunches = 9.8 (sec |
| Maximum intensity | 4 x 1013 ppp (protons/spill) |
| Total beam power | 404 kW at maximum intensity |

Worst case design 1e13 particles per batch. plus factor of 2 for off-center beam.

Time Domain Calculations:

Ib(t) =            n=1e13 particles/80 bunches q=1.602e-19 C/particle (=2e-9sec

Ib(0)  = 3.994Amps.

The output of the BPM is dIb/dt times the length of the detector L/2c that is 5.25 in (0.13335 m).

= 222.25e-12 sec

=

plot of dI(t)/dt using the following  matlab program

function numi_i()

sigmat=2.0e-9;

t=8e-10:8e-10:100e-9;

l=0.13335; c=3e8;

n=1.0e13/(80); q=1.602e-19;

%detector length

bpm=l/(2*c);

% derivative of bunch

%

dib1=((n*q*(10e-9 - t))/(sigmat^3*sqrt(2*pi))).*exp(-0.5*((10e-9 - t)/sigmat).^2);

dib2=((n*q*(28.87e-9 - t))/(sigmat^3*sqrt(2*pi))).*exp(-0.5*((28.87e-9 - t)/sigmat).^2);

%di_bpm1=dib1+dib2;

di_bpm1=dib1*bpm+dib2*bpm;

plot(t,di_bpm1);figure;grid on;

dIb(0)/dt = 0.52 Amps peak-peak or 0.26 Amps peak.

The load resistor is 25 Ohms. The voltage generated is

dIb(0)/dt * R = 0.52 Amps peak-peak * 25 = 13 Volts peak-peak.

Cable attenuation is .848 dB/100 ft. The shortest cable length is 175 ft (1.48 dB attn) and the longest is 767 ft  (6.5 dB attn)

dIb(0)/dt * Ro // Rl * Cable attn = 13 Volts peak-peak * 0.8433= 10.96 Volts peak-peak  and 13V p-p * 0.473 = 6.15 Volts peak-peak.

A safety margin of 2 should be added for cases where beam is off-center near one of the plates, doubling the total voltage seen by the receiver.

Vin = 21.92 Vp-p and 12.3 Vp-p for 1e13 particles per batch.

The signals at the coupled output are attenuated by 11 dB. This makes the signal levels at the input of the op-amp

Vopamp = 6.17 Vp-p and 3.4 Vp-p.

Frequency Domain calculations :

In a batch the average current  determined by the number of protons, their charge and the batch period

Iav =

Iav= 1e13 * 1.602e-19/ 1.6e-6 = 1.001 Amps

I(f) =

The output of the BPM is the derivative of I(f)

=

Calculated values for dI/df case:

For 53.1 MHz and 106.2MHz

I(f) 53.1 = 261.9e-3* 0.895 = 234.48e-3 Amps;  I(f)106.2 = 524.58e-3 * 0.411 = 215.6e-3 Amps.

Consider a cable with its 25 Ω termination, the load voltage is

V(f) 53.1 = 5.86 Volts peak and V(f) 106.2 = 5.39 Volts peak