ARM TR-024

UV-B Radiometer (UVB) Handbook

November 2004

T. Stoffel

Work supported by the U.S.  Department of Energy,

Office of Science, Office of Biological and Environmental Research

Contents

1.	General Overview	1

2.	Contacts	1

3.	Deployment Locations and History	1

4.	Near-Real-Time Data Plots	1

5.	Data Description and Examples	2

6.	Data Quality	4

7.	Instrument Details	5

# General Overview

A UV-B radiometer is used to measure ultraviolet solar irradiance on a planar surface. Accurate measurements of UV are needed for erythemal (sunburning) and other biological concerns, as well as to help understand the primary source of energy input to the earth-atmosphere-ocean system. The ARM Program uses UVB radiometers for ground-based in situ UV radiation measurements collected by the ARCS/SKYRAD instrument platforms in the TWP and NSA regions. UVB instruments are not currently installed in the SGP region.

# Contacts

## Mentor

Tom Stoffel
National Renewable Energy Laboratory
1617 Cole Blvd.
Golden, CO 80401-3393
Phone:  303-384-6395
Fax:  303-384-6391
Email: tom_stoffel@nrel.gov

## Instrument Developer

Solar Light Co., Inc. 
721 Oak Lane
Philadelphia, PA 19126
Phone:  215-927-4206
Fax:  215-927-6347
E-mail:  info@solar.com
Web:  http://www.solar.com

# Deployment Locations and History

Solar Light 501A UV-Biometers are currently deployed in the Tropical Western Pacific (TWP) and the North Slope of Alaska (NSA).

# Near-Real-Time Data Plots

Barrow, Alaska Quicklooks

SHEBA Ice Camp Quicklooks

Atqasuk, Alaska Quicklooks

TWP Data Quicklooks are available from DQ HandS (Data Quality Health and Status).

# Data Description and Examples

See Near-Real-Time Data Plots, Section 4, and Data Quality Health and Status, Section 6.1.

## Data File Contents

### Primary Variables and Expected Uncertainty

A UV-B Radiometer measures the amount of ultraviolet radiation incident on a planar surface. Measurements are reported in milliwatts per square meter (mW/m2).

#### Definition of Uncertainty

Data are returned in units of MED/HR. Specifications state the instrument has a measurement range of 0 - 10 MED/HR, and a resolution better than 0.01 MED/HR. 1 MED/HR is approximately equal to 5.83x10-6 watts/cm2 (0.0583 W/m2). Using this conversion, the range is 0 to 583 mW/m2, and the resolution is approximately 0.58 mW/m2.

Data logger and sampling uncertainties may be considered separately.

The model 501A UV-Biometer is capable of providing 1-minute averages of down welling total hemispheric ultraviolet-B irradiance measurements within +/- 5% at the 95% confidence level. This estimate assumes the instrument has been calibrated within the previous year, is properly installed, and is adequately maintained.

There are a number of conditions under which UV-B measurements may be incorrect. The data user should examine the data quality flags and data quality reports (DQR) to determine whether significant malfunctions have occurred due to problems with an instrument or data logger. The most frequent sources of error are described briefly here.

Because the 501A is an active instrument, a failure in the site power or in the DC power supply will cause faulty data.

Contaminated optical surfaces can reduce or enhance the apparent UV at the detector. Reduction of the dome transmittance due to contamination will decrease the signal from a radiometer. Under some circumstances, the presence of dew, frost, snow or ice can enhance the amount of radiation reaching the detector.

The downwelling hemispheric ultraviolet irradiance will be in error if the UV radiometer detector is not horizontal. Tilting toward or away from the sun by as little as one degree can bias the instrument's signal.

Local obstructions to the sun such as trees, buildings, power lines, or towers can shade or reflect sunlight naturally available to the instrument. The locations of all measurement stations have been selected to eliminate or greatly reduce the effects of such obstructions to accurate irradiance measurements. Generally, deviations of the local horizon within an elevation angle of 5° will not result in measurable differences for a UV-Biometer.

### Secondary/Underlying Variables

This section is not applicable to this instrument.

### Diagnostic Variables

This section is not applicable to this instrument.

### Data Quality Flags

Data are clipped below 0 W/m2 and above 1.2 W/m2 (TWP), or 0.15 W/m2 (NSA).

### Dimension Variables

This section is not applicable to this instrument.

## Annotated Examples

This section is not applicable to this instrument.

## User Notes and Known Problems

This section is not applicable to this instrument.

## Frequently Asked Questions

How often should the outer dome be cleaned?

Research-quality irradiance measurements require frequent cleaning of all optical surfaces. The frequency and amounts of contamination are site dependent. ARM suggests cleaning all exposed optical surfaces every day, preferably early in the morning. Domes should also be wiped clean after precipitation events.

What is the maximum signal cable length?

From the instrument documentation: "50 ft. (15m) standard, if longer the voltage drop of the temperature controller power must be taken into account."

When should the desiccant be changed?

When the indicating paper in the desiccator plug turns pink, the desiccant needs to be recharged or replaced.

What is the maximum irradiance expected from a UVB radiometer measurement?

Typical signal levels for the TWP region peak around 200 mW/m2. For NSA, peak values are over 100 mW/m2.

Does it matter where I install a UVB radiometer?

Yes. The instrument should be installed in a location with a free horizon -- there should be no significant obstructions to the hemispheric field of view of the sky. Obstructions to the sky view can artificially lower the readings by shading the instrument. Some obstructions can artificially increase the readings by reflecting direct UV onto the sensor.

Some examples of objects to avoid when positioning an instrument include poles, roof vents, elevator shafts, trees, light-colored buildings or buildings with glass windows. Basically anything that extends significantly above the height of the instrument sensor.

Special care should be taken to stay away from objects that may shade the direct sunlight at any time during the year. A solar calculator can be used to calculate the extreme positions of the sun around June 22 and December 22. On these dates, the sun rises and sets farthest north and south, respectively, of any time of year. Placing the instrument so that no obstructions lie to the east or west between these extreme sunrise and sunset positions will help avoid unwanted shading of the sensor.

# Data Quality

## Data Quality Health and Status

The following links go to current data quality health and status results.

- DQ HandS (Data Quality Health and Status)

- NCVweb for interactive data plotting using.

The tables and graphs shown contain the techniques used by ARM's data quality analysts, instrument mentors, and site scientists to monitor and diagnose data quality.

## Data Reviews by Instrument Mentor

Data are routinely compared with total solar for qualitative similarity.

## Data Assessments by Site Scientist/Data Quality Office

All DQ Office and most Site Scientist techniques for checking have been incorporated within DQ HandS and can be viewed there.

## Value-Added Procedures and Quality Measurement Experiments

No VAPs or QMEs are anticipated for the UV-B Radiometer. However, many of the scientific needs of the ARM Program are met through the analysis and processing of existing data products into "value-added" products or VAPs. Despite extensive instrumentation deployed at the ARM CART sites, there will always be quantities of interest that are either impractical or impossible to measure directly or routinely. Physical models using ARM instrument data as inputs are implemented as VAPs and can help fill some of the unmet measurement needs of the program. Conversely, ARM produces some VAPs not in order to fill unmet measurement needs, but instead to improve the quality of existing measurements. In addition, when more than one measurement is available, ARM also produces "best estimate" VAPs. A special class of VAP called a Quality Measurement Experiment (QME) does not output geophysical parameters of scientific interest. Rather, a QME adds value to the input datastreams by providing for continuous assessment of the quality of the input data based on internal consistency checks, comparisons between independent similar measurements, or comparisons between measurement with modeled results, and so forth. For more information, see the VAPs and QMEs web page.

# Instrument Details

## Detailed Description

### List of Components

A UV-B radiometer is the lowest component-level device used by an ARM instrument platform. Data are collected by the ARCS SKYRAD data logger. Please refer to ARCS/SKYRAD for a description of data loggers used in these systems.

### System Configuration and Measurement Methods

The ARM Program uses Solar Light Co., Inc. Model 501A UV-Biometers for UVB irradiance measurements. The 501A is a Robertson-Berger type meter, using a GaAs diode under a phosphor-coated glass detector as its primary means of sensing UV radiation. A quartz dome protects the sensor from the elements. The instrument has a metal body with a white powder-coated exterior, a spirit level for accurately determining the horizontal orientation, adjustable leveling screws and an indicating desiccator which can be readily inspected.

This active device requires DC power to operate the sensor circuitry (5.5 to 25 V, 5mA max) and temperature stabilization (11 to 15 V, 1 A max). The output signals include a voltage proportional to the UV irradiance reaching the detector, and a separate voltage output proportional to the sensor temperature. The latter is used to correct the UV signal for temperature dependence. A shielded signal cable connects the 501A to a data logger. The data logger is programmable and provides 1-minute statistics of analog to digital conversions made every second (see SKYRAD information for details).

UV-B 501As are located to provide maximum solar access (i.e., the local horizon is unobstructed to the extent possible). The radiometer is carefully leveled to provide accurate downwelling UV irradiance.

The UVB radiometer is located on the top platform of the SKYRAD stand of the ARCS installations in TWP and NSA. The instrument sits on a raised mount which brings the sensor surface to the same level as the other SKYRAD sensors. This is done to minimize shading and reflections from the other instruments.

Data are collected by the SKYRAD data logger. Sampling occurs every 1 second, with data reported as 1-minute averages of 1-second data. Minimum, maximum and standard deviation values are also reported for each minute. Three instantaneous data readings are recorded in the a-level data streams for 10, 30 and 50 seconds after the minute. Calibration factors are multiplied in by the data logger; data are reported in mW/m2.

### Specifications

The field of view of the UVB radiometer is the full sky hemisphere (2-pi steradian). The instrument is operated in a level position. The output range is 0 to 583 mW/m2, and the resolution is approximately 0.58 mW/m2.

The nominal sensor temperature is 25° C. A Peltier element is used to heat or cool the sensor to the nominal temperature. Temperature stabilization is ±1° C, with a settling time less than 10 minutes for a 20° C change. The heater is capable of ranges up to 50° C above ambient; cooling is to 20° C below ambient.

## Theory of Operation

The theory of operation for the 501A UV-Biometer is the same as that of the Robertson-Berger meter. Solar light enters the detector through the quartz dome and input filter. The partially filtered light, containing the full UV spectrum, excites the phosphor. The visible light emitted by the phosphor is detected by the GaAsP diode, which produces a current proportional to the amount of UV. This current is converted to a voltage and amplified to an output between 0 and 2.5V. The detector temperature is converted to a linear output voltage in the range of ±2.5V, with 1.0V equal to 25° C, and a slope of 20 mV/°C. The temperature of the diode and phosphor are controlled to 25° C (±1°C) by a Peltier element.

## Calibration

### Theory

Instrument calibration should be performed by the manufacturer on a yearly cycle.

### Procedures

This section is not applicable to this instrument.

### History

Calibration history information for specific instruments is available from the Instrument Mentor.

## Operation and Maintenance

This section is not applicable to this instrument.

### User Manual

This section is not applicable to this instrument.

### Routine and Corrective Maintenance Documentation

Desiccators are changed when necessary. Desiccator indicator changes color from blue to pink when saturated.

Instrument dome is cleaned with distilled water and lint-free cloth each morning, and after precipitation events.

Instrument level is checked -- and corrected if necessary -- at least monthly.

### Software Documentation

This section is not applicable to this instrument.

### Additional Documentation

This section is not applicable to this instrument.

## Glossary

Erythema - an approximation of the relative biological effectiveness (ability to cause sunburn) of different wavelengths in the UVA and UVB regions of the spectrum.

Irradiance - the amount of energy (light, heat, UV) per unit time per unit area.

MED/HR - Minimum Erythema Dose per Hour. A measure of biological effectiveness of UV irradiation. One MED/HR is approximately equal to 5.83 * 10-6 watts/cm2.

Radiance - the amount of energy per unit time per unit solid angle per unit area in the direction of travel.

Ultraviolet (UV) - the region of the electromagnetic spectrum that falls outside of the visible, just beyond violet. Defined as wavelengths between 200 and 390 nanometers.

UVA - the region of the electromagnetic spectrum between 320 and 390 nanometers.

UVB - the region of the electromagnetic spectrum between 290 and 320 nanometers.

wavelength - simply, the length between consecutive peaks or troughs in any wave. Expressed in units of length, i.e., nanometers (nm). In electromagnetic radiation, wavelength defines the color of light if it is visible, or the type of radiation (infra-red, ultraviolet, etc.) if it is outside of the visible region of the electromagnetic spectrum. Mathematically, the wavelength times the frequency of the radiation measured in cycles-per-second, or Hertz) is equal to the speed of light.

Also see the ARM Glossary.

## Acronyms

ARCS:  Atmospheric Radiation and Cloud System

DC:  Direct Current

DQR:  Data Quality Report

MED/HR:  Minimum Erythema Dose per Hour

NSA:  North Slope of Alaska

SGP:  Southern Great Plains

SKYRAD:  SKY RADiation Instrument Systems. Measures down welling radiation

TWP:  Tropical Western Pacific

UV:  Ultraviolet radiation

Also see the ARM Acronyms and Abbreviations.

## Citable References

Berger, D.  1976.  The sunburning ultraviolet meter: design and performance.  Photochem. Photobiol. 24, 587-593.

Blumthaler, M., and W. Ambach.  1986.  Messungen der Temperaturkoeffizienten des Robertson-Berger Sunburn Meters und des Eppley UV-radiometers. Arch. Met. Geophys. Biocl., Ser. B, 36, 357-363.

Blumthaler, M., W. Ambach, M. Morys, and J. Slomka.  1989.  Comparison of Robertson-Berger UV Meters From Innsbruck and Belsk.  Publs. Inst. Geophys. Pol. Acad. Sc., D-32 (230), Warsaw.

McKinlay, A., and B.L. Diffey.  1987.  A reference action spectrum for ultra-violet induced erythema in human skin.  In Human Exposure to Ultraviolet Radiation: Risks and Regulations (Edited by W.F. Passchier and B.F.M. Bosnajakovic), pp83-87, Elsevier, Amsterdam.

Solar Light Co., Inc.  1991.  UV-Biometer Model 501A Version 3 User's Manual.