Laser Safety Program Safety and Environmental Protection Program SAIC-Frederick, NCI-Frederick Created June 1996 Revised October 2003

CONTENTS

1.	PURPOSE	  1

2.	SCOPE AND APPLICATION	  1

3.	REFERENCES	  1

4.	DEFINITIONS	  2

5.	RESPONSIBILITIES	 3

5.1	Managers and Supervisors	  3

5.2	Laser Safety Officer (LSO)	  3

5.3	Occupational Health Services	  4

5.4	Laser Operators	  4

6.	LASER/LASER SYSTEM HAZARD CLASSIFICATIONS	 4

6.1	Class-I Laser/Laser System	  5

6.2	Class-II Laser/Laser (Low-Power) System	  5

6.3	Class-III Laser/Laser (Medium-Power) System	 5

6.4	Class-IV Laser/Laser (High-Power) System	 5

6.5	Embedded Laser/Laser System	  6

6.6	Fiber Optics and Laser/Laser System Classifications	 6

7.	HAZARDS ASSOCIATED WITH A LASER/LASER SYSTEM	 7

8.	REQUIREMENTS	  8

8.1	Education and Training	  8

8.2	Modifications	  8

8.3	Hazardous Materials Control	  8

8.4	Engineering and Administrative Controls	 8

8.4.1	Laser-Controlled Area (LCA)	 9

8.4.2	Laser Interlocks	 10

8.4.3	Nominal Hazard Zones (NHZs)	 10

8.4.4	Temporary Laser-Controlled Area	 10

8.5	Personal Protective Equipment (PPE)	 11

8.6	Lasers in Construction	 12

8.7	Safety Signs and Labels	 12

8.8	Laser Inspections	 13

8.9	Medical Surveillance	 13

APPENDICES

APPENDIX A — Laser Registration Form	 14

APPENDIX B — Laser Safety Self-Assessment Worksheet	 16

APPENDIX C — Laser Safety Signs	 17

APPENDIX D — Laser Inspection Worksheet	 18

- PURPOSE

This section establishes and implements requirements and assigns responsibilities for NCI-Frederick laser safety in accordance with the requirements of American National Standards Institute (ANSI) Z136.1(2000), American National Standard for the Safe Use of Lasers;

- SCOPE AND APPLICATION

This section applies to lasers/laser systems used at NCI-Frederick whether purchased, borrowed, fabricated, or brought in for use by others.  It is intended for use by NCI-Frederick employees and outside personnel who work with or on lasers/laser systems.  This section addresses the following laser-safety elements:

M	Responsibilities

M	Associated nonbeam hazards

M	Education and training

M	Laser/laser system classifications

M	Engineering and administrative controls

M	Personal protective equipment (PPE)

M	Inventory and inspection

M	Medical surveillance.

- REFERENCES

  - American National Standards Institute:

M	Z136.1(2000), American National Standard for the Safe Use of Lasers.

  - Code of Federal Regulations, Title 21:

M	Part 1040.10, Laser Products.

M	Part 1040.11, Specific Purpose Laser Products.

  - Code of Federal Regulations, Title 29:

M	Part 1910.147, The Control of Hazardous Energy.

M	Part 1910.133, Eye and Face Protection.

M	Part 1926.54, Non-ionizing Radiation.

  - Occupational Safety and Health Administration PUB 8-1.7, Guidelines for Laser Safety and Hazard Assessment.

  - Laser Institute of America, Guide for the Selection of Laser Eye Protection.

- DEFINITIONS

Accessible Emissions Limit (AEL).  The maximum accessible emission level permitted within a particular class.

Beam.  A collection of rays that may be parallel, divergent, or convergent.

Embedded Laser.  A laser housed in a system with a classification less than the laser classification, due to the engineering features limiting accessible emission.

Fail-Safe Interlock.  An interlock where the failure of a single mechanical or electrical component of the interlock will cause the system to go into, or remain in, a safe mode.

High Power.  For the purposes of this section, refers to laser emissions greater than 0.5 W.

Laser.  A device that produces an intense, coherent, directional beam of light by stimulating electronic or molecular transitions to lower energy levels.

Laser Safety Officer (LSO).  An individual, designated by the Manager, Safety and Environmental Protection Program, who has the background and training to evaluate and control laser hazards.

Laser System.  An assembly of electrical, mechanical, and optical components that includes a laser.

Maximum Permissible Exposure (MPE).  The level of laser radiation a person may be exposed to without hazardous effects or adverse biological changes in the eye or skin.

Nominal Hazard Zone (NHZ).  A physical space where the level of direct, reflected, or scattered radiation during normal operation exceeds the applicable MPE.  Exposure levels beyond the boundary of the NHZ are below the appropriate MPE level.

Reflection, Diffuse. A change in the spatial distribution of a beam of radiation when it is reflected in many directions by a surface or by a medium.

Reflection, Specular. A mirror-like deviation of radiation following incidence on a surface.

- RESPONSIBILITIES

  - Managers and Supervisors have responsibility for:

      - Assuming overall responsibility for the safe operation of all lasers operated by their personnel.

      - Ensuring that laser operators are trained in the safe operation of the laser for which they are registered.

      - Ensuring all laser operators have satisfied the medical requirements of

Subsection .

      - Completing a Laser Registration Form (Ref. Appendix A) for each laser used in their area and maintaining a copy.

      - Upon a change in status of any laser under his/her responsibility, completing a Laser Registration Form and obtaining the required approval signature from the appropriate manager.

      - Requesting a change of registration from the LSO if the laser is removed from its primary facility/area for more than six months.

      - Performing periodic self-assessments on lasers/laser operations and documenting the self-assessments on Appendix B, Laser Safety Self-Assessment Worksheet.

      - Ensuring Class-IIIb and -IV laser operations have written operating and alignment procedures, and that these procedures are approved by the Laser Safety Officer.

5.2	The Laser Safety Officer (LSO) has responsibility for:

[1]	Annually reviewing and, when necessary, revising this Program to ensure 		compliance with all applicable federal and state regulations.

[2]	Serving as a liaison with NCI-Frederick personnel in matters regarding laser 		safety.

[3]	Assisting line management with interpretation of laser safety requirements.

[4]	Maintaining copies of the Laser Registration Forms (ref.Appendix A) for all 		lasers/laser systems in use at NCI-Frederick.

[5]	Recommending measures to line management for the control of nonbeam 		hazards.

[6]	Conducting annual inspections of each registered laser at NCI-Frederick and 		conducting additional laser inspections at the request of laser operators.

[7]	Maintaining current inspection records for each registered laser at NCI-		Frederick.

[8]	Working with line management to develop, schedule, present, and document 		laser safety training for all employees and outside personnel who work with or 		on a laser/laser system.  (The training must meet the requirements of 		ANSI Z136.1.)

5.3	Occupational Health Services has responsibility for maintaining laser eye exam 	information as part of the medical files for all registered NCI-Frederick laser operators.

5.4	Laser Operators have responsibility for:

[1]	Operating all laser/laser systems in accordance with the manufacturer's 		requirements and/or an approved operating and alignment procedure.

[2]	Verifying laser-controlled-area (LCA) interlocks are functioning properly before 		conducting any laser operations.

[3]	Working with the LSO to identify risks for the documentation of a nominal 		hazard zone (NHZ) for each Class-IIIb and -IV laser.

[4]	Upon receipt of a laser into the work place, contacting the LSO for an 		inspection of the laser for conformance with the Code of Federal Regulations 		(CFR), Title 21, Parts 1040.10, Laser Products, and 1040.11, Specific Purpose 		Laser Products, and for registration into the computerized laser database.

[5]	Wearing dedicated eye protection for CO2, argon, and krypton lasers and such free 		of scratches.

- LASER/LASER SYSTEM HAZARD CLASSIFICATIONS

The primary method for determining controls for a laser/laser system is based on the laser classification.  Classification is based on the laser's ability to cause biological damage.  This section covers laser, laser system, and optical-fiber-communications-system (OFCS) classifications.

  - Class-I Laser/Laser System

A Class-I laser/laser system is considered to be incapable of producing damaging laser emission levels and is therefore exempt from most control measures or other forms of surveillance.  A Class-I system with an embedded laser is covered separately in Subsection .

  - Class-II Laser/Laser (Low-Power) System

A Class-II laser/laser system emits light in the visible portion of the spectrum (0.4-0.7 μm).  Eye protection is normally afforded by the involuntary human aversion response to bright radiant sources.  However, this laser/laser system classification may present some risk potential if the aversion response is overcome and the beam is viewed directly for extended periods of time.

The Class-IIa laser/laser system is a special classification where the laser/laser system emits light in the visible spectrum, the system is intended for a specific use, and the output beam is not intended to be viewed directly.  This laser/laser system must not exceed the Class-I Accessible Emmission Limit for an exposure duration of less than 1000 seconds (16.7 min).  This classification includes bar-code scanners commonly used for inventory control.

6.3	Class-III Laser/Laser (Medium-Power) System

A Class-III laser/laser system may be hazardous under direct-beam and specular-reflection viewing conditions, but the diffuse reflection is usually not a hazard.  The Class-III classification is divided into two subclasses, Class IIIa and IIIb, each with different control requirements.  A Class-III laser/laser system should not present a fire hazard.

6.4	Class-IV Laser/Laser (High-Power) System

A Class-IV laser/laser system is a hazard to eye and skin from the direct beam and diffuse beam reflections.  A Class-IV laser/laser system can present a fire hazard.  Any laser/laser system or modified laser/laser system of unknown power must be operated as a Class-IV laser/laser system until reclassified in accordance with ANSI Z136.1.

6.5	Embedded Laser/Laser System

An embedded laser/laser system is normally a Class-I or -IIa laser/laser system until opened for servicing or maintenance.  At which time, it is then reclassified, depending on the type of laser inside.  Laser printers, which are an example of a Class-I embedded laser, are exempt from the requirements of this section.

Any laser/laser system can be converted to a Class-I enclosed laser/laser system by incorporating appropriate engineering controls in the laser/laser system design and obtaining approval from the LSO.  These engineering controls enclose the laser, preventing personnel exposure to laser emissions, while permitting unrestricted access into the area, and include the following;

M	Protective housing

M	Interlocks (interlocks may be defeatable for maintenance purposes)

M	Fail-safe design

M	Attenuated viewing windows (portals)

M	Safety signs and labels (ref. Appendix C, Laser Safety Signs and Labels).

6.6	Fiber Optics and Laser/Laser System Classifications

A fiber-optic laser/laser system is a laser/laser system that employs optical cable to transmit the laser light.  This laser/laser system is considered an enclosed system with the optical cable forming part of the enclosure.  Disconnection of a connector may result in accessible laser emissions above the applicable Class-I AEL.

Fiber-optic laser systems are divided into four classifications as listed on the following page (OFCS Service-Group Ratings).

OFCS Service-Group Ratings

Rating 1

O Any OFCS containing a light-emitting diode (LED) or laser/laser system that cannot emit accessible levels of optical emission in excess of the service-group rating 1 AEL for an emission duration of 100 seconds. It is exempt from all control measures and any form of surveillance.

Rating 2

O Any OFCS containing an LED or laser/laser system emitting in the visible light region (O.4-0.7 μm) that emits a total accessible radiant power exceeding the appropriate service-group rating 1 AEL for the maximum possible emission duration inherent in the design of the system, but not exceeding the service-group rating 1 AEL for an exposure duration of 0.25 seconds or less.

Rating 3a

O Near-infrared continuous wave or repetitively pulsed OFCSs that emit accessible radiant power in excess of the service-group rating 1 AEL for an emission duration of 100 seconds, but cannot emit an average radiant power in excess of the AEL for service-group rating 3a. O Infrared continuous wave or repetitively pulsed OFCSs that emit accessible radiant power in excess of the service-group rating 1 AEL for an emission duration of 10 seconds, but that cannot emit an average radiant power in excess of the AEL for service-group rating 3a. O Visible continuous wave or repetitively pulsed OFCSs that emit accessible radiant power in excess of the service-group rating 2 AEL for an emission duration greater than 0.25 seconds, but cannot emit an average radiant power in excess of the AEL for service-group rating 3a.

Rating 3b

O Near-infrared continuous wave or repetitively pulsed OFCSs that emit accessible radiant power in excess of the service-group rating 3a AEL for an emission duration of 100 seconds, but cannot emit an average radiant power in excess of 0.5 W. O Infrared continuous wave or repetitively pulsed OFCSs, that emit accessible radiant power in excess of the service-group rating 3a AEL for an emission duration of 10 seconds, but cannot emit an average radiant power in excess of 0.5 W. O Visible continuous wave or repetitively pulsed OFCSs that emit accessible radiant power in excess of the service-group rating 3a AEL for an emission duration of 100 seconds, but cannot emit an average radiant power in excess of 0.5 W.

- HAZARDS ASSOCIATED WITH A LASER/LASER SYSTEM

Improperly used laser beams can ignite flammable materials or volatilize toxic or hazardous fumes, gases, debris, and radiation.  In maintenance or normal use, the equipment and optical apparatus required to produce and control the laser energy may introduce additional risks associated with high voltage, high pressure, cryogenics, noise, radiation, and toxic fluids.  In addition to the beam hazards, flashlamps can produce nonionizing radiation hazards.  Microwave, ultraviolet, and radio frequency emmissions may be generated by a laser/laser system or support equipment.

- REQUIREMENTS

  - Education and Training

    - All personnel who use laser/laser systems in NCI-Frederick facilities/areas must be trained on the appropriate operating characteristics and safety hazards.  Job-specific training must be completed and documented.

  - Modifications

Modifications to commercial laser/laser systems must be evaluated and approved by  the LSO.

  - Hazardous Materials Control

An operator shall bring into the Nominal Hazard Zone only those hazardous materials that are needed for a particular laser operation.  It is the operator's responsibility to ensure these materials are properly used, stored, and controlled.  The operator shall not allow laser beams and reflections to impinge on combustible materials, explosives, flammable liquids or gases, or substances that decompose into highly-toxic products under elevated temperature, without providing adequate controls.  The operator shall conduct or sponsor LSO-approved tests that establish the effects of beam interactions with hazardous materials.  Test results must then be used to determine safe parameters for laser operation.

8.4	Engineering and Administrative Controls

When operating lasers/laser systems, hazards must be evaluated prior to operation.  Engineering and administrative controls, as determined by the LSO, are required with the various  laser classifications to provide a safe work environment consistent with the requirements of ANSI Z136.1.  These controls must take into account the hazard potential of the laser/laser system, the environment in which the laser/laser system is used, and the personnel who may use or be exposed to the laser/laser system.  Use of embedded Class-III and -IV laser/laser systems will require different controls during normal operation than during alignment and maintenance.  The existence of other hazards in the laser environment may require additional controls.

All Class-II and above laser/laser systems must have a written, LSO-approved alignment procedure and documented Nominal Hazard Zone.  Class-IIIb and -IV laser/laser systems must have a standard operating procedure.  Since each laser/laser system is unique, the LSO shall work with the user to develop a standard operating procedure that contains the necessary administrative controls and safety requirements.  The procedure must include, but not be limited to, the following:

M	Necessary requirements for laser setup, adjustment, and operation.  Such activities must be performed by authorized personnel only.

M	A caution statement indicating that direct viewing of the laser beam (and in some cases, the reflected beam) is extremely dangerous.

M	A requirement that the laser must be flashed once after the switch is turned off to remove most of the residual charge from the capacitors.  (Lasers in hot cells are exempt from this requirement.)

If the levels of accessible fiber-optical cable emissions exceed Class-I AEL, appropriate engineering and administrative controls, consistent with the hazards classification, must be applied.

When hazardous levels of laser emissions (greater than or equal to the MPE) may be present at the end of optical fibers, a label with the appropriate signal word "CAUTION" (OFCS Service Group Rating 2 and OFCS Service Group Rating 3a) or "DANGER" (OFCS Service Group Rating 3b) (ref. Appendix C) must be attached and the optical fiber capped off when not in use.

8.4.1	Laser-Controlled Area (LCA)

An LCA is a location in which a Class-IIIb or -IV laser-associated experiment or operation is conducted.  The LCA may be a laboratory or a portion of a laboratory.  The LCA is used as an engineering control mechanism for limiting access to a laser operation and must include active and administrative controls to limit access and exposure to nonessential personnel.  An LCA must include the following:

M	A danger sign (ref. Appendix C)

M	Red warning light

M	Emergency shutdown feature (such as a panic button)

M	Door interlocks (with optional door interlock bypass).

8.4.2	Laser Interlocks

All Class-IIIb and -IV lasers must include the following engineering controls (as a minimum):

M	External interlock connector

M	Case interlocks

M	Master keyswitch.

8.4.3	Nominal Hazard Zones (NHZs)

M	During Class-IIIb or -IV laser operations, beam hazards must be limited to NHZs.

M	NHZs must encompass all areas within an LCA where the MPE may be exceeded.

M	NHZs must be determined and documented by the LSO.

8.4.4	Temporary Laser-Controlled Area

A temporary laser-controlled area must be established for short-duration (less than six months) laser operations and where implementing an LCA is not practical.  These areas are normally controlled through LSO-approved administrative controls.  A notice sign (ref. Appendix C) will be used to indicate the entrance into a temporary laser-controlled area.  Where maintenance or repair to a laser is taking place, a danger sign is needed for any other temporary LCAs.

8.5	Personal Protective Equipment (PPE)

When other control measures do not provide adequate means to prevent access to direct or reflected beams at levels above the MPE,  PPE must be used as specified by the LSO.

The requirements for PPE will be identified in a standard operating procedure.  Impervious gloves, face shields, and disposable coveralls may be required for handling certain types of dyes and solvents.  Personal protective equipment may be required when operating ultraviolet or Class-IV lasers.  Personal protective equipment used with Class-IV laser/laser systems must be fire resistant.

Specific to Laser Eyewear.  Where possible, laser eyewear must be selected that meets or exceeds ANSI Z87.1, American National Standard for Occupational and Educational Eye and Face Protection.  The LSO shall consider the following criteria  and work with the laser operators to determine the appropriate eyewear.

M	Wavelength

M	MPE and optical density (OD)

M	Visible light transmission

M	Lens material

M	Mechanical construction

M	Task restrictions.

Laser eyewear must be labeled and include:

M	the wavelength(s) for which the eyewear is designed (Dλ);

M	the attenuation at each wavelength (OD); and

M	the last LSO inspection date.

Each set of laser eyewear must be inspected by the LSO prior to use and annually thereafter.

8.6	Lasers in Construction

The following requirements pertain only to lasers/laser operations involved in construction.  Where conflicts occur between these requirements and ANSI Z136.1, the most conservative requirement applies.

      - Special Maximum Permissible Exposure limits include:

M	Direct staring:  1 μW/cm2

M	Incidental:  1 m/W/cm2

M	Diffuse reflected:  2.5 W/cm2.

      - Protective eyewear must be worn by personnel operating a Class-IIIb or -IV laser.

      - Beam detection must be only by mechanical or electronic means.

      - The beam must be positioned below eye level whenever possible.

      - A temporary laser-controlled area must be posted, and where possible, constructed using physical barriers.

8.7	Safety Signs and Labels

All lasers, laser systems, and entrances to laser-controlled areas and Nominal Hazard Zones must be posted with signs or labels (as applicable) appropriate to the class of laser present.  For example:

M	A CAUTION sign or label must be used for Class-II and -IIIA lasers that do not exceed the appropriate Maximum Permissible Exposure for irradiance.

M	A DANGER sign or label must be used for Class-IIIa lasers that exceeds the Maximum Permissible Exposure  for irradiance, and Class-IIIb and -IV lasers.

M	A NOTICE sign must be posted outside a temporary laser-controlled area  when maintenance is being performed upon Class-IIIb or -IV lasers.

8.8	Laser Inspections

[1]	The LSO shall inspect all laser/laser systems upon initial receipt prior to 		operation and annually thereafter.  Inspections must be documented using 		Appendix D, Laser Inspection Worksheet.

[2]	A laser inspection must include, but is not limited to:

M	Labeling (i.e., warning, aperture, access panel, and eyewear labels)

M	Protective housing

M	Safety interlocks on housing and access panels

M	Remote interlock (Class IIIb and IV only)

M	Key control (Class IIIb and IV only)

M	Emissions indicator

M	Beam attenuator

M	Operating controls

M	Manual reset (Class IV only)

M	Associated hazards.

8.9	Medical Surveillance

Prior to the operation of any Class-IIIb of -IV laser/laser system, the laser operator must obtain an initial dilated pupil eye exam for baseline.  Laser eye exams should include history, visual acuity, retinal exam, and corneal observations (scars, burns) and are scheduled by Occupational Health Services.

If an operator suspects an eye strike OHS must be informed, and a new dilated pupil eye exam must be completed to determine whether any eye damage has occurred.

Upon termination of employment, a final laser eye exam must be completed to determine if any undetected eye damage has occurred during employment.  The laser eye exam request must be turned in to Occupational Health Services at least two weeks prior to the termination date to allow for scheduling.

NOTE There is no requirement for periodic laser eye exams.

Appendix A

Page 1 of 2

(Form is shown smaller than actual size.)

| LASER REGISTRATION FORM |
| --- |
| Check one: | G Add to database G Change/update database |  |
| Laser type: |  |
| Manufacturer: |  |
| Property No.: |  |
| Model No.: |  |
| Serial No.: |  |
| Wavelength range: (Circle appropriate unit) |  | ηmeters / μmeters |
| Maximum peak power: (Circle appropriate unit) |  | W(s) / mW(s) / μW(s) |
| Maximum peak energy: (Circle appropriate unit) |  | Joule(s) / mJoule(s) / μJoule(s) |
| Minimum pulse duration: (Circle appropriate unit) |  | μsecond(s) / ηsecond(s) / ρsecond(s) |
| Minimum pulse repetition frequency: (Circle appropriate unit) |  | Pulses: per second / per minute |
| Primary location: |  |
| Laser class (check one): | G Class-I Embedded Laser G Class II G Class IIIa G Class IIIb G Class IV |
| Safety inspection | Initial date: |  | Last date: |  | Due date: |  |
| Position: |  | Signature: |  | Date: |  |
| Additional information. Include name(s) of laser operator(s) with type of training and date(s) received. |  |

Appendix A

Page 2 of 2

Instructions for Completing Laser Registration Form

Check one.  Identify whether this is an addition or update to the database.

Laser type.  Identify the lasing medium used for the laser in question (i.e., HeNe, Nd:YAG).

Manufacturer.  Identify the laser manufacturer.

Property No.  Identify the laser-head property number.

Model No.  Identify the laser-head model number.

Serial No.  Identify the laser-head serial number.

Wavelength range.  If the laser has a frequency mixing device present in the same case, identify the operating wavelengths leaving the laser case.  Circle the appropriate unit at right.  [This can vary from a single wavelength for lasers (such as a HeNe) to multiple wavelengths for ion and dye lasers.]  Enter N/A if not applicable.

Maximum peak power.  For CW lasers, identify the maximum peak power for repetition rates.  Enter N/A for non-CW lasers.  Circle appropriate unit at right.

Maximum peak energy.  Applicable to pulsed lasers only —   Circle appropriate unit at right.  Enter N/A for CW-only lasers.

Minimum Pulse duration.  Applicable to pulsed lasers only — Circle appropriate unit at right.  Enter N/A for CW-only lasers.

Minimum pulse repetition frequency (PRF).  Applicable to pulsed lasers only — Identify all repetition rates at which the laser will operate.  Enter N/A for CW-only lasers.  Circle appropriate unit at right.

Primary location.  Identify the primary location of the laser.  Be specific to the building number where it is located.

NOTE If this location is changed for more than six months, an update form must be submitted to include this information.

Laser class.  Check appropriate class.  If the system contains an embedded laser, enter the appropriate laser classification for the embedded laser in the space provided.

Safety inspection.  Inspect the laser for conformance to 21 CFR-1040.  LSO — Sign this section upon inspection completion and enter the due date for the next inspection (12 months from last laser inspection date).

Position/Signature/Date.  Responsible Manager — Sign and date to indicate that the information on the form is valid.

Additional information:  Include information regarding specific dyes, gases, etc., used with the laser.  Include descriptions; serial, property, and model numbers of support equipment necessary to make the laser lase (such as power supplies, flow controls).

Appendix B

Page 1 of 1

(Form is shown smaller than actual size.)

| LASER SAFETY SELF-ASSESSMENT WORKSHEET |
| --- |
| Area: |  |
| Supervisor (Name and Signature): |  | Date: |  |
| Item | Supervisor Initials |
| Training |
| Laser operators are trained in general laser safety. |  |
| Laser operators are trained to job-specific functions. |  |
| Required Inspections |
| Laser-control-area interlock inspections are current. |  |
| Lasers in use are registered in the NCI-FCRDC database. |  |
| Laser inspections for lasers in use are current. |  |
| Engineering Controls |
| Laser-control-area interlocks are operational. |  |
| Laser-control-area interlock bypass systems are being used appropriately. |  |
| Controls are appropriate to hazards. |  |
| Class-I embedded laser system requirements met. |  |
| Personnel Protective Equipment (PPE) |
| Any LSO-required eye protection available and properly labeled. |  |
| The appropriate PPE available for nonbeam hazards. |  |
| Administrative Controls |
| Current alignment procedure available. |  |
| The alignment procedure covers all nonbeam hazards; appropriate controls in place. |  |
| Nominal hazard zones are identified for Class-IIIb and -IV laser operations. |  |
| Alignment procedures are available for Class-IIIb and -IV laser operations. |  |
| Operating procedures are available for Class-IIIb and -IV laser operations. |  |

Appendix C

Page 1 of 1

Laser Safety Signs and Labels

Appendix D

Page 1 of 1

(Form is shown smaller than actual size.)

| LASER INSPECTION WORKSHEET |
| --- |
| Laser type: |  |
| Laser serial No.: |  |
| Laser Safety Officer (Name and Signature): |  | Date: |  |
| Item | LSO's Initials |
| Labeling |
| Warning Label — Laser(s) systems greater than Class I are labeled with a warning label. |  |
| Aperture Label — Laser(s) systems greater than Class-IIe systems have aperture warning labels for each aperture. |  |
| Access Panel Label — Access panels that are not interlocked or have defeatable interlocks are properly labeled. |  |
| Protective Housing |
| The housing is at ground potential with the laser energized. |  |
| The housing has no light leaks that would expose the operator to laser or flashlamp emission levels greater than Class-I AEL. |  |
| If the laser has no housing (as in the case of some experimental systems) the proper controls and procedures are in place to protect personnel. |  |
| Safety Interlocks (i.e., on housing and access panels) |
| Have been tested for operation. |  |
| Fail-safe design (for Class-IIIb and -IV lasers or Class-I systems with an embedded Class-IIIb or -IV laser) has no more than one snap-action spring switch. |  |
| Remote Interlock (Classes IIIb and IV only) — It has been tested for operation with LCA interlock system. |  |
| Key Control (Classes IIIb and IV only) |
| The key cannot be removed when the laser is on. |  |
| Key control is functional. |  |
| Emission Indicator — The indicator is functional. |  |
| Beam Attenuator — Either a mechanical or electrical device that reduces emissions to below Class-I AEL is installed. |  |
| Operating Controls — They are positioned such that the operator is not exposed to the beam during laser operation. |  |
| Manual Reset — Class IV only — A manual reset is required to reenergize the laser if power is lost for greater than 5 seconds. |  |

| LASER SAFETY SELF-ASSESSMENT WORKSHEET |
| --- |
| Area: |  |
| Supervisor (Name and Signature): |  | Date: |  |
| Item | Supervisor Initials |
| Training |
| Laser operators are trained in general laser safety. |  |
| Laser operators are trained to job-specific functions. |  |
| Required Inspections |
| Laser-control-area interlock inspections are current. |  |
| Lasers in use are registered in the NCI-Frederick database. |  |
| Laser inspections for lasers in use are current. |  |
| Engineering Controls |
| Laser-control-area interlocks are operational. |  |
| Laser-control-area interlock bypass systems are being used appropriately. |  |
| Controls are appropriate to hazards. |  |
| Class-I embedded laser system requirements met. |  |
| Personnel Protective Equipment (PPE) |
| Any LSO-required eye protection available and properly labeled. |  |
| The appropriate PPE available for nonbeam hazards. |  |
| Administrative Controls |
| Current alignment procedure available. |  |
| The alignment procedure covers all nonbeam hazards; appropriate controls in place. |  |
| Nominal hazard zones are identified for Class-IIIb and -IV laser operations. |  |
| Alignment procedures are available for Class-IIIb and -IV laser operations. |  |
| Operating procedures are available for Class-IIIb and -IV laser operations. |  |

| LASER REGISTRATION FORM |
| --- |
| Check one: | G Add to database G Change/update database |  |
| Laser type: |  |
| Manufacturer: |  |
| Property No.: |  |
| Model No.: |  |
| Serial No.: |  |
| Wavelength range: (Circle appropriate unit) |  | ηmeters / μmeters |
| Maximum peak power: (Circle appropriate unit) |  | W(s) / mW(s) / μW(s) |
| Maximum peak energy: (Circle appropriate unit) |  | Joule(s) / mJoule(s) / μJoule(s) |
| Minimum pulse duration: (Circle appropriate unit) |  | μsecond(s) / ηsecond(s) / ρsecond(s) |
| Minimum pulse repetition frequency: (Circle appropriate unit) |  | Pulses: per second / per minute |
| Primary location: |  |
| Laser class (check one): | G Class-I Embedded Laser G Class II G Class IIIa G Class IIIb G Class IV |
| Safety inspection | Initial date: |  | Last date: |  | Due date: |  |
| Position: |  | Signature: |  | Date: |  |
| Additional information. Include name(s) of laser operator(s) with type of training and date(s) received. |  |
| LASER INSPECTION WORKSHEET |
| Laser type: |  |
| Laser serial No.: |  |
| Laser Safety Officer (Name and Signature): |  | Date: |  |
| Item | LSO's Initials |
| Labeling |
| Warning Label — Laser(s) systems greater than Class I are labeled with a warning label. |  |
| Aperture Label — Laser(s) systems greater than Class-IIe systems have aperture warning labels for each aperture. |  |
| Access Panel Label — Access panels that are not interlocked or have defeatable interlocks are properly labeled. |  |
| Protective Housing |
| The housing is at ground potential with the laser energized. |  |
| The housing has no light leaks that would expose the operator to laser or flashlamp emission levels greater than Class-I AEL. |  |
| If the laser has no housing (as in the case of some experimental systems) the proper controls and procedures are in place to protect personnel. |  |
| Safety Interlocks (i.e., on housing and access panels) |
| Have been tested for operation. |  |
| Fail-safe design (for Class-IIIb and -IV lasers or Class-I systems with an embedded Class-IIIb or -IV laser) has no more than one snap-action spring switch. |  |
| Remote Interlock (Classes IIIb and IV only) — It has been tested for operation with LCA interlock system. |  |
| Key Control (Classes IIIb and IV only) |
| The key cannot be removed when the laser is on. |  |
| Key control is functional. |  |
| Emission Indicator — The indicator is functional. |  |
| Beam Attenuator — Either a mechanical or electrical device that reduces emissions to below Class-I AEL is installed. |  |
| Operating Controls — They are positioned such that the operator is not exposed to the beam during laser operation. |  |
| Manual Reset — Class IV only — A manual reset is required to reenergize the laser if power is lost for greater than 5 seconds. |  |

Instructions for Completing Laser Registration Form

Check one.  Identify whether this is an addition or update to the database.

Laser type.  Identify the lasing medium used for the laser in question (i.e., HeNe, Dye).

Manufacturer.  Identify the laser manufacturer.

Property No.  Identify the laser-head property number.

Model No.  Identify the laser-head model number.

Serial No.  Identify the laser-head serial number.

Wavelength range.  If the laser has a frequency mixing device present in the same case, identify the operating wavelengths leaving the laser case.  Circle the appropriate unit at right.  [This can vary from a single wavelength for lasers (such as a HeNe) to multiple wavelengths for ion and dye lasers.]  Enter N/A if not applicable.

Maximum peak power.  For CW lasers, identify the maximum peak power for repetition rates.  Enter N/A for non-CW lasers.  Circle appropriate unit at right.

Maximum peak energy.  Applicable to pulsed lasers only —   Circle appropriate unit at right.  Enter N/A for CW-only lasers.

Minimum Pulse duration.  Applicable to pulsed lasers only — Circle appropriate unit at right.  Enter N/A for CW-only lasers.

Minimum pulse repetition frequency (PRF).  Applicable to pulsed lasers only — Identify all repetition rates at which the laser will operate.  Enter N/A for CW-only lasers.  Circle appropriate unit at right.

Primary location.  Identify the primary location of the laser.  Be specific to the building number where it is located.

NOTE If this location is changed for more than six months, an update form must be submitted to include this information.

Laser class.  Check appropriate class.  If the system contains an embedded laser, enter the appropriate laser classification for the embedded laser in the space provided.

Safety inspection.  Inspect the laser for conformance to 21 CFR-1040.  LSO — Sign this section upon inspection completion and enter the due date for the next inspection (12 months from last laser inspection date).

Position/Signature/Date.  Responsible Manager — Sign and date to indicate that the information on the form is valid.

Additional information:  Include information regarding specific dyes, gases, etc., used with the laser.  Include descriptions; serial, property, and model numbers of support equipment necessary to make the laser lase (such as power supplies, flow controls).