November 14, 2006

Operational Acceptance Test (OAT) Plan

for

Automated Surface Observing System (ASOS) Software Version (V) 2.79C

Introduction and Background

ASOS Operational Trouble Report (OTR) # 1078, dated 09/18/06, identifies a condition where an erroneous Ice Free Wind Sensor (IFWS) maintenance indicator ($) is appended to the Aviation Routine Weather Report (METAR) or Aviation Selected Special Weather Report (SPECI) reports.  This can occur when the IFWS correctly reports a sensor path error for path 5 (one of the 6 paths in the IFWS).  This path error is incorrectly interpreted by the Acquisition Control Unit (ACU) software and an erroneous $ is appended to the next METAR or SPECI.  Although the IFWS is operating normally (it only needs 4 of the 6 paths for a VALID wind report) and no data quality error exists, this erroneous $ may mask other potential ASOS problems.  In addition, the ASOS Operations and Monitoring Center (AOMC) has experienced an increased workload in responding to the erroneous $ indicators and, in some cases, IFWS which were functioning properly were replaced unnecessarily.

During September 2006, 220 sites (68% of the sites with an IFWS and V2.79B installed) flagged this error at least once – for a total of 495 errors.  Twenty sites flagged the error five or more times and four of those flagged the error 10 or more times.  The problem is present in the V2.79, V2.79A, and V2.79B ACU software versions.

The ASOS Configuration Control Board (ACCB) has judged this problem warrants immediate attention and revised software (V2.79C) is now available for evaluation.  With the exception of the fix for the IFWS path error maintenance flag problem, V2.79C is identical to V2.79B.

Previous Testing and OAT Schedule

A V2.79C System Test (ST) began on October 13, 2006, using test ASOS systems at Weather Service Headquarters and at the Sterling Research and Development Center in Sterling, Virginia, and was completed successfully on October 25, 2006.  The fix for the erroneous $ problem was confirmed during the ST.  The V2.79C OAT will begin during the week of November 13, 2006, and will last approximately two weeks.  If no problems are noted during the OAT, V2.79C will be released immediately for all operational ASOSs with an IFWS currently installed.

OAT Objectives

During the OAT, the following must be verified with V2.79C installed:

Stable ASOS operation - The availability of ASOS observations/products and the frequency of warmstarts must be consistent with or improved over performance with previous software versions.   There must be no spontaneous coldstarts.

Operator interfaces - ASOS must respond properly to operator commands.

Production, transmission, and archiving of observations/products - Observations must be representative of conditions (within the limitations of the ASOS sensors), correctly formatted, transmitted successfully, and archived by ASOS.

Interfacing with both NWS and non-NWS systems - ASOS must correctly transmit information for display on the FAA ASOS Controller Equipment (ACE) - verification by FAA tower staff.  ASOS must correctly incorporate Runway Visual Range (RVR) information from the FAA New Generation RVR (NGRVR) system into the ASOS METAR/SPECI reports.  ASOS must correctly incorporate thunderstorm information from the FAA Automated Lightning Detection and Reporting System (ALDARS) system into the ASOS METAR/SPECI reports.

IFWS maintenance indicator $ – There must be no erroneous $'s for the IFWS.  The AOMC will monitor the site performance and report any cases of erroneous $'s.

Methodology

Since all field ASOSs are operational, testing will be limited to performance monitoring during routine operations.  No on-site test activities other than monitoring observations/products and reporting/documenting any ASOS problems are required.

ASOS Electronics Technicians (ET) will install the software at the test sites using the procedures in Modification Note 80A.  The V2.79C software is available on the OPS12 ftp site and Modification Note 80A is posted on the ASOS ET website.

Since all OAT sites (except Anchorage, AK) already have V2.79A or V2.79B installed, and there are no changes to the Operator Interface Device (OID) displays and observer procedures with V2.79C, no training of FAA or FAA contract observers is required before installation of the software at FAA-staffed sites.  Although training is not required for the NWS observers at Anchorage, training will be required for the FAA tower staff.  When V2.79C is installed at Anchorage (PANC), it must also be installed at Lake Hood (PALH) to ensure consistency in the Anchorage tower OIDs.

For sites with FAA, FAA contract, or NWS contract staff, the NWS site Focal Point will monitor ASOS performance and review observations, reporting any ASOS problems to the OAT Coordinator.  Responsibility for the evaluation may be assigned to the on-site FAA or contract staff (at the discretion of the NWS Region and with the agreement of the FAA) but the NWS site Focal Point remains responsible for reporting problems to the OAT Coordinator:

Bryan Moore – Test and Evaluation Branch (OPS24)

301-713-0326 x176

Bryan.Moore@noaa.gov

For NWS-staffed sites, the test site Focal Point (or on-site staff, under the direction of the test site Focal Point) will monitor ASOS performance and review observations, reporting any ASOS problems to the OAT Coordinator.

For unstaffed sites, the NWS site Focal point will monitor ASOS performance and review observations remotely, reporting any problems to the OAT Coordinator.

The AOMC will report, to the OAT Coordinator, any cases when test site ASOS observations/products are not received at the NWS Telecommunication Gateway as expected.

At any time during the OAT, if the test site Focal Point judges ASOS performance to be unacceptable, he/she is authorized to have the ET re-install the previous software version to ensure continued satisfactory support of NWS and FAA operations.

The OAT Coordinator will conduct weekly conference calls with participation by Regional Focal Points and a Test Review Group made up of NWS and FAA national and regional representatives (Site Focal Points may participate at the invitation of the Regional Focal Points).  Times and specific details will be provided (by e-mail) to the conference call participants.

At the conclusion of the evaluation, the OAT Coordinator will provide test results and an implementation recommendation to the ASOS Configuration Control Board Chair.

OAT Sites and Site Focal Points

The OAT sites were selected from those experiencing the IFWS erroneous path error $ problem.  The OAT sites and Site Focal Points are:

KDDH	Bennington, VT		John Durant (at ALY)		518-435-9571

KPWM	Portland, ME		Jim Mansfield (at GYX)		207-688-3223 x225

KGFK	Grand Forks, ND		Mike Lukasz 			701-795-5198 x372

KHYR	Hayward, WI		Donald Price (at DLH)		218-729-0645

KITR	Burlington, CO		Grady Bonsall (at GLD)		785-899-2360 x381

KLHX	La Junta, CO		Chris Price (at PUB)		719-948-9429

KLND	Lander, WY		Rob McFall (at RIW)		307-857-3898 x372

KMFI	Marshfield, WI		Rob Hoag (at GRB)		920-494-5845 x372

KMHE	Mitchell, SD		Don Beck (at FSD)		605-330-4244 x372

KOFK	Norfolk, NE		Larry Howard (at OAX)		402-359-2394 x381

KVPZ	Valparaiso, IN		Walter Cowan (at LOT)		815-834-0600 x372

KDHT	Dalhart, TX		Ken Hunter (at AMA)		806-335-1835 x260

KGUY	Guymon, OK		Ken Hunter (at AMA)		806-335-1835 x260

KBKE	Baker, OR		Michael Pereira (at BOI)		208-334-9865

KCEC	Crescent City, CA		Donald Dodd (at EKA)		707-443-5610

KDLS	The Dalles, OR		Johnny Blagg (at PDT)		541-276-7832

KINW	Winslow, AZ		Steven Kusyj (at FGZ)		928-556-9161 x260

KMYV	Marysville, CA		Douglas Swann (at STO)		916-979-3051

KOKB	Oceanside, CA		Michael Lauderdale (at SGX)	858-675-8700 x260

KRNM	Ramona, CA		Michael Lauderdale (at SGX)	858-675-8700 x260

KSFO	San Francisco, CA		Wayne Bailey (at MTR)  	831-656-1710

KSHN	Shelton, WA		Kris Johnson (at SEA)		206-526-6090 x260

KSJN	St. Johns, AZ		Steven Kusyj (at FGZ)		928-556-9161 x260

PALH	Lake Hood, AK		Jim Hunter			907-271-5124

PANC	Anchorage, AK		Jim Hunter			907-271-5124

Specific site characteristics and interfaces are listed in the table below.

.

ASOS V2.79C OAT Site Characteristics and Interfaces

| SID | Name | FAA Service Level | Staffing | DCPs | Multiple Sensors | Comms | ZR | TSTM/ ALDARS | GTA/ ATIS | ACE | RVR | WSP | IFW | AWPAG | Current Software Version |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| KDDH | Bennington, VT | D | --- | 1 | --- | ADAS | ZR | ALDARS | GTA | --- | --- | --- | IFW | --- | 2.79B |
| KPWM | Portland, ME | C | PT | 2 | M/B | ADAS | ZR | ALDARS | --- | --- | --- | --- | IFW | AWPAG | 2.79B |
| KGFK | Grand Forks, ND | B | FT | 1 | --- | ADAS | ZR | ALDARS | --- | --- | --- | --- | IFW | AWPAG | 2.79B |
| KHYR | Hayward, WI | D | --- | 1 | --- | ADAS | ZR | ALDARS | --- | -- | --- | --- | IFW | --- | 2.79B |
| KITR | Burlington, CO | D | --- | 1 | --- | ADAS | ZR | ALDARS | GTA | --- | --- | --- | IFW | --- | 2.79B |
| KLHX | La Junta, CO | D | --- | 1 | --- | ADAS | ZR | ALDARS | GTA | --- | --- | --- | IFW | --- | 2.79B |
| KLND | Lander, WY | D | --- | 1 | --- | DIAL | --- | TS | --- | --- | --- | --- | IFW | AWPAG | 2.79B |
| KMFI | Marshfield, WI | D | --- | 1 | --- | ADAS | ZR | ALDARS | GTA | --- | --- | --- | IFW | --- | 2.79B |
| KMHE | Mitchell, SD | D | --- | 1 | --- | ADAS | ZR | ALDARS | --- | --- | --- | --- | IFW | --- | 2.79A |
| KOFK | Norfolk, NE | C | --- | 1 | --- | ADAS | ZR | ALDARS | GTA | --- | --- | --- | IFW | AWPAG | 2.79B |
| KVPZ | Valparaiso, IN | D | --- | 1 | --- | DIAL | ZR | --- | --- | --- | --- | --- | IFW | --- | 2.79B |

ASOS V2.79C OAT Site Characteristics and Interfaces

| SID | Name | FAA Service Level | Staffing | DCPs | Multiple Sensors | Comms | ZR | TSTM/ ALDARS | GTA/ ATIS | ACE | RVR | WSP | IFW | AWPAG | Current Software Version |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| KDHT | Dalhart, TX | D | --- | 1 | --- | ADAS | ZR | ALDARS | GTA | --- | --- | --- | IFW | --- | 2.79B |
| KGUY | Guymon, OK | D | --- | SCA | --- | ADAS | --- | ALDARS | GTA | --- | --- | --- | IFW | AWPAG | 2.79B |
| KBKE | Baker, OR | D | --- | 1 | --- | ADAS | ZR | ALDARS | --- | --- | --- | --- | IFW | --- | 2.79B |
| KCEC | Crescent City, CA | D | --- | 1 | --- | ADAS | --- | ALDARS | GTA | --- | --- | --- | IFW | --- | 2.79B |
| KDLS | The Dalles, OR | D | --- | 1 | --- | ADAS | ZR | ALDARS | GTA | --- | --- | --- | IFW | --- | 2.79B |
| KINW | Winslow, AZ | D | --- | 1 | --- | DIAL | ZR | --- | GTA | --- | --- | --- | IFW | AWPAG | 2.79B |
| KMYV | Marysville, CA | D | --- | 1 | --- | ADAS | --- | ALDARS | GTA | --- | --- | --- | IFW | --- | 2.79B |
| KOKB | Oceanside, CA | D | --- | 1 | --- | ADAS | --- | ALDARS | GTA | --- | --- | --- | IFW | --- | 2.79B |
| KRMN | Ramona, CA | C | PT | 1 | --- | ADAS | --- | ALDARS | --- | --- | --- | --- | IFW | AWPAG | 2.79B |
| KSFO | Sam Francisco, CA | A | FT | 3 | M/B | DIAL | --- | --- | --- | ACE | RVR | --- | IFW | AWPAG | 2.79B |
| KSHN | Shelton, WA | D | --- | 1 | --- | ADAS | ZR | ALDARS | GTA | --- | --- | --- | IFW | AWPAG | 2.79B |
| KSJN | St. Johns, AZ | D | --- | 1 | --- | ADAS | ZR | ALDARS | GTA | --- | --- | --- | IFW | --- | 2.79B |
| PALH | Lake Hood, AK | C | --- | 2 | Remote wind | ADAS | ZR | --- | --- | --- | --- | --- | --- | --- | 2.7B |
| PANC | Anchorage, AK | A | FT | 2 | M | ADAS | ZR | --- | --- | --- | RVR | --- | IFW | AWPAG | 2.7B |