U.S. DEPARTMENT OF COMMERCE

National Oceanic and Atmospheric Administration

NATIONAL WEATHER SERVICE

1325 East-West Highway

Silver Spring, Maryland 20910-3283

MEMORANDUM FOR:   Distribution

FROM:		    W/OPS2 – John Van Kuren (Signed November 14, 2006)

SUBJECT:		    Initiation of the Operational Acceptance Test (OAT) for

Automated Surface Observing System (ASOS) Software

Version (V) 2.79C

ASOS Operational Trouble Report (OTR) #1078 dated 09/18/06 identifies a condition where an erroneous Ice Free Wind (IFW) sensor maintenance flag ($) is appended to Aviation Routine Weather Reports (METAR) or Aviation Selected Special Weather Reports (SPECI).  This occurs when the IFW sensor correctly reports an error for path 5 (one of the six paths in the IFW sensor), but the path error is incorrectly interpreted by the Acquisition Control Unit (ACU) software and an erroneous $ is appended to the next ASOS METAR or SPECI.  Although the IFW sensor is operating normally (it only needs data from four of the six paths for a valid wind report) and no data quality error exists, this erroneous $ might mask other potential ASOS problems.  In addition, the ASOS Operations and Monitoring Center (AOMC) has experienced an increased workload in responding to each erroneous $, and in some cases, IFW sensors which were functioning properly were replaced unnecessarily.  The problem is present in V2.79 and subsequent ACU software versions.

V2.79C (10/05/06) is identical to V2.79B (09/19/05) with the exception of a fix for the

erroneous $ problem.  The fix has been confirmed in a System Test (ST) at the Sterling Research and Development Center (SR&DC), Sterling, VA, and the V2.79C software was also subjected to a battery of regression tests to confirm ASOS functionality.

The V2.79C OAT plan is attached and is also posted at:

http://www.nws.noaa.gov/ops2/ops24/documents/asos_V2-79C.htm

The V2.79C software is available on the OPS12 ftp site and ASOS Modification Note 80A is posted on the ASOS Electronics Technician website.

Installation at the OAT sites listed below is authorized to begin during the week of November 13, and evaluation of ASOS performance by NWS Headquarters, on-site, or nearby staff will continue for approximately two weeks.  A Test Review Group, made up of representatives from Observing Services Division (OS7), Observing Systems Branch (OPS22), Maintenance Branch (OPS12), and Aviation Services Branch (OS23) will evaluate any reported problems as well as overall performance.  If no problems are noted during the OAT, V2.79C will be released immediately for all operational ASOSs with an IFW sensor currently installed.

Installation of V2.79C at the following OAT sites is now approved:

Eastern Region				Southern Region

KDDH	Bennington, VT			KDHT	Dalhart, TX

KPWM	Portland, ME			KGUY	Guymon, OK

Central Region				Western Region

KBKE	Baker, OR

KGFK	Grand Forks, ND			KCEC	Crescent City, CA

KHYR	Hayward, WI			KDLS	The Dalles, OR

KITR	Burlington, CO			KINW	Winslow, AZ

KLHX	La Junta, CO			KMYV	Marysville, CA

KLND	Lander, WY			KOKB	Oceanside, CA

KMFI	Marshfield, WI			KRNM	Ramona, CA

KMHE	Mitchell, SD			KSFO	San Francisco, CA

KOFK	Norfolk, NE			KSHN	Shelton, WA

KVPZ	Valparaiso, IN			KSJN	St. Johns, AZ

Alaska Region

PALH	Anchorage (Lake Hood), AK

PANC	Anchorage, AK

If you have any questions or comments, please contact the test coordinator Bryan Moore, W/OPS24, at 301-713-0326 x176.

Distribution (email):

FAA - Calvin Smith

CIO12 – K. Conaty

OS7 – J. Heil

OS23 – K. Schlachter

OST11 – B. Giza

OPS12 – A.Wissman

OPS12 – G. Dalyai

OPS22 – D. Mannarano

ER41 – H. Machado

CR43 – B. Brashears

SR41 – M. Garcia

WR4 – G. Deiotte

AR42 – J. Hunter

PR – A. Lowe

CC:

FAA - Bing Huang

E/CC11 - M. Urzen

CIO12 – S. Murphy

OS12 – L. Kozlosky

OPS – J. McNulty

OPS1 – M. Paese

OPS12 – A. Leonardo

OPS13 – M. Detommaso

OPS22 – J. Facundo

OPS22 – R. Parry

OPS23 – R. Thomas

OPS23 – P. Hoch

OPS24 – J. Dinges

OPS24 – M. Buckingham

OPS24 – J. Lee

ER42 – K. Murray

CR1 – T. Townsend

SR11 – V. Murphy

SR41 – J. Villescaz

WR42 – J. Lachacz

AR41 – E. Doerr