Pixel Modules

There are 1744 modules in the pixel detector. A schematic view of a barrel pixel module is shown in Figure M1 and a picture of a disk pixel module in Figure M2. A pixel module consists of the following components:

- a sensor tile 63.4 x 24.4 mm2, about 250 microns thick

- 16 front-end electronics chips, each with 2880 electronics channels

- bump bonds (PbSn or In) that connect the electronics channels to pixel sensor elements

- a copper-on-kapton flexible printed circuit board (flex-hybrid)

- a module-control-chip (MCC) attached to the flex-hybrid

- copper-on-kapton pigtail with connector (barrel modules) or a micro-cable (disk modules) connected to the flex-hybrid

Figure M1 A schematic view of barrel pixel module.

Sensor tiles and electronics chips are connected by bump bonding technology to form bare modules. Both solder(PbSn) and indium bump technologies were used to make pixel modules. The production statistics of bare modules are summarized in Table M1.

|  | In | PbSn | Total |
| --- | --- | --- | --- |
| Assembled | 1468 | 1157 | 2625 |
| Accepted | 1296 | 1122 | 2418 |
| Rejected | 172 | 35 | 207 |

Table M1 Summary of production bare modules.

The number of accepted modules includes reworked modules (13% for In and 8% for PbSn) in which a faulty front-end chip was removed and replaced.

Figure M2 A picture of a pixel disk module.

Flex hybrids with attached and tested MCC chips were glued to accepted bare modules. Losses occurred during this step, leaving 2312 modules for electrical and final mechanical characterization.

Pixel modules were ranked according to electrical testing and mechanical criteria. Extensive electrical testing was done at room temperature and at -10oC, approximately the operating temperature. Thermal cycling was performed on each module prior to completion of electrical testing. The ranking also took into account repair procedure (wire bond replacement or other) and other features. The ranking procedure was used to separate the modules into those acceptable for the B-layer r(highest quality) and for the outer barrel and disk regions. Modules with the best ranking were loaded on barrel staves and disk sectors, leaving modules with the poorest ranking as spares. A brief summary of the yields and ranking after testing is given in Table M2.

|  | In | Solder | Total |
| --- | --- | --- | --- |
|  | Number | % | Number | % | Number | % |
| Assembled | 1190 |  | 1122 |  | 2312 |  |
| Accepted | 1025 | 86.1 | 1075 | 95.8 | 2100 | 90.8 |
| B-layer quality | 281 | 23.6 | 445 | 39.7 | 726 | 31.4 |
| Not B-layer quality | 744 | 62.5 | 630 | 56.1 | 1374 | 59.4 |
| Rejected | 165 | 13.9 | 47 | 4.2 | 212 | 9.2 |

Table M2 Summary of yield and ranking of tested pixel modules.