Requirements/Specifications for a Low-Energy Proton Source

Technical:

- Primary beam to be singly ionized H+ (protons).  Ion source capable of producing ion beams of other elements desired.  Two complete, identical ion sources required. (One will serve as a “quick-replacement “spare).

- Proton Energy Range: 200 eV to 30 KeV with delta E 50 eV.

- Ion mass/velocity analysis capability.

- Beam current density range; 1E8 to 5E11 protons/sq.cm/s.

- Include personnel safety and EMI shielding, as required, for ion source.  Ion source should produce microamp current at the source.

- Minimum raster-scanned beam coverage area of 12 inch diameter on target with coverage uniformity of 95%, or better.  Minimum separation distance between source and target is 4 feet, maximum separation is 6 feet.

- Control function switchable between computer and analog controller

- All mating flanges to be Conflat-design, no larger than 10” OD.

System:

- Projected maintenance interval to be not less than 1000 hours of operating time.

- Non-restricted system mounting orientation.

- Support (ancillary) hardware includes but is not limited to:

  - All necessary system control and monitoring equipment

  - Beam line support system, if required

  - Faraday Cup, controller and readout.

  - All required power supplies and cables.

- Remote operation (from approximately 100 feet)

- All required system cables and connectors.

- All equipment to operate on 110VAC.

General:

- The contractor shall deliver Line Item 1 to NASA’s George C. Marshall Space Flight Center (MSFC) within 6-months after receipt of award.  As part of Line Item 1, the contractor shall provide a maintenance kit (spare parts, special tools, etc) as well as two copies of operating manuals, service manuals, drawings, parts lists, wiring diagrams/lists, and system installation drawings including system power requirements.  The offeror shall provide any and all maintenance (including parts and/or labor) that is covered in any maintenance agreement that comes standard with the purchase of Line Item 1.

- Under Line Item 2, the contractor shall install Line Item 1 in Building 4605 at MSFC and successfully complete an acceptance test within one week after delivery.  The acceptance test shall demonstrate Line Item 1 meets or exceeds the requirements/specifications stated herein.  The contactor shall also provide a minimum of one day of training at MSFC for up to 6 people.  The training shall cover the procedures on how to operate Line Item 1.

- Under Line Item 3, the contractor shall provide all parts and labor required to maintain Line Item 1 of not less than three years after successful completion of the acceptance test, when combined with any maintenance agreement that comes standard with the purchase of Line Item 1.  The minimum maintenance required to be performed on Line Item 1 shall be consistent with the recommendations and warranty requirements of the manufacturer.  The maintenance required to be performed under Line Item 3 is limited to the period of time between the not otherwise covered by any maintenance agreement that comes standard with the purchase of Line Item 1.