VAISALA PTB-201 SETUP

FOR ALL CAMPBELL LOGGERS

The Barometers are serial devices so to connect you will need:

A computer

A 9-pin straight through serial cable

A 12-volt supply for power to the Barometer (logger)

A serial communication programs such as HyperTerminal

A special built connector so 12V and ground can remain on the logger and

TX/RX/GND can be connected to the communications port of laptop.

- Apply Power to the Barometer and start HyperTerminal.  If this is the first time setting up communications to the Barometer to change settings the following HyperTerminal settings should be used:

  - Baud Rate: 1200

  - Data Bits:   7

  - Parity:        Even

  - Stop Bits:   1

  - Flow control: None

If system already in use the following settings should be used:

SGP/TWP		NSA

  - Baud Rate: 1200		1200

  - Data Bits:   8			7

  - Parity:        None		Even

  - Stop Bits:   1			1

  - Flow control: None		None

- The following commands will change the settings in the Barometer to match what the Campbell Dataloggers need: DO NOT TYPE QUOTATIONS

  - Type “OPEN”

  - Type “ECHO ON”

  - Type “FORM <CR>”

The unit will return the following:

“\PPPP.PP\ \uuuu\\r\n”

?

Or “\PPPP.P\ \uuuu\\r\n

?

The cursor will be after the question mark

Type “\PPPP.P\ \uuuu\\r\n<CR>”

There is a space between the 2nd and 3rd backslash.  You MUST type the format after the question mark regardless of the current form.  Typing FORM <CR> puts barometer into the “change format mode”

  - Type “UNIT hPa” (probably not needed—but good to check)

  - Type “SERI 1200 N 8 1”

  - Type “ECHO OFF”

  - Type “?” and make sure the following are all the same:

Software version             PTB 200 / 2.03

Baud Parity Data Stop     1200    N 8 1

Output format		     \PPPP.P\\uuuu\\r\n

Pressure unit                    hPa

Temperature unit             'C

Sending mode                  STOP

Address                            0

Sleep mode                      OFF

Measurement time          64 x 25 milliseconds

Filter                                OFF / FAST

Output interval                 0 s

Calibration day code	    92261

Offset drift comp.	    ON

Offset drift correction	    -0.056

Multipoint correction 	    ON

  - If any of the 2nd – 11th settings are not the same as above, see manual to set as above.  The last 4 settings may differ between barometers.

  - Type “RESET” to save all settings.

  - Type “CLOSE”.