# NuMI BPM Sign Convention and Signal Path

P. Prieto

4/28/04

This note describes the sign convention used in the NuMI BPM system and the labeling signal channels between the BPM’s in the tunnel and the electronics in MI60S, MI60N, and MI65.

# System Convention:

The sign convention in NuMI BPM system matches the sign convention of BPM’s located in the Main Injector. Measured positions delivered from the VME front-end system have a positive sign for beam positions radially outside at horizontal locations and positions above center for vertical locations. This fixes the polarity of signals as POSITIVE  values corresponding to beam UP or OUT and NEGATIVE values corresponding to beam DOWN or IN.

This same convention extends into  the beam line going toward the NuMI target.   Measured positions at horizontal locations will have a positive sign for beam to the right of center (radially outside) when looking down the NuMi tunnel facing the target. Vertical locations will have a positive sign for beam positions above center. This means that POSITIVE measurement values correspond to beam UP and to the RIGHT (as defined facing the NuMI target with your back to the Main Injector) and NEGATIVE values correspond to beam DOWN and to the LEFT.

BPM:

NuMI BPM’s are two-electrode electrostatic-type pickups with a diagonally split cylinder of 3.85 in. diameter and 8.0 in. in length. The same design, rotated appropriately, is used for both horizontal and vertical locations. There is a doghouse on the body of the pickup where a ceramic feed-thru connects each of the pickup plates to an N connector. Not all BPMs are installed in the beam line with the dog-house in the same upstream/downstream, right/left, or top/bottom orientation. Nevertheless, at ALL vertical installations the upper and lower N connectors attach to the top and bottom electrodes respectively and similarly, at ALL horizontal installations, the radial inside and outside connectors attach to the inside and outside electrodes respectively.

Cable Connections to BPM:

The BPM signal cable labeling convention is a three row label with the first row indicating the cable source, the second row the cable destination rack and the third row the signal name. At horizontal and vertical locations, cables are labeled for example, hp106 BPMA and hp106 BPMB indicating the BPM name and the plate to which its attached.

Care must be taken when connecting the cables to the BPMs in the tunnel that the polarity convention is maintained regardless of the BPM Nconnector doghouse orientation. At vertical BPM locations, the doghouse will face out either the right or left side of the BPM. In either case, the BPMB labeled cable should be attached to the upper N connector and the BPMA labeled cable to the lower N connector.  At horizontal locations, the doghouse will face either up or down. In either case, as you look down the beam line facing the NuMI target, the BPMB labeled cable should be attached to the right-hand N connector and the BPMA labeled cable to the left-hand N connector.

The cables from 13 BPM’s go to MI60N and 12 BPM’s to MI65 building.

The connections for MI0South originate from MI BPM’s 02, 604, 605 and 606. Signal splitters have been added to the existing RG-8 cables using RG-58 cables going to rack XXX where an 8 slot VME^$X crate with a single DDC card downconverts the MI signals.

Building and Top Plate Cabling:

For MI60N (rack #MI60203) and MI65 (rack #MI65105) there is a rack assigned and a top-plate  N-type feed-thru. Each of the BPM signal cables from the tunnel terminates at the rack top plate. A drawing will be made of the top plate assignments.

An N-to-SMA  transition cable, type LMR-195 and 36 inches in length, connects the signals from the top plate to the transition module.

Transition and Calibration Module:

The NuMi PBM transition module (dwg. 9820-ED-3701755) is a four-channel board capable of handling pairs of signals from two BPM’s. A VME-style crate will hold up to 7 transition modules in addition to one calibration board that distributes the 53 MHz and 2.5 MHz calibration signals to each of the transition modules. Inputs to the transition board are made through four, rear edge-mounted SMA connectors. Input SMA in the rear of the transition module are not labeled but the top SMA corresponds to channel A1, second channel B1, third channel A2 and fourth SMA channel B4. Therefore cables with labels BPMA or BPMB will be matched to either the A channel or B channels in the transition board.  Outputs from the transition board going to the DDC board are from front panel SMA’s  ,top to bottom, labeled A1, B1, A2, B2.

DDC:

The Echotek GC814 digital receiver board has eight independent digitizer-down-converter-FIFO channels with front panel inputs labeled 0-7. Transition module “B” outputs connect to odd-numbered Echotek input channels and “A” outputs connect to even-numbered Echotek input channels. This maps the BPM pickup “B” plate to the upper input and the “A”plate to the lower input of an Echotek even-odd channel pair.

Signal cables between the transition module and the Echotek DDC board will be labeled with the BPM name at both ends, such as VP118-A and VP118-B. These cables will be type FMR-195, 36 inches long, with SMA’s at each end.

Software:

NuMi BPM software is based on Recycler software and is currently being implemented.

# Text boxes

TOP/RIGHT

A

A

B

Position

(A-B)/(A+B) *

*Scale Factor + Offset

Software

GC-814

Transition

Module

BOTTOM/

LEFT

B

0

1

2

3

6

7