July 22, 1994

Contract # TW0592

SPECIAL SPECIFICATION

SECTION 15493

LAMINAR FLOW BENCHES/FUME HOOD

Page

PART 1 - GENERAL

1.01 Description.	2

1.02Related Work..	2

1.03Quality Assurance.	2

1.04References.	2

1.05Submittals.	2

1.06 Delivery, Storage, and Handling.	3

1.07 Project/Site Conditions.	4

PART 2 - PRODUCTS

2.01Acceptable Manufacturers.	5

2.02 Equipment.	5

PART 3 - EXECUTION

3.01 Inspection.	8

3.02 Installation/Application/Erection.	8

3.03 Field Quality Control.	8

3.04 Adjusting and Cleaning	9

3.05 Commissioning.	9

3.06 Demonstration.	9

SPECIAL SPECIFICATION

SECTION 15493

LAMINAR FLOW BENCHES/FUME HOOD

PART 1 - GENERAL

1.01	DESCRIPTION

A.  	This Section covers requirements for providing laminar flow benches/fume hood with pre-piped and pre-wired connections.

1.02	RELATED WORK

A.  	Section 01300

B.  	Section 15010

C.  	Section 15892

D.  	Section 15994

E.  	Section 15995

1.03	QUALITY ASSURANCE

A.  	Certify the laminar flow bench performance according to Federal Standard 209.

1.04	REFERENCES

A.  	Federal Standard 209

B.  	National Fire Protection Association (NFPA) 70, National Electrical Code (NEC).

1.05	SUBMITTALS

A.  	Make submittals according to this Section, Section 15010, and Section 01300.

B.  	Product data for each laminar flow benches/fume hood:

1.  	Unit design requirements.

2.  	Unit features.

3.  	Weights (shipping, installed, and operating).

4.  	Furnished accessories.

5.  	Installation and start-up instructions.

C. 	Shop drawings detailing fabrication and installation of laminar flow benches/fume hood including plans, elevations, sections, component details, attachments, color finishes, and other construction elements.  Include the following:

1.  	Dimensions.

2.  	Weight loadings and distribution.

3.  	Clearances for maintenance and operation.

4.  	Size and location of field connections.

D.  	Wiring diagrams detailing power and control wiring.  Differentiate clearly between manufacturer-installed wiring and field-installed wiring.

E.  	Maintenance data for each laminar flow benches/fume hood.

F. 	Certification of performance from the factory.

G.  	Certification letter from manufacturer stating that the laminar flow benches/fume hood have been through start-up procedures and that they are functioning properly.

1.06	DELIVERY, STORAGE, AND HANDLING

A.  	See Section 15010.

B.  	Protect laminar flow benches/fume hood from damage by factory packing.

C.  	Reject any damaged laminar flow benches/fume hood upon arrival.

D.  	Store laminar flow benches/fume hood to prevent damage, and protect from weather, dirt, fumes, water, and construction debris.  Provide a clean dry space.

E.  	Handle laminar flow benches/fume hood according to the manufacturer's rigging and installation instructions for unloading and transporting into the final location.

1.07	PROJECT/SITE CONDITIONS

A.  	Coordinate layout and installation of the laminar flow benches/fume hood with the piping system and exhaust ductwork.  Revise locations to suit field conditions and as approved by the Architect/Engineer (A/E).

B.  	Coordinate installation with laboratory furniture installation.

PART 2 - PRODUCTS

2.01	ACCEPTABLE MANUFACTURERS

A.  	Liberty Industries, Inc. or approved substitute

2.02	EQUIPMENT

A.  	General

1.  	Provide a blower, filter unit, sash (or cover), work surface, frame/chassis, side panels, base cabinets (or support legs), lights, power cord, and services for the laminar flow benches/fume hood as indicated in this Section.  Provide features and services as indicated in Tables 1, 2, and 3.

2.  	Provide UL labeled components for the units.

3.  	Provide a unit that delivers air across the entire cross-section in accordance with Federal Standard 209d (90 +/- 20 fpm).

B.  	Physical Characteristics

1.  	Construct the blower to be dynamically balanced without the use of speed controls.  Provide a 115-Volt, 1-phase, 60-Hertz motor for the blower.  Size the blower assembly to provide air flow with a minimum 50 percent increase in the HEPA filter pressure drop.

2.  	Provide vibration isolation and seismic snubbers for the blower according to Section 15200-S.

3.  	Design the laminar flow bench noise generation to be less than the Occupational Safety and Health Administration's required level.

4. 	Provide industry-standard, 99.99 percent efficient, HEPA filter units per IES-RP-CC001.3 which are inter-changeable with filters of the same size on other laminar flow benches and the laminar flow fume hood.

5.  	Provide minimum 16-gage (ga.) metal; use 304 or 316 stainless steel where stainless steel is identified.

6.  	Provide pre-piped and labeled piping connections, so all connections are made exterior to the laminar flow bench.  Provide pre-wired power connections with a pre-wired, 3-prong, 8-foot long, power cord.

7.  	Construct the laminar flow bench exterior of welded cold-rolled steel or stainless steel component parts as indicated in Table 1.  Provide access panels which are attached with screws to allow for removal of the panels.  Provide front, rear, and side access panels for maintenance.  Provide for front removal of the filters through the access panels.  Provide steel component parts with an acid, alkali, and solvent resistant finish on both exterior and interior surfaces after fabrication of all cold-rolled steel parts and prior to final assembly.  Provide approved finish color.

8.  	Provide at the top of the laminar flow bench a fluorescent or incandescent light fixture as indicated in Table 2.  Provide a fixture shielded from the hood interior by a tempered glass panel sealed to the hood by an extruded vinyl channel.  Provide electronic ballasts and T-8 lamps for fluorescent fixtures.

9. 	Provide a sash (or cover) for the hood as indicated in Table 1.  Provide a 16-mil transparent vinyl plastic film with snaps or 1/4-in. transparent lexan for the sash as indicated in Table 1.  Provide a stainless steel piano hinge full length for the sash (or cover) which permits the sash (or cover) to be lifted from the front of the laminar flow bench for easier access during equipment placement.  Provide hooks, clamps, or snaps to hold the sash (or cover) up during this equipment placement.

10. 	Construct the laminar flow benches/fume hood to permit passage through a 3-foot by 7-foot personnel door.

11. 	Design the laminar flow fume hood exhaust to be down through the working surface and out the back of the base cabinet.

12. 	Interior Surface Construction

a. 		Provide laminar flow bench interior materials as indicated in Table 1.  Provide minimum 16-ga. metal, 304 or 316 stainless steel where stainless steel is indicated.  Weld stainless steel materials together.  Triple weld polypropylene materials together.  Provide neoprene gaskets to join dissimilar materials.

b. 		For areas exposed to the laminar flow bench interior, use stainless steel fasteners on removable parts.

c. 		Provide a smooth finish on the interior surfaces to reduce dirt accumulation.

d. 		Provide 14-ga. reinforced stainless steel for the work surface as indicated in Table 1.

e. 		Provide reinforced working surface capable of supporting 100 pounds per square foot without damage or warping.

13. 	Features

a. 		Provide features as indicated in Tables 2 and 3.

b. 		Provide a Dwyer "Magnehelic" Gauge or approved substitute for the air filter gauge.  Provide a scale range which exceeds the system operating range by 100 percent both above and below the operating range limits unless noted otherwise on the Drawings.  Provide 2, 5, or 10 unit scale divisions as appropriate to system range.  Provide in. w.g. divisions.

14. 	Service Locations

a. 		See Table 3.

b. 		Locate the duplexes, remote valve operators, and light switches on the laminar flow bench front exterior surface as identified in Table 3.  Replace duplexes with a power strip having equal number of outlets and same protection, if so desired.

c. 		Provide 3/8-in. service valves with remote serrated hose connection located on the laminar flow bench work surface.

PART 3 - EXECUTION

3.01	INSPECTION

A. 	Inspect areas and conditions, with Installer present, for compliance with requirements for installation tolerances, and other conditions affecting performance of the laminar flow benches/fume hood.

B.  	Inspect rough-in for utilities to verify actual locations of connections prior to installation.

C. 	Do not proceed until unsatisfactory conditions have been corrected.

3.02	INSTALLATION/APPLICATION/ERECTION

A.  	General

1.  	Install laminar flow benches/fume hood level and plumb, in accordance with manufacturer's written instructions.

2.  	Arrange installation of units to provide access space for service and maintenance.

B. 	Connections

1.  	Piping installation requirements are specified in other Division 15 sections.  The Drawings indicate the general arrangement of piping, valves, fittings, and specialties.

2.  	Electrical Connections:  The following requirements apply:

a. 		Connect unit components to ground in accordance with the NEC.

3.03	FIELD QUALITY CONTROL

A.  	Arrange and pay for a factory-authorized service representative to perform the following:

1.  	Inspect the field assembly of components and installation of the laminar flow benches/fume hood including piping and electrical connections.

2.  	Inspect the blower operation and confirm proper installation of the HEPA filters.

3.  	Prepare a written report on findings and recommended corrective actions.

3.04	ADJUSTING AND CLEANING

A.  	Adjust sashes for proper operation.

B.  	Clean unit laminar flow bench interiors to remove foreign material and construction dirt and dust.

3.05	COMMISSIONING

A.  	Provide labor, material, equipment, etc., required to facilitate the commissioning process.  Perform form tests and verification procedures required by the commissioning process when requested by the Commissioning Authority and directed by the General Contractor.  See Section 15995.

3.06	DEMONSTRATION

A.  	Provide a factory-authorized service representative supervise start-up of equipment, to demonstrate operation, and to train personnel.

B.  	See Section 15995.

END OF SECTION

EXCEPT FOR ATTACHMENTS

LAMINAR FLOW BENCHES/FUME HOOD SPECIFICATION TABLES 1, 2 & 3

LAMINAR

FLOW

BENCHES/FUME

HOOD

SPECIFICATION

TABLE 1

Dimensions                              Construction                                        Basis of Design

Equip.Room                                                            Material

No.

LFB-

B =     PP = polypropylene,

Back,   PVC = polyvinyl

B =     chloride,

blue    S = slate,

oil-    SS = stainless

based   steel,

paint,  VPF = snapdown vinyl plastic

ER =    film,

epoxy   WS = white

(chemicalenameled

resistant)steel

resin,

Fr =

Front,

G = glass,

L = lexan,

PG = plexiglas,

PL =

plastic

laminate,

Maximum  Maximum Working Working  Opening   Sash  Working Frame/Drain    Side       Base                 Filter

Unit     Unit  Surface Surface   Height    or   SurfaceChassis        Panels    Cabinet                 Unit

Width    Height  Height  Depth    Inches  Cover                                     or                 Housing

Inches   Inches  Inches Inches                                                   Support

Legs

1     1404     62.5     102      35      30       36    Fr, L     SS     SS    -       G          SS                    SS

2     1303     62.5     102      35      30       36     Fr,      SS     SS    -       G          SS                    SS

VPF

3     1303     62.5     102      35      30       36     Fr,      SS     SS    -       G          SS                    SS

VPF

4     1303     50.5     102      35      30       36     Fr,      SS     SS    -       G          SS                    SS

VPF

5     1303     50.5     102      35      30       36     Fr,      SS     SS    -       G          SS                    SS

VPF

6     1303     50.5     102      35      30       36     Fr,      SS     SS    -       G          SS                    SS

VPF

7     1303     50.5     102      35      30       36     Fr,      SS     SS    -       G          SS                    SS

VPF

8     1303     50.5     102      35      30       36     Fr,      SS     SS    -       G          SS                    SS

VPF

9     1401     98.5     102      35      30       36     Fr,      SS     SS    -       G          SS                    SS

VPF

10    1401     98.5     102      35      30       36     Fr,      SS     SS    -       G          SS                    SS

VPF

11    1401     98.5     102      35      30       36     Fr,      SS     SS    -       G          SS                    SS

VPF

12    1106     62.5     102      30      24       22     Fr,      SS     SS    -       G          WS                    WS

VPF

19    1504     74.5     102      30      30       30      -       SS     SS    -       L          WS                    WS

20    1504     74.5     102      30      30       30      -       SS     SS    -       L          WS                    WS

21    2207     74.5     102      30      36       21      -       PL     SS    -       PG         B                     B

22    2207     74.5     102      30      36       21      -       PL     SS    -       PG         B                     B

23    2207     74.5     102      36      36       36      -       PL     SS    -       PG         WS                    WS

24    2207     74.5     102      30      36       21      -       PL     SS    -       PG         B                     B

25    2207     74.5     102      30      30       21      -       PL     SS    -       PG         B                     B

26    2205     74.5     102      36      30       36      -       PL     SS    -       PG         WS                    WS

27    2205     74.5     102      36      30       36      -       PL     SS    -       PG         WS                    WS

LFH-   1403     74.5     102      36      28       30     VPF      SS     WS    -       SS         WS                    WS

1

LAMINAR

FLOW

BENCHES/FUME

HOOD

SPECIFICATION

TABLE 2

Equip.   Room  Features

No.

LFB-

Support                                                           Air Flow                       Air        Air      Work Surface

Structure                         Lighting      Sink    Drain      Pattern                       Intake    Exhaust     Perforated

Location   Location

(125

footcandles

at the work

surface)

Support    Base

Legs   Cabinet

Y = Yes  Y = Yes  D =  Doors  F =            Y = Yes  Y = Yes  D = Down      I = Air Flow     T = Top    T = Top    Y = Yes

Dr =        Fluorescent                                    Into

N = No   N = No   Drawers                    N = No   N = No   H =               Bench        B =        B =        N = No

Ds = Doors  I =                              Horizontal    O = Air Flow     Bottom     Bottom

&           Incandescent                                       Out of

Shelves                                                    Bench

S =                                                        E = Exhaust

Shelves

1      1404     N        Y          D            F           N        N           D              O             T          -            N

2      1303     N        Y          D            F           N        N           D              O             T          -            N

3      1303     N        Y          D            F           N        N           D              O             T          -            N

4      1303     N        Y          D            F           N        N           D              O             T          -            N

5      1303     N        Y          D            F           N        N           D              O             T          -            N

6      1303     N        Y          D            F           N        N           D              O             T          -            N

7      1303     N        Y          D            F           N        N           D              O             T          -            N

8      1303     N        Y          D            F           N        N           D              O             T          -            N

9      1401     N        Y          D            F           N        N           D              O             T          -            N

10     1401     N        Y          D            F           N        N           D              O             T          -            N

11     1401     N        Y          D            F           N        N           D              O             T          -            N

12     1106     Y        N          -            F           N        N           D              O             T          -            N

19     1504     Y        N          -            F           N        N           H              O             B          -            N

20     1504     Y        N           -           F           N        N           H              O             B          -            N

21     2207     N        Y          D            F           N        N           H              O             B          -            N

22     2207     N        Y           D           F           N        N           H              O             B          -            N

23     2207     Y        N          -            F           N        N           D              O             T          -            N

24     2207     N        Y          D            F           N        N           H              O             B          -            N

25     2207     N        Y          D            F           N        N           H              O             B          -            N

26     2205     Y        N          -            F           N        N           H              O             B          -            N

27     2205     Y        N          -            F           N        N           H              O             B          -            N

LFH-1    1403     N        Y          D            F           N        N           N              E             T          B            Y

LAMINAR

FLOW

BENCHES/FUME

HOOD

SPECIFICATION

TABLE

3

Features          Service

Equip.   Room                                                  Locations

No.

LFB-

Air Intake Filters                      Frame  Ar =   Duplex,   with on

Groundedargon, FP = foot pedal operator,airthe

Y = Yes, N = No                               BS =   LS = lightflow right,

blower switch,   rate 2R

switch,N2 =      gauge = 2

DPLX   nitrogen, L =  on

= 120  NPW =     1    the

V,     non-      on   right,

GFCI,  potable   the  etc.

Duplex,water,    left,

NDPLX  R'stat = Rheostat2L

= 120  to        = 2

V,     control   on

Non-   laminar   the

GFCI,  flow      left,

DI =   bench     etc.

deionizedair flowR =

water, rate      1

Roughing 30% NBS 99.99%Y = Grounding    N2       Ar    Duster

Filter EfficientEfficientYesTerminals                  Gun

on                                   Y =

0.3  N =                              Yes

micron No                             N = No

and

smaller

particles

per

IES-RP-CC001.3

DI    NPW    R'stat  NDPLX           DPLX             LS        BS      1    1404      N      N    Y      Y        L        -       -

Y      -       -       -              2L              -         L       L     2       1303    N    N      Y        Y        L       L

-      Y       -       -              -               2L        -       L     L        3     1303  N      N        Y        Y       L

L      -       Y       -              -               -         2L      -     L        L      4   1303    N        N        Y       Y

L      L       -       Y              -               -         -       2L    -        L      L    5     1303      N        N       Y

Y      L       L       -              Y               -         -       -     2L       -      L    L      6       1303      N       N

Y      Y       L       L              -               Y         -       -     -        2L     -    L      L        7       1303     N

N      Y       Y       L              L               -         Y       -     -        -      2L   -      L        L        8      1303

N      N       Y       Y              L               L         -       Y     -        -      -    2L     -        L        L       9

1401    N       N       Y              Y               L         L       -     Y        -      -    -      4L       -        L       L

10    1401     N       N              Y               Y         L       L     -        Y      -    -      -        4L       -       L

L      11     1401     N              N               Y         Y       L     L        -      Y    -      -        -        4L      -

L      L       12     1106            N               N         Y       Y     L        L      -    Y      -        -        -       -

2L     L       L       19            1504             N         N       N     N        -      L    -      Y        -        -       -

-      L       L       L              20             1504       N       N     N        N      -    L      -        Y        -       -

-      -       L       L              L               21       2207     N     N        Y      N    -      L        -        N       -

-      R       L       -              L               L         22     2207   N        N      Y    N      -        L        -       N

-      -       R       L              -               L         L       23   2207      N      N    Y      N        -        L       -

N      -       -       -              -               -         L       L     24      2207    N    N      Y        N        -       L

-      N       -       -              R               L         -       L     L        25    2207  N      N        Y        N       -

L      -       N       -              -               R          L      -     L        L      26  2205    N        N        Y       N

-      -       -       N              -               -         -       L     -        L      L    27    2205      N        N       Y

N      -       -       -              N               -         -       -     L        -      L    L    LFH-1     1403      N       N

Y      Y       L       L              L               N         R       R     -        -      2L   L      L