Excerpted from Transcript provided by

NEAL R. GROSS

COURT REPORTERS AND TRANSCRIBERS

1323 RHODE ISLAND AVE., N.W.

(202) 234-4433	WASHINGTON, D.C.  20005-3701	www.nealrgross.com

“Failures of endovascular grafts that have been observed clinically”

By Geoff White, MD

DR. WHITE:  Okay.  We are moving now from theory to reality, and as far as technology is concerned, this is where East meets West; California method here, and East Coast over here.

What I am going to do is just show some slides illustrating failure modes that have been encountered clinically in our own practice and elsewhere, using some of our own slides, and some that I have stolen from various publications.

So if I use your slide here, thank you very much.  Then I am going to go to the computer slides and just give a quick summary as to what I see as being the most important issues.

To start with the most common failure mode we have encountered is probably endoleak, and I think it is a good place to start because it illustrates one very important feature, and that is endoleak may occur because of device failure.

It may occur because of poor patient selection, and what somebody recently termed as patient failure.  It may occur because of a procedural failure by the implanting physician, and it may occur because of diverse, anatomic variations within the patient, as we have already touched on.

And just a quick summary of the current classification that we have been using, Type One of the attachment slides; and Type Two, any retrograde type of flow from side branch vessels, and Type Three, a defect in the graft wall or disconnection.

This type probably mainly depends on the patient, but it obviously can also occur because of anatomic factors.  To confuse all of this, we have variations in graft construction, which can give the appearance of endoleak, and with the earlier version of the endograft, we have this mixture of perhaps porosity of the graft fabric, and little channels through the graft where the sutures go through, which on early CT scans gave it this appearance of like a jet of blood coming out form the graft.

And this we had termed Type Four endoleak, because it seemed to close reliably within at least a month, and often quite earlier, but we are still not quite sure whether there may be long term consequences of these little channels.

This is an illustration of another type of endoleak, and I am going to run through these fairly quickly.  This is one of the earlier seam separations, and I think this was the stentograph, where the seam was sewn and the suture broke with perforation of the graft limb.

Here is another similar one, again with either the stent or the vanguard, showing leakage outside the graft, and this is the appearance after explant of the graft, and after attempted repair with another graft inside.  And again you can see the separation of the sutured seam.

Here is another type of graft erosion, or tear, or hole, a fairly subtle appearance on this view, but when you hold it up to the light, you can see the holes within the graft, which obviously can become significant channels, and can lead to repressurization of the graft, and much of the aneurysm.

And I think it was Juan Parodi earlier who emphasized the fact that once the aneurysm is deemed depressurized, it may be far more liable to early rupture when it is repressurized.

In the direction of the metal components, and the fabric components, they obviously can have the same effect.  And with all modular grafts which are constructed inside the body by connecting various pieces, we can later have disconnection.

This is a typical CT scan appearance, where one of the graft limbs is appearing sideways, like somebody crossing their knees, with the endoleak associated as you see here.  Here is the thrombus and here is the leak within the aneurysm sac.

This is how it appears on the plain x-ray, which is extremely valuable for diagnosing this kind of dislocation, and where you see the limb here and the trunk here and these two points should be in contact with each other.

And here is the angiographic appearance with contrast coming down and appearing in the aneurysm sac at that disconnection point.  Here is another form of disconnection, where approximal extension has become disconnected from the graft, and you see the two similar sized fragments here.

These are both the approximal ends of the graft and not limbs, and here they are laying side by side, instead of one inside the other.  Another common form of graft failure or complication is this kink of the limb seam, where the metallic fragments are angulated and very close together.

And this typically would lead to occlusion of thrombosis of a graft limb, and or in some cases partly thrombosis, following by embolization, leading to trash foot and other complications.

There have been attempts made to define this more scientifically if you like; buckling and any increase in angulation, kinking and acute localized angulation, and this comes from a publication of Peter Harris, the Journal of Endovascular Surgery.

Importantly, they showed that late complications, that is, endoleak or lumen collusion, developed in 19 percent of the cases, all of whom exhibited graft buckling or kinking in a series of Vanguard grafts.

So it was a very important association with those early signs, with late adverse effects on the patients.  Here is another example of graft buckling or kinking from our own practice, associated with displacement of the graft from the distal or the approximal neck of the aneurysm.

And you can see that it slipped down 5 or 6 centimeters from the neck.  There is still good blood from the channel, and here is the plain x-ray of that patient showing you the angulation of the metallic components of the graft.

We have also shown and published a variety of these grafts shortened considerably during the implant procedure, and with both the Vanguard and the aneurylextial (phonetic) stent, we observed a shortening of more than 10 percent, and close to 50 percent of case, probably caused by the forces needed to push the graft out of the deployment system so that the graft in the loader is this long, and in the patient is this long.

And that is another possible cause for acute or light device failure.  This slide was meant to demonstrate the fracture of a Y-Form, and I think it illustrates why it is an important point about Y-Form fractures, and those can be very difficult to detect.

And in this case, we see that a tubular graft has dislocated from the distal aspect of the makeup of the so-called distal makeup on an aneurysm.  This is one of our early cases from about 1993, and I will come back to this later on the computer slides just to discuss some of the complexities of diagnosis and analysis.

The other thing that is important to stress on these is variations in the angle between the various wires, which we feel may be associated with external pressure, but it may also be associated with variations in the compliance or the oscillations, or whatever you like to term it.

This slide was just meant to show that analysis and searching for these Y-Form fractures can be extremely difficult.  Here is an electro x-ray of a patient with one of the Baxter endografts in place, and you can see the individual Y-Forms.

And to look at these is a very painstaking effort to determine whether or not these are actually fractured or not.  So a lot of these are probably being missed on diagnosis today for all graft types.

And it is far more difficult to see the ones made of nyanom (phonetic).  In some cases, poor apposition of the graft, either because it has been placed too low, or because of thrombus, may lead to the complication which we now term endo-tension.

And that is defined as continued or recurrent pressurization of the aneurysm sac without any evidence of endoleak.  And here you see a patient whose aneurysm was slightly more than five centimeters at the time of implant.

And two years later the aneurysm is now close to seven centimeters in diameter, with no evidence of endoleak on any of the imaging studies that have been done.

So persistent or reoccurring pressurization without endoleak, and here is that same patient with an open surgical conversion, and these are the typical appearances we found now in seven cases.

And with liquified thrombus, or semi-liquid, or plasma-jelled type thrombus, surrounding the graft, and this liquified component, and in most cases we feel that this is due to an endoposed layer of thrombus at the neck, and this thrombus, which you can see here after the graft being removed, probably transmits pressure in males who have osmotic effects leading to activity within the thrombus and within the sac.

Whether or not this can be tested in a pre-clinical means is very doubtful.  However, the factors that lead to it may be identified, and perhaps can be tested.

We have heard a lot about variations in anatomy, and I would just like to finish off the slides just by talking a little bit about causal factors in variations in anatomy.

Somebody asked earlier about the relevant incidents of the taper or reverse taper in a series of patients that we analyzed last year, and this was in 8 percent of the cases that we actually treated with an endograft, and probably far higher than in those that are not treated.

We have defined six different types of neck, and we have shown that in almost every case it can be fitted with this classification.  And we have used that to define what we feel is a practical anatomic risk grading.

And in essence this depends on the neck and the iliac arteries, or the sac, with all favorable necks being defined -- correction, with all favorable necks being defined as Grade I, and unfavorable necks as Grade II, and severely unfavorable as Grade III.

Whereas, if the iliacs or the sac are unfavorable, that led to it being changed from 1A to a 1B.  And another way to put this, Grade IA, the neck was favorable, and the iliacs and sac were favorable.

And Grade 1B, the neck favorable, and iliacs and the sac unfavorable; and we come down to Grade III, the neck was severely unfavorable, and the iliacs and the sac were unfavorable.

Now,I mentioned this because I think this might have relevance to pre-clinical testing, because I believe that we should have difficult grades of difficulty in the models that we use for our pre-clinical testing.

We should be testing grafts in the Type 1-A anatomy as most of us have seen today, but we also should have different degrees of difficulty within the anatomic testing to simulate the human variation that we see.

Now, these ones would not fit the criteria for any graft at present, and you can see in Grade III the incidents of earlier outcome failure; that is, endoleak was 50 percent.  But in fact all of these were sealed and were cured by subsequent removals.

And we are now awaiting the long term results of studies such as this.  I would quickly now just kind of like cross to the other screen, because I would like to try and bring together what I have said from the clinical situation, to how it pertains to this meeting today.  If we could have the next slide.

In an outline, I am just going to reiterate the frequent modes of endograft failure, and state which I think are the most critical modes of failure, and thereby come to some form of classification.

And talk a little bit about causing factors, and emphasize the relationship of device failure to patient, and any other factors, and I will try to do all of that in about three minutes.  Next slide, please.

The classification of failure may be according to many factors, and we can classify according to the component or device affected, and we can classify according to the cause, and we can differentiate device factors versus patient factors.

We can talk about the timing, early versus late, the severity, and particularly with respect to clinical effects, and there are probably many other ways of classifying device failure.

To date, most testing has really been relevant to acute device failure.  That is, being able to deploy the graft and analyze acute failure within the model, usually with fairly simple anatomy.

And the things that are tested there, first, whether or not the device will deploy, which we are not ready to discuss today; and failure of the device to attach to the aortic wall; and failure of the device to seal or exclude the aneurysm; disconnection of modular components; and occlusion or thrombosis of the graft.

So I think that this represents a useful framework for a subsequent discussion.  Next slide.  If we now look at the chronic or long term modes of device failure, these are the ones which we really have difficulty in devising preclinical testing for, because it really pertains to device durability.

But these same factors really come up here.  The failure of the device to maintain its fixation to the aortic wall, and leading to displacement, loss of device seal, or it may have been previously sealed, disconnection of modular components, and again occlusion, or thrombosis of the graft again.

And now device materials failure or durability rears its ugly head.  Next slide.  And when we are looking particularly at those device and materials durability factors, we are coming back to most of the things that we have already talked about today.

Fractures of the stents, or Y-Forms, and which hasn't really been mentioned too much, either due to fatigue or corrosion, or to most likely a combination of these.

Breakage or disconnections of sutures, which I didn't illustrate on those slides, but it has also been a frequent failure mode.  Graft fabric or material tears, metal erosion through the fabric, and device angulation or kinking.  Next slide.

The critical effects of device failure are really those that affect the patient in the most critical ways, and causes, such as failure of the seal, or fabric tears, or device shift, or stent fractures, are really only relevant to the patient if they cause either an aneurysm enlargement, or aneurysm rupture.

Likewise, device kinking, buckling, and so on, is also only relevant to the patient if it is a precursor to these factors, or if it causes arterial blockage, or embolus.

So in looking at the critical effects, we are really talking about the effects on the patient rather than the bench top testing.  But naturally we need to devise bench top testing which can lead to earlier recognition of these potential failure modes.  Next slide, please.

I would like to just quickly give a patient illustration, which I also used at the recent conference in Ireland, because I think it really shows the multifactorial nature very nicely.

This is a true case, and it involves a 73 year old male, who we initially treated in 1993 for an abdominal aortic aneurysm.  At that time, we only had two grafts available in a workable fashion, and so he received the tube endograft.

And not surprisingly, two years later, he presented with distal endoleak, and the graft had pulled out of the distal end, and it was similar to that slide that I showed you where the graft had pulled out distally.

And we put a second redistal tube extension in to seal it off.  Two years again later, he had a Type I distal endoleak, which occurred during an exercise test for his cardiac status, and this was treated by a third endograft.  Next slide.

It may on the face of it seem very simple on what caused failure in this case, but here are the factors that really need to be analyzed in determining that, and I think that is part of the problem that we are struggling with here, because it is a very multifactorial question.

The patient selection and device type obviously paid a role, because we think that distal disconnection may not have occurred if we used bifurcated graft rather than a tube graft.

However, they weren't available.  Changes in the patient's anatomy may have contributed because the distal nick may have expanded, leading to that disconnection.

Activity and reactions in the  aneurysm sac seemed to be an important component, because the patient actually developed a coagolapathy at the time of his second endoleak.

Endoleak forces, and the effect of exercise and extraneous factors, and perhaps there was also a device material failure, and interaction of the metal, and the fabric, and the graft.  Next slide.

The patient factors, I think, we have done to death already, and so I won't dwell on those.  Next slide.  Device materials failure.  Again, all of these are the things that can lead to long term failure in the durability of devices.  Next slide.

And to finish off, there have been some attempts within the various bodies to try and define these failure modes in a mode that is relative to reporting.

One is to classify endograft deformity, and you see an initial attempt here, which is due to be published in the SVS standards document hopefully later this year in the Journal of Vascular Surgery, with a grading of angulation or kinking of the graft from zero to three.

And on the final slide, a further classification of device failure according to clinical severity, which from the patient's point of view is obviously the most relevant.

And with Grade Zero being no adverse or clinical effects, and down to Grade Three, resulting in conversion, rupture, major complications, or death of the patient.

So I hope that will serve as a useful introduction to this section on modes of failure that have been observed in practical and clinical experience to date.  Thank you.