Chapter	7

APPLIED RADIATION MEASUREMENTS

John M. Keller

Oak Ridge National Laboratory

INTRODUCTION

The four major radiation measurements systems found in most radioanalytical chemistry laboratories include gas-flow proportional counters for alpha and beta counting, gamma spectrometry systems, alpha spectrometry systems, and liquid scintillation counters. The practical application of each of these counting systems will be discussed in detail. Specialized counting equipment for the measurement of neutrons and x-rays will be introduced along some non-electronic techniques used for radiation measurements.

DETECTOR SELECTION

The types of radiation detection systems employed by radioanalytical chemistry laboratories have changed considerably over the last sixty years and there has been significant improvements observed since the early 80's.  Advancements in the areas of material science, electronics, and computer technology have resulted in more sensitive, reliable, and user friendly laboratory instruments.  The main discussions in this section will focus on four primary radiation measurement systems considered to be absolutely necessary for any modern counting laboratory.  These basic components of a counting laboratory include gas-flow proportional counters, liquid scintillation counters (LSC), gamma spectrometry systems and alpha spectrometry systems.  These four counting systems provide the tools to measure and identify most forms of radiation from alpha and beta decay, including both charged particles and gamma photons.

MEASUREMENT OF CHARGED PARTICLES

The most useful charged particles for radiochemical measurements include the following particles emitted from the nucleus; negatrons (negative electrons or - particles), positrons (positive electrons or + particles), and alpha particles which consist of two protons and two neutrons or the nucleus of a 4He atom.  There are other sources of charged particles that an analyst needs to be aware of because of the potential impact on radiochemical measurements. For example, charged particles such as conversion and Auger electrons would be difficult to distinguish from - particles with most beta counting systems and without proper correction could bias final activity measurements

An important distinction between electrons emitted from the nucleus and alpha particles is the energy distribution of the charged particles emitted during the decay process.  The energy from both - and + particles are distributed over an energy continuum ranging from 0 KeV to some maximum energy, usually referred to as Emax.  In theory, the value for Emax from a - or  energy spectrum could be used to identify a radionuclide, however, in practice it is difficult to get an accurate enough measurement of Emax to be used for conclusive identification. In contrast, the alpha particles are emitted at discrete energies that are characteristic of the parent radionuclide and are much easier to use in practice for radionuclide identification.

GAS-FLOW PROPORTIONAL COUNTERS

There are many different types of radiation detectors that use a counting gas but all gas-filled or gas-flow counters are based upon a similar design.  All of these detector systems apply a high voltage across a volume of gas in some type of containment system.  The gas containment systems for these types of counters are available as either sealed detectors or can be designed to allow continuous flow of the gas.  When the high voltage is applied across the gas an electrical field is formed between two electrodes.  Then when radiation passes through this electrical field with enough energy (~30-35 eV) to ionize the counting gas a flow of current is created between the electrodes.  Based upon detector design, this electrical signal can be measured as either a current, voltage pulse, or accumulated charge which can be related back to the incident radiation.

There are three basic types of gas-filled detectors

| Detector Type | Applied Potential (Volts) | Characteristics |
| --- | --- | --- |
| Ionization Chamber | ~250-500 | no gas amplification pulse height ( energy useful for , , and  simple, low cost e.g., pocket dosimeter |
| Proportional Counter | ~500-750 | gas amplification, 103 - 104 pulse height ( energy useful for  and  relatively simple, higher cost e.g., laboratory instruments |
| Geiger-Müller Counter | ~1000-1500 | gas amplification, no energy discrimination useful for , , and  simple, low cost e.g., survey meters |

Although, almost any gas will work with an ionization chamber, the counting gas for a proportional counter must have specific characteristics for proper operation.  The most commonly used gas for a proportional counter is a mixture of 90% Argon with 10% methane.  This counting gas mixture is referred to as P-10 counting gas and is readily available at a reasonable cost.  For a gas-flow Geiger-Müller or G-M counter another standard counting gas used is called Q-gas, which is a mixture of 97% helium and 3% isobutene.

The objective of this chapter is to concentrate on the application of laboratory counting systems and although there are some useful applications for ionization chambers and Geiger-Müller counters in the laboratory, the real laboratory workhorse is the proportional counter; specifically the gas-flow proportional counter.  There are several types of gas-flow proportional counters frequently used in modern counting laboratories.  The first is a classical 2 gas-flow proportional counter that is usually available with a thin Mylar window or as a windowless detector.  These 2 counters are typically shielded with a few inches of lead and may use guard detectors to reduce background from cosmic radiation.  The next type of system usually available is similar to the 2 counter but also includes an automatic sample changer.  This type of stack counter is normally used to count samples that do not require a long counting time (i.e., 30 min. to a few hours) and are good for counting plates, filters and smears up to about two inches in diameter.  A third type, which is more expensive and may only be found in larger laboratories, uses multiple detectors to simultaneously count multiple samples.   The standard configurations available commercially have a modular design with groups of four detectors arranged in a slide out drawer.  These multiple detector counting systems can be obtained with as few as four detectors and some can be expanded up to 64 detectors, however, due to the high cost most systems are sold with 8–16 detectors.  This type of counting system is very useful for low-level samples that require a long count time (i.e., 24-48 hours per sample), however, the laboratory needs a fairly high sample load to justify the high instrument cost.

The selection of an operating voltage is critical for proper operation of a gas-flow proportional counter.  A gas-flow proportional counter can be operated at a bias voltage (~600-800V with P-10 gas) where only alpha particles are detected, which is commonly referred to as “alpha only” mode.  If the bias voltage is increased to ~1500 V with P-10 gas the counter will respond to both alpha and beta particles.  Operation at this higher voltage is referred to as “simultaneous” or “+” mode. The optimal operating voltage for each mode is determined from the alpha/beta plateau curves which are plots of counting rate versus bias voltage with an alpha or beta source.  Although an alpha/beta plateau curve can be generated with a mixed source containing both alpha and beta activity, there are advantages to generating two sets of data with both a “pure” alpha emitter (238Pu, 244Cm, etc.) and a “pure” beta emitter (i.e., 63Ni, 90Sr, 99Tc, etc.).  When the two sets of data are plotted on the same graph the analyst can easily choose a bias voltage for “alpha only” mode that is below the point where beta particles just start to be detected.  Selection of the bias voltage for counting alpha particles only by this technique will minimize beta cross talk.  Selection of the bias voltage for + counting is not as critical and should be about the midpoint of the + plateau.

Once the bias voltages are selected the next step is to determine the counting efficiency for alpha and beta measurements.  The alpha counting efficiency with a proportional counter is essentially constant over the energy range of 4-6 MeV which includes most of the alpha emitting radionuclides of interest to most radiochemistry laboratories. (Note: This energy range includes all the naturally occurring alpha emitting isotopes of radium, thorium, and uranium, along with the plutonium and americium present environmentally from the global fallout of nuclear weapon testing.  There are also a few naturally occurring rare earth elements (e.g., 147Sm, 144Nd, 152Gd, etc.) that emit alpha particles in the 2-3 MeV range, however, these radionuclides have very long half-lives and are not generally of much interest to most radiochemistry laboratories.) Since the alpha particle will deposit most of its energy within the active volume of a proportional counter the alpha efficiency is a function of the source-detector geometry minus any losses due to mass absorption from the detector windows.  The typical counting efficiency, which is the ratio of count rate to disintegration rate (cpm/dpm), for alpha counting with a windowless 2 gas-flow proportional counter is ~51.5%. The additional 1.5% over the expected 50% from the 2 geometry is due to backscattering of the alpha particle from the source mounting material, which for many laboratories is simply a stainless steel plate.  The alpha efficiency for a similar proportional counter with a thin Mylar™ window would be in the range of 30-40% due to mass absorption by the window material.

Beta particles are more penetrating than alpha particles and are less affected by mass absorption.  Beta particles may also escape the active volume of a proportional counter before all the energy from the  particle is deposited within the detector.  In addition, there is a higher probability of backscatter with beta particles which, depending upon the source backing material, can have a significant impact on the overall counting response with a proportional counter.  Due to both the energy distribution of beta emissions and the interactions of the beta particles with the counting gas, the beta counting efficiency for a proportional counter can be fairly complex.  In general, the beta counting efficiency is more of a function of the average beta energy and can range from 50-90% with a windowless gas-flow proportional counter.

Since the alpha efficiency is fairly constant over the range of 4-6 MeV, the activity of any alpha emitting radionuclide in this energy range is frequently calculated with an estimate of the counting efficiency determined with 241Am (5.50 MeV).  In other words, radionuclides ranging from 238U (4.02 MeV) to 244Cm (5.80 MeV) would all have about the same alpha counting efficiency with a gas-flow proportional counter.  However, in contrast to the alpha emitting radionuclides, the beta counting efficiency with a gas-flow proportional counter is significantly different for each beta emitting radionuclide.  Therefore, to quantify the activity for 63Ni (66.9 KeV) or 99Tc (292 KeV) would require a separate standard of each radionuclide be available for the determination of the respective counting efficiencies.

Two measurements routinely performed by radiochemistry laboratories on a gas-flow proportional counter are gross alpha and gross beta.  These measurements are performed on a sample solution after the initial sample preparation but before any chemical separations. The gross alpha and beta results are routinely reported in units of activity with requires a known counting efficiency.  This is not a problem for gross alpha measurements since the alpha counting efficiency is fairly constant over a fairly wide energy range, however, as discussed previously the beta counting efficiency is a function of the average beta energy so there is no single beta counting efficiency to use for the calculation of the gross beta activity.  In practice, the beta efficiency for gross beta measurements is commonly determined with a source prepared from 90Sr (546 KeV) in secular equilibrium with 90Y (2281 KeV).  This method only provides a counter efficiency for the 90Sr/90Y pair, so any user of the gross beta activity data needs to understand these limitations.

These gross activity measurements, along with the gamma spectrum from the same sample preparation, can give the radiochemist a good starting point to determine which additional analyses may be necessary.  Also, this information helps determine the aliquots needed for any required radiochemical separations.  These gross activities can also be used as a data quality check to determine if all significant alpha and beta emitters have been identified in the sample.

Samples for gross alpha and beta measurements frequently involve dissolving solid materials such as a soil or the sample may simply be an acidified ground water. In either case, there can be a significant amount of dissolved solids in the sample solution.  Also, some radiochemical separation methods add carriers or other reagents that may increase the dissolved solids within a sample solution.  When these sample solutions are evaporated onto a counting planchet there is a buildup of solid residue of the surface of the plate that can reduce the count rate due to self-absorption or mass-absorption.  Self-absorption is a source of counting error that must be minimized or accounted for with suitable correction curves.  A suggested upper limit for sample residue weight on a counting plate or planchet is about 5 mg/cm2 for alpha counting and about 10 mg/cm2 for beta counting.  For example, an upper limit for sample loading on a standard 2-inch (~20 cm2) counting planchet for alpha and beta counting would be 100 mg and 200 mg, respectively.  If the residue loading is above these limits the sample should be diluted and a new planchet prepared.  For sample loading below these upper limits the data can be corrected for the reduced count rate from the effects of self-absorption.  Therefore, each gas-flow proportional counter should have a set of self-absorption curves showing the sample loading or residue (mg) vs. the counting efficiency (cpm/dpm) prepared with both standard alpha and beta solutions.  These self-absorption curves can be used to extrapolate the correct counting efficiency up to the recommended upper limit for sample loading.

SOURCE PREPARATION

There are many techniques for preparing a counting source suitable for alpha and beta measurements and the analyst needs to understand the advantages, disadvantages, and limitations of these different approaches.  These techniques range from very simple evaporation methods to more involved procedures using electroplating or vacuum sublimation.  An ideal source for counting alpha or beta should have the following characteristics:

- Low solids content to minimize mass absorption problems, the material to be counted should be evenly distributed over the source backing material (i.e., thin stainless steel disk) in a very thin layer of low mass.

- The alpha or beta activity needs to be firmly fixed to the counting source to avoid contamination of the counting system and to ensure the general safety of the analyst.  Some laboratories cover or seal the source with a thin material such as Mylar™ to provide an additional layer of safety, especially when working with alpha emitters.

- For beta counting the source backing material should be a thin low Z material to reduce the effect of backscatter.

- For some counting systems (especially windowless gas-flow proportional counters) the source backing material needs to be conductive to avoid polarizing the source and creating a static charge within the counting gas volume.

LIQUID SCINTILLATION COUNTERS

Liquid scintillation counting (LSC) has been around for over fifty years and is probably the most widely used technique for the measurement of radioactivity.  Traditional LSC applications in radiochemistry laboratories are minor compared to the wide use LSC techniques in the life sciences (biological and medical sciences) for the measurement of labeled proteins,  genetic materials (RNA and DNA), and other types of biological tracer work.  Liquid scintillation analysis (LSA) techniques are also very useful for the analysis of smear samples collected for health physics or radiation protection applications.  Although liquid scintillation counting can be used for the measurement of radioisotopes in any liquid sample, LSC is the preferred method for both low energy beta emitters (i.e., 3H, Emax= 18.6 KeV) and volatile radionuclides (e.g., 99Tc as TcO42-) that could be easily lost preparing a planchet for a gas-flow proportional counter.  Some of the radionuclides routinely measured by liquid scintillation counting are listed in the following table.

| Radionuclide | Emax (KeV) | T1/2 | Volatile Forms |
| --- | --- | --- | --- |
| 3H | 18.6 | 12.3 y | H2O, H2 |
| 14C | 156.5 | 5730 y | CO2 |
| 32P | 1709.0 | 14.3 d |  |
| 35S | 167.4 | 87.2 d | SO3, H2S |
| 36Cl | 709.0 | 301000 y | Cl2, HCl |
| 45Ca | 258.0 | 163 d |  |
| 63Ni | 66.9 | 100 y | NiCl2, Ni(NO3)2 |
| 99Tc | 292.0 | 213000 y | TcO42- |

The counting efficiency for charged particles with a liquid scintillation counter can approach 4 geometry.  The LSC efficiency for alpha particles is typically > 95%, whereas the counting efficiency for beta particles is generally > 90% for betas with an Emax > 250 KeV.

Sample preparation for Liquid Scintillation Counting

To achieve good analytical results on a (pure( beta emitter it is usually necessary to chemically separate the radionuclide of interest from all other radioactive species prior to analysis.

Distillation, wet oxidation, combustion (dry oxidation)

(erenkov counting

The emission of (erenkov radiation occurs when a charged particle enters a media (i.e., water) at a speed that exceeds the velocity of light in the same media.  The energy lost as the velocity of the charged particle slows to the speed of light is emitted as (erenkov radiation, which at high enough activity levels can be visually observed as a bluish-white light.  This blueish-white light can be used to measure high energy beta particles with a liquid scintillation counter.  The production of (erenkov radiation is a function of the refractive index of the medium.  For water, which has a refractive index of 1.33, the threshold beta energy to produce (erenkov radiation is ~263 KeV.  Therefore, the beta emitting radionuclides needs to have an Emax well above the threshold of 263 KeV for (erenkov counting to be practical.

ALPHA SPECTROMETRY

GAMMA SPECTROMETRY

It was discovered early in the history of radiation measurements that certain materials would give off a flash of light when exposed to radiation. It was also recognized early on that some materials were more sensitive to different forms of radiation.  As technology improved these scintillation materials were coupled to photomultiplier tubes to measure the flashes of light. Inorganic scintillation detectors based on alkali halides were used in most of the early gamma spectrometry systems.  The most common scintillation crystals used in these systems were sodium iodide doped with thallium and were referred to as NaI(Tl) detectors.

Lithium drifted semiconductor detectors were the next generation of detectors for gamma spectrometry.  These detectors were based on both germanium and silicon and were designated as Ge(Li) and Si(Li), respectively.  The lithium drifted detectors provided higher resolution but at a much higher cost for an efficiency similar to the NaI(Tl) detectors.  Historically the Ge(Li) detectors were much more useful for gamma spectrometry and the Si(Li) detectors were better for low energy gamma-rays and x-rays.  The Si(Li) detectors continue to be the best choice for x-ray spectrometry in modern counting laboratories.

In addition to cost, the three important performance characteristics of a semiconductor detector for gamma spectrometry include the efficiency, resolution, and peak-to-Compton ratio.  These detector characteristics should be optimized for each particular application.  For example, a NaI(Tl) detectors have excellent efficiency but poor resolution and may be ideal for simple spectra or an online monitor system but are not very useful for applications involving complex spectra.

The formal definition for resolution is E/E, or the peak width divided by the peak centroid energy.  The resolution is typically reported as a percentage for NaI(Tl) detectors and as the full-width-half-maximum (FWHM) for semiconductor detectors.  The FWHM is the width of the peak in units of energy measured at ½ of the peak height.  The resolution is very important for the interpretation of complex spectra and radionuclide identification.  The FWHM for a full energy peak is directly proportional to the peak energy.

A standard method as described in the IEEE Standard Test Procedures (ANSI/IEEE 325-1996) was developed to define the relative detection efficiencies for coaxial detectors which is suppose to permit the comparison of detectors for different sources.  The relative efficiency for a coaxial detector is defined as follows:

where,

HPGe 	=	absolute efficiency at 1332 KeV for a coaxial detector measured 25 cm from the source.

NaI(Tl) 	=	absolute efficiency at 1332 KeV for a 3" x 3" NaI(Tl) scintillation detector measured 25 cm from the source.

The absolute efficiency is the ratio of the number of counts measured in the 1332 KeV photopeak divided by the number of gamma rays at 1332 KeV emitted from the 60Co source during the same time interval.

A comparison of resolution and efficiency

| Energy (KeV) | NaI(Tl) (30mm x 30mm) | HPGe (~30% Relative Eff.) |
| --- | --- | --- |
|  | abs (%) | FWHM (KeV) | abs (%) | FWHM (KeV) |
| 60 | (241Am) | 50 | ~8 (14%) | 22.4 | 1.2 |
| 662 | (137Cs) | 32 | ~50 (7.5%) | 5.0 | 1.5 |
| 1332 | (60Co) | 25 | ~95 (7.1%) | 2.5 | 1.9 |

The efficiencies listed in this table represent nominal values that would be expected for a point source at contact with the detector, which is basically a 2 geometry.

The third performance characteristic that needs to be considered before selecting a semiconductor detector for gamma spectrometry is the Peak-to-Compton ratio.  The Peak-to-Compton ratio is an important indication of the detectors performance at lower energies, this is especially important in the presence of higher energy gamma radiation such as 60Co (1332 KeV) where the Compton scattering become more predominate within the detector volume.  The resulting Compton continuum starts at ~256 KeV below the photopeak and increases the spectral background across the entire low energy region below the Compton edge.  For 60Co the Compton edge is observed at ~1076 KeV (e.g., 1332-256).  Most detector manufactures report the Peak-to-Compton ratio using the 60Co photopeak at 1332 KeV, which would be the count rate measured at 1332 KeV divided by the average count rate at 1076 KeV.  This ratio typically ranges from 40 to 70 for coaxial HPGe detectors currently produced.  In general, for detectors with a similar relative efficiency, the Peak-to Compton ratio is higher for detectors with better resolution.

The most common type of detector currently used for gamma spectrometry is the coaxial high purity germanium (HPGe) detector.  The HPGe detectors are usually available with either a p-type or n-type crystal.  The efficiency of a p-type detector typically drops off quickly below ~100 KeV, whereas, the n-type detectors can actually be used down to about 3 KeV with a beryllium window.  The best choice for these types of detectors depends on the application.

CHARACTERISTICS OF A GAMMA SPECTRUM

There are three ways a gamma photon can interact with matter and each type of interaction must be understood to properly interpret gamma spectral data with either a NaI(Tl) or HPGe detector.  The three modes of interaction are photoelectric absorption, Compton scattering, and pair production.  The probability or cross-section for each of these interactions is a function of the energy for the incident gamma photon.  For gamma spectrometry the photoelectric absorption is the preferred mode of interaction since it results in a full-energy photopeak in the spectrum that is used to quantify and identify the radionuclide.  If the photon interacts with the detector material by Compton scattering there is a loss of energy by the incident photon that is transferred to a recoil electron.  Since this energy loss is a function of the scatter angle, and all angles are possible, this scattering interaction results in the Compton continuum that is observed in the gamma spectrum as an increase in the spectral background.  The third mode of interaction is pair production which can only occur with incident gamma photons that have energy greater than 1022 KeV.  The process of pair production absorbs all the incident gamma photon energy and results in the creation of an electron-positron pair within the active volume of the detector.  Then as the positron loses energy it will combine with an electron which results in the annihilation of the charged particles and the release of two annihilation photons that have an energy of 511 KeV each.  Depending upon the detector size it is possible for either one or both of the 511 KeV annihilation photons to escape from the active volume of the detector.   If both of the annihilation photons escape the active detector volume a “double escape peak” would be observed in the gamma spectrum that would be located 1022 KeV below the full-energy peak.  It is also possible to have a “single escape peak” located 511 KeV below the full energy peak if only one of the annihilation photons escape the detector volume.

To ensure a low background detectors used for gamma spectrometry are typically surrounded with at least lead shielding.  Also, the crystals used in both NaI(Tl) and HPGe detectors are encased in an aluminum shell.  These various materials around the detector can interact with the radiation from a sample source and cause additional background structure in the gamma spectrum.  The same interactions that occur within the detector are also experienced with these materials surrounding the detector; however, these interactions external to the detector have a different impact on the spectral background.  For example, Compton scattering within the lead shielding will result in a broad backscatter peak that is located in the lower energy region (< ~250 KeV) of the spectrum.  The location of the peak maximum (EBS) for this backscatter peak, which is observed at a backscatter angle of , can be estimated from the following equation:

Where, E is the energy in KeV for the incident gamma photon.  For instance, the backscatter peak for 137Cs (662 KeV) would be ~184 KeV.

For a gamma photons with an energy greater than 1022 KeV, the photon interaction with the lead shielding can result in pair production.  The annihilation of the positron produced during this pair production process can result in a 511 KeV gamma that escapes from the lead shielding and ends up being superimposed on the gamma spectrum.  Another source of pair production within the lead shielding is cosmic radiation.  The sample source counted for the gamma spectrum can also contribute to the 511 KeV gamma peak if there are any positron emitters present in the sample. This source of 511 KeV gamma-rays results for the absorption of the positron, and subsequent annihilation, by the aluminum casing surrounding the detector crystal.

The photoelectric absorption of radiation by the lead shielding and other material surrounding the detector can result in characteristic x-rays in the lower energy region of the gamma spectrum.  For example, the lead K and K x-rays are almost always present as part of the background in a gamma spectrum.  There are commercially available gamma detector shields that are lined with cadmium and copper to reduce the level of lead x-rays in the background.  In addition to the lead x-rays, characteristic x-ray from any material located near the source and detector can end up in the background spectrum.

Environmental and other low level samples may require very long counting times to satisfy the low detection limits needed for regulatory applications and other special projects.  The background spectra associated with these long counting periods will show additional structure from the natural background radiation surrounding the gamma spectrometer system. Naturally radioactive trace elements present as impurities (i.e., uranium decay daughters, 40K, etc.) in the shielding material may contribute to the background and can become significant with long counting periods.  One final source of background radiation that must be considered with long counting periods are prompt -rays from the thermal neutron capture of cosmic ray neutrons within the detector and shielding materials.  A excellent indicator for this type of background problem is cadmium which is frequently present in graded shields (Cd-Cu lined lead) for gamma spectrometry systems.  One of the naturally occurring isotopes of cadmium (113Cd) has an extremely high cross section (~20,000 barns) for thermal neutron capture and will yield a 558 KeV prompt -ray.  There can also be prompt -rays from the lead and even the germanium crystal, but well below the level experienced with cadmium.  In general, any background structure due to prompt -rays should be significantly lower than the 511 KeV annihilation peak and the 1460 KeV from 40K.

X-RAY SPECTROMETRY

solid state detectors (energy dispersive)

crystal spectrometers (wavelength dispersive)

MEASUREMENT OF NEUTRONS

Thermal neutrons

10B(n,a)7Li, 3He(n,p)3H, 6Li(n,a)3H

Fast Neutrons

VARIOUS FACTORS THAT MAY IMPACT RESULTS

Background, Scatter

Mass Absorption

Interference

NUCLIDE IDENTIFICATION

Gamma emitters

Alpha emitters

(Pure( Beta emitters

REFERENCES

Evans, R. D., The Atomic Nucleus, New York: McGraw-Hill, 1958.

Horrocks, D. L., Applications of Liquid Scintillation Counting, New York: Academic, 1974.

Friedlander, G., Kennedy, J. W., Macias, E. S., Miller, J. M., Nuclear and Radiochemistry, 3rd ed. New York: Wiley-Interscience, 1981.

Tsoulfanidis, N., Measurement and Detection of Radiation, New York: McGraw-Hill, 1983.

Knoll, G. F. Radiation Detection and Measurement, 2nd ed. New York: Wiley, 1989.

Ehmann, W. D. and Vance, D. E., Radiochemistry and Nuclear Methods of Analysis, New York: Wiley-Interscience, 1991.