# Sterile Filtration:  Process Optimization Through Real Compliance

by Hans G. Schroeder, Ph.D.

International Consultants Association

Encinitas, CA

As a general rule, the regulations imposed on the pharmaceutical industry offer clear advantages to the patients who depend on their products.  Enforcement of the regulations and associated guidelines provides assurance to the consumer that approved and acceptable quality levels are met, and that similar products from different manufacturers are, in essence, equivalent in safety and efficacy.  The main benefit of compliance to the industry is the corresponding license to stay in business with a minimum of citations or regulatory actions by the enforcing agency.

Over time, industry has learned to stay in compliance by following good manufacturing practice regulations to the letter.  It is, in effect, impossible to run a pharmaceutical operation without a sound set of written standard operating procedures and programs that assure important contributions to the expected product quality, such as operator training, internal audits, validation of critical processes,  and change control.  Compliance is obviously a vital component of our industry, so important that an entire journal is dedicated to this topic (1).  Many articles reflect the fact that current regulations are subject to changes as well as interpretation, and suggest their readers update their compliance programs by reviewing current FDA 483 observations, Warning Letters and guidance provided by publications like the Pink Sheet and the GMP Monitor (2,3).  While such “Regulatory Compliance” may keep inspectors happy, it unfortunately also leads to constant and costly knee-jerk reactions to avoid similar observations in the next inspection of their own facility.

A viable, and, on the long run more cost effective, alternative was recently presented by Madsen (3).  He proposes that a state of process control well above regulatory compliance can be achieved by applying fundamental scientific principles and technically sound operating, control and documentation procedures.  Such an approach must, of course, keep current regulations in mind, and, if properly implemented, regulatory compliance becomes, for practical purposes, a beneficial side effect such “Real Compliance.”

According to Madsen, key elements required to achieve real compliance include an adequate number of trained operators, a technically sound set of instructions and operating procedures, dependable redundancy, appropriate verification and documentation and an immune system that corrects process variations and discrepancies before they result in defective product.  Once a real compliance state of control is reached, the process can be monitored by meaningful feedback loops, including warning indicators and alarms, but in-process checks become confirmatory rather than controlling.  Unfortunately, the direct outcome of some processes cannot be monitored on line and in real time.  Good examples are sterilization processes, for which the incubation time associated with commonly practiced microbial monitoring represents an unavoidable delay.  In addition, many quality attribute tests, sterility included, are destructive in nature.  Thus, particularly in such cases, a real state of control depends heavily on sound characterization and understanding of the process, which, in turn, makes the success of such processes most dependent on a sound validation program.

# Process Validation

Many approaches to validation have been presented in the technical literature over the years, but, once the glitter is removed, a sound process validation program can be reduced to responsibly pondering, substantiating and documenting the following four questions:

- What is the desired outcome or what is the objective of this process step?

- Which of the process variables will affect the successful outcome of the step?

- How are these process variables monitored and controlled?

- Will the control limits imposed assure that the objective is met consistently?

By analyzing these questions it is possible to determine if a process is in real compliance, even if no regulatory citations are issued against the manner in which the process is currently run.  Sterile filtration will be discussed to illustrate this approach.

# Objective of the Sterile Filtration Process

By implication, the obvious objective of a sterile filtration step is the removal of any viable organisms that may be present in the bulk product solution as manufactured.  This is typically accomplished by the use of “Sterilizing Grade” membrane filters, defined in the FDA “Aseptic Guideline” (4) as those capable of totally retaining a challenge level of 10^7 cfu / sq cm at a differential pressure of 30 psi.  The FDA currently accepts B. diminuta at this challenge level as a worst case model, but it is also acceptable to use natural flora.

# Analysis of Variables

The variables recognized to have an impact on the outcome of a sterile filtration exercise have been extensively discussed in the technical literature and were recently reviewed by Jornitz et al (5,6).  These variables include the physical-chemical nature of the carrier liquid, the hydrodynamic conditions and the duration of the process, qualitative and quantitative aspects of the bioburden, and last, but certainly not least, the list of variables must include the very filter chosen for the process.  Variables directly associated with the filter itself include the type, its membrane material and any surface treatment, the actual pore size (as opposed to an arbitrarily assigned pore size rating), pore morphology, thickness and tortuosity, and the degree of “integrity” of the filter, as determined by the actual size of the largest pores and defects, if any, present on the filter used.

Retention has long been recognized to depend on complex interactions between these variables.  The membrane and its surface treatment, in combination with the contaminant and the carrier liquid, typically define the retention mechanism, while the integrity determines the extent of the retention capability.  In recognition of the impact of such variables, the FDA Aseptic Guideline expects the validation to be run simulating the actual processing conditions, using the actual product to be filtered or a suitably similar surrogate as a carrier vehicle.  In a series of validation runs, the bioburden is also defined and kept relatively constant, and, of course, the filter type is fixed by the user’s choice according to whatever filter selection criteria were followed.  This reduces the long list of variables that have an impact on the retention to one, namely, the integrity of the filter.  Hence, the long-standing emphasis on establishing a sound correlation between a non-destructive integrity test and bacterial challenge tests as a basis for validating the retention capability of membrane filters.

Note that in an actual filtration process involving the filter type validated, most of the variables discussed above will become fixed parameters.  This is true even for the integrity of the filter, since a minimum acceptable integrity parameter will have to be met before it is placed into service.  On the other hand, the cell count and nature of the bioburden will likely be different in actual filtration settings, particularly if a model such as B. diminuta was used in the validation runs.  In light of recent publications that document the passage of presumably smaller organisms through B. diminuta retentive membranes (7,8,9), this could adversely affect the expected retention in an actual filtration setting.  For this reason, validations using naturally occurring flora will likely eventually be preferred over any model organism testing, even if natural flora may also vary from time to time.  To reach concentrations of 10^7 / sq cm using natural flora it may also be necessary to increase the inherent bioburden of the actual product by additional steps, such as allowing incubation beyond the regular processing time or the concentration of cells by cross-flow filtration.

# Monitoring the Integrity of Sterilizing Grade Membrane Filters

As has been extensively discussed in the technical literature, currently practiced membrane filter integrity tests are based on the observation that membrane samples wetted out with a suitable liquid are essentially impermeable to viscous flow of gases at low pressures.  As the applied pressure is increased, the liquid within the largest pores present in the sample will eventually start to dislodge, allowing the test gas to flow through them forming visually observable streams of bubbles.  The pressure at which bubbling commences is referred to as the bubble point of the membrane, a value which has been successfully used to characterize the integrity of small area membrane samples, such as the 47 or 90 mm discs often used for retention validation tests.

When passage is observed in a given set of retention validation runs it must be attributed to the larger pores, no matter what the retention mechanism is, since the actual retention mechanism is not a variable in the retention validation studies.  This can readily corroborated by subjecting membrane samples of successively higher bubble-point values to challenge studies using a defined contaminant such as B. diminuta.  Typically, a clear retention pattern will be observed.  At higher bubble-point values, the retention becomes more “reliable,” and a pressure, or a narrow range, above which “total” (better than the challenge level or 10^7 cfu / sq cm) retention is observed, can readily be identified.  This pressure is often referred to as the “critical” bubble-point value.

Since integrity is the only true variable that will affect the outcome of a given set of retention validation runs, the FDA Aseptic Guideline rightfully expects the retention capability of the filter to be correlated to a meaningful integrity test.  Even prior to this official request, technical support literature and the so-called Validation Guides published by the filter manufacturers have shown B. diminuta retention data as a function of the bubble-point pressure of their sterilizing grade membrane products.  Johnston (10) reviewed the retention data presented in various studies and concluded that for most types of membrane, a 10% reduction in the bubble-point will reduce the retention capability by 2 orders of magnitude.  This means that if a given type of membrane has a critical bubble-point of 50 psig, passage of up to the order of 100 cfu per square centimeter in a 10^7 cfu / sq cm challenge test of a 45 psig bubble-point sample could be observed.  If such a disc were to be built into an otherwise “good” cartridge, challenge of the cartridge to the same bacterial load would again result in passage of a similar number of organisms, since a chain can not be stronger than its weakest link.

The FDA Aseptic Guideline also requires that the integrity test be appropriately scaled up for the filter to be used in the production setting.  The integrity of sterilizing filter cartridges typically used in the production setting is often established through quantitative gas flow measurements, often referred to as diffusive flow tests, rather than the go/no-go visual bubble point approach.  If the flow of the test gas at a pre-established pressure exceeds a value deemed acceptable, the filter is considered defective and will be excluded from use.  While the approach as such is reasonably sound, it has been a long-standing tradition that filter cartridges are to be integrity tested at 80% of the critical bubble-point value.  Such a lower test pressure can obviously assure only that the cartridge is no less than 80% as tight as it really needs to be, allowing non-retentive areas, such as described above, to go undetected.

Poor wetting of a membrane filter element may result in failing to meet the integrity test acceptance criterion.  These so-called “wetting problems” can be distinguished from actual defects by re-testing the filter after extended flushing or re-wetting it with another liquid of lower surface tension, but it is not very practical to perform such additional tests in the production setting.  Thus, to circumvent the issue, some filter users opt not to perform a pre-filtration integrity test.  This practice is safe.  After all, the retention observed in validation runs is correlated to the integrity test results of clean filters, thus, to expect any retention assurance, the filter must be tested for integrity prior to becoming plugged with the contaminants it retains during its use.  This is particularly important since it has been reported that marginal and substandard membrane filters can plug up during use to the point that after the filtration an acceptable integrity test result is observed (11).  Since untested filters may not even have met an already compromised 80%-specification, a fair number of viable cells may well make it through the membrane during the early part of the filtration process.

Adequacy of Integrity Control

Traditions such as lowering the test pressure to improve cartridge yields or omitting pre-filtration integrity tests to circumvent possible wetting problems are not safe, because the filter may indeed not be any tighter than what corresponds to the pressure to which it was tested.  As discussed above, this action will unquestionably reduce the sterility assurance of the lot of product filtered.  Further, deliberately settling for an avoidable compromise in the sterility assurance level is in strong contradiction with the FDA “Validation” Guideline (12), which specifically states:  “Each step of the manufacturing process must be controlled to maximize the probability that the finished product meets all quality design specifications.”  Nevertheless, the FDA condones such practices, making them acceptable in terms of regulatory compliance.

The FDA “Validation Guideline” also specifically states that “In assessing the suitability of a given piece of equipment, it is usually insufficient to rely solely upon the representations of the equipment supplier,” presumably in order to increase the validity of validations and reduce the potential for conflict of interest.  Again, sterile filtration stands out as a regulatory exception, as a high percentage of filter users rely heavily on data generated on a contract basis by filter manufacturers, even product-specific validation data.  The FDA “Aseptic Guideline,” in recognition that “Some of the more complex filter validation tests may be beyond the capabilities of some filter users,” accepts this practice at the expense of complacency among filter users.  This exemption has kept the number of adequately trained and knowledgeable filter users at such a low level that filter manufacturers are often called upon to answer questions the FDA has about their client’s process.  The fact that some users may not be able to answer such questions on their own is a poor reflection of the level of control our industry has managed to acquire over such a critical process step.

The technical flaws in the currently accepted integrity test practices make it mathematically impossible to consistently expect the advertised retention level of better than 10^7 cfu / sq cm.  In fact, a few filter users have observed and reported discrepancies when attempting to reproduce manufacturers’ retention data on their own (13,14,15), and even the FDA has reported observing such discrepancies in validation and actual production filtration settings (16,17).  The possibility that occasionally a batch of filtered product may be lost for failing the sterility test on account of poor filter retention performance cannot be ruled out.  None-the-less, it has to be recognized that the process of sterile filtration serves the pharmaceutical industry incredibly well on a day-to-day basis.  The reasons for this include the fact that on account of sound hygiene and microbial control practices, including pre-filtration and redundant filtration, the bioburden level that reaches the final filter in most production lots is well below the 10^7 / sq cm level used in the validation studies, allowing for ample safety margins.  Further, in many cases, the actual filter used might well be far more integral than the test at 80% of the critical bubble-point would indicate.  And last, the statistical limitation of the sterility test would most likely cover for imperfections in the sterile filtration process, much like it did for other sterilization processes prior to the mandated implementation sound validation programs in the 1970’s.

Discussion

For all the above reasons, it must be concluded that sterile filtration is not in a state of real compliance.  In fact, the state of the art of this process is similar to the state of the art that prevailed for other sterilization processes over 25 years ago, prior to the general requirement for process validation.  Although the validation of sterile filtration is obviously required by cGMP regulations, the currently accepted approach differs considerably from the validation of other processes.  For instance, the validated cycle time is a parameter as key to moist heat sterilization as the critical bubble point pressure for a given filter-product combination is for sterile filtration.  While a reduction to 80% of the validated cycle time for the corresponding production sterilization run would be totally out of the question, the critical bubble-point for most products sterilized by filtration is often not even taken into consideration.

This lack of understanding is akin to accepting the validity of statements to the effect that “A 20-minute cycle will sterilize most loads” found in the instruction manuals of most old autoclaves.  When such statements were found to be misleading, rightfully so, the FDA did not question the validity of moist heat as a sterilization tool, but “encouraged” autoclave users to gain better understanding and control of key processing parameters through targeted validation studies.  Some of the cycles in use at the time were indeed found to be marginal, if not unsafe, when validated against what are now commonly accepted sterility assurance level targets.  Corrections were made and “real compliance” control was gained through the appropriately run validation studies.  The net effect was an overall improvement in the sterility assurance of terminally sterilized products.

The validity of sterile filtration as a tool must also not be questioned, but some of the currently accepted practices need to be revised in order to achieve a state of real compliance.  A key element in the success of such a revision is a more responsible involvement in sterile filtration issues by members of our own filter user community.  Validation of autoclaves by the user was also beyond the capability of some members of our industry 25 years ago, but the first Parenteral Drug Association Technical Report on Validation of Moist Heat Sterilization was instrumental in assisting autoclave users in this task.  The more recent Technical Report No. 26 on Sterile Filtration can help the filter user overcome most of the shortcomings discussed.  This industry guidance document, prepared with cooperation from the FDA, membrane filter manufacturers and, albeit, relatively few members of the filter user community, represents a great step towards providing the basic tools required by filter users to bring sterile filtration into a controlled state of real compliance.  The document encourages filter users to become more aware and more involved in the filter selection process, which ultimately will reduce the bias and potential for conflict of interest that shadows the present approach to filter validation.  Most importantly, the document encourages filter users to take more responsibility for product-specific retention test and to make sure that the membranes used for such tests are commensurate with the minimum integrity specified by the filter manufacturer or their own production test parameters.

It should be evident from the above discussions that the integrity test parameter selection must be taken more seriously than it is in current practice in order to optimize the sterility assurance of sterile products produced by aseptic processing techniques that involve sterile filtration.  Other than convenience, mainly for the filter manufacturer, there is no excuse to test the integrity of sterilizing filters in the production setting at pressures below the critical bubble point value.  As discussed elsewhere in greater detail (18), if the critical bubble-point for a given filter-product combination is not known, the lowest bubble-point value at which the desired, typically total, retention is observed in a series of product-specific challenge tests must be used as the diffusive flow test pressure for cartridge filters to be used in production.  This is the case even if the manufacturer has qualified the membrane as sterilizing grade by means of B. diminuta challenges.  After all, the retention may be affected by the nature of the product and is a function of the true integrity of the membrane, not an arbitrary rating.  Further, the retention capability at lower bubble-point values was not established, and unknown retention performance must be avoided to err on the safe side.  For the reasons discussed above, the integrity of filter cartridges used in the production setting must also be verified before and after the filtration process to assure retention equivalent to the level observed in the validation runs.

For decades, filter users have been lead to believe that sterilizing filters are capable of retaining to the specified level of 10^7 cfu / sq cm of filter area, but this goal can not be substantiated in light of current practices.  By bringing the process into a controlled state of real compliance, this retention level can more likely be achieved and validated appropriately, making products sterilized by filtration that much safer.  The fact that there have been no major catastrophes that can be traced to sterile filtration should, by rights, not stand in the way of a safer real compliance approach.  However, progress in this area might be slow.  As Madsen points out, “The primary obstacles to real compliance are inertia, people and the regulators…” (3).

References

- Journal of GXP Compliance, Institute of Validation Technology

- J. Berger, “Assessing the Quality Assurance of your Aseptic Processing Operations”, American Pharmaceutical Review, 4, (4), 31 (2001).

- R. E. Madsen, “Real Compliance and How to Achieve It”, PDA J. Pharm. Sci. Technol. 55 (2): 59 (2001).

- FDA "Guideline on Sterile Drug Products Produced by Aseptic Processing", Division of Manufacturing and Product Quality, Rockville, MD, June 1987.

- M. W. Jornitz, J. P. Agalloco, J. E. Akers, R. E. Madsen and T. H. Meltzer, “Filter Integrity Testing in Liquid Applications, Revisited” (Part I), Pharm. Technol. 25, (10), 34 (2001).

- M. W. Jornitz, J. P. Agalloco, J. E. Akers, R. E. Madsen and T. H. Meltzer, “Filter Integrity Testing in Liquid Applications, Revisited” (Part II), Pharm. Technol. 25, (11), 24 (2001).

- S. Sundaram, J. Eisenhuth, G. Howard Jr., and H. Brandwein, “Retention of Water-Borne Bacteria by Membrane Filters Part I, Bacterial Challenge Tests on 0.2 and 0.22 Micron Rated Filters”, PDA J. Pharm. Sci. Technol. 55 (2): 65 (2001).

- S. Sundaram, S. Mallick, J. Eisenhuth, G. Howard Jr., and H. Brandwein, “Retention of Water-Borne Bacteria by Membrane Filters Part II, SEM and FAME Characterization of Bacterial Species Recovered Downstream of 0.2 / 0.22 Micron Rated Filters”, PDA J. Pharm. Sci. Technol. 55 (2): 87 (2001).

- S. Sundaram, J. Eisenhuth, G. Howard Jr., and H. Brandwein, “Retention of Water-Borne Bacteria by Membrane Filters Part III, Bacterial Challenge Tests on 0.1Micron Rated Filters”, PDA J. Pharm. Sci. Technol. 55 (2): 114 (2001).

- P. R. Johnston, Fluid Sterilization by Filtration, Interpharm Press, Buffalo Grove, IL; 1992.

- T. H. Meltzer, “The Need for Initial or Prefiltration Integrity Testing”, Ultrapure Water 5 (1), (1988).

- FDA “Guideline on General Principles of Process Validation,” Division of Manufacturing and Product Quality, Rockville, MD, May 1987.

- K. H. Wallhausser, "Is the Removal of Microorganisms by Filtration Really a Sterilization Method?," J. Parenteral Drug Assoc., 33 (3), 156 (1979).

- K. H. Wallhausser, "Bacterial Filtration in Practice", Drugs Made in Germany, 19, 85 (1976).

- C. E. Weatherford, "A User's Perspective on Bacterial Retention Validation of Sterilizing Grade Cartridges," Pharm. Tech. Conference Proceedings, page 222-230 (1989).

- E. M. Fry, "FDA Update on Aseptic Processing Guideline," J. Parenteral Sci. Technol. 41 (2) 56 (1987).

- P. Cooney, Oral Presentation, FDA / PDA Sterile Filtration Symposium, Bethesda, MD (1995).

- H. G. Schroeder, “Rationalization and Valid Scale-Up of Integrity Test Parameters for Sterilizing Grade Filters,” PDA J. Pharm. Sci. Technol. 55 (2): 134 (2001).