## Sheet1
Troyonic BLM Tests for AD
using BLM test I-to-V amplifier
[15 empty] | Max % Change/100V => | 10 | 2 | 2 | 2
BLM |  | Exposure |  |  | Meter Reading |  |  | Leakage |  |  | Measured Current (pA) |  |  | Leakage Current | Leakage | Operating Voltage Slope Acceptable at: |  |  |  | Sensitivity @ 2000 volts
SN | Comments | Rate |  |  | @ Detector Bias (volts) |  |  | reading @ |  |  | @ Detector Bias (volts) |  |  | at 2000 volts | Current | [6 empty] | PASS/FAIL
(mmddyysn) |  | (R/hr) | 10 | 100 | 1000 | 1500 | 2000 | 2000 volts | 10 | 100 | 1000 | 1500 | 2000 | (pA) | OK? | 100 volts | 1000 volts | 1500 volts | 2000 volts | (nC/mR) | (R/Coul)
05270401 | BNL mods | 100 | -0.2142 | -0.2189 | -0.2315 | -0.2406 | -0.262 | 0 | 2142 | 2189 | 2315 | 2406 | 2620 | 0 | YES | YES | YES | YES | YES | 0.09432 | 10602205.2587 | PASS
05270402 | BNL mods | 100 | -0.2268 | -0.2395 | -0.2458 | -0.257 | -0.2801 | 0 | 2268 | 2395 | 2458 | 2570 | 2801 | 0 | YES | YES | YES | YES | YES | 0.10084 | 9917093.1017 | PASS
06100401 | BNL mods | 100 | -0.2219 | -0.2247 | -0.2545 | -0.3077 | -0.4114 | -0.0006 | 2219 | 2247 | 2545 | 3077 | 4114 | 6 | YES | YES | YES | NO | NO | 0.1481 | 6752012.0996 | FAIL
06100402 | BNL mods | 100 | -0.2114 | -0.2141 | -0.2361 | -0.2694 | -0.3318 | 0 | 2114 | 2141 | 2361 | 2694 | 3318 | 0 | YES | YES | YES | NO | NO | 0.11945 | 8371843.8149 | FAIL
06100403 | BNL mods | 100 | -0.2273 | -0.2281 | -0.239 | -0.2496 | -0.2655 | 0 | 2273 | 2281 | 2390 | 2496 | 2655 | 0 | YES | YES | YES | YES | YES | 0.09558 | 10462439.841 | PASS
06100404 | BNL mods | 100 | -0.2175 | -0.2215 | -0.2463 | -0.2897 | -0.3792 | -0.0001 | 2175 | 2215 | 2463 | 2897 | 3792 | 1 | YES | YES | YES | NO | NO | 0.13651 | 7325363.338 | FAIL
06100405 | BNL mods | 100 | -0.2319 | -0.2374 | -0.2499 | -0.259 | -0.2782 | 0.0001 | 2319 | 2374 | 2499 | 2590 | 2782 | -1 | YES | YES | YES | YES | YES | 0.10015 | 9984823.0689 | PASS
06110401 | BNL mods | 100 | -0.2171 | -0.2206 | -0.2462 | -0.2887 | -0.3681 | -0.0001 | 2171 | 2206 | 2462 | 2887 | 3681 | 1 | YES | YES | YES | NO | NO | 0.13252 | 7546258.565 | FAIL
06110402 | BNL mods | 100 | -0.2224 | -0.2262 | -0.2483 | -0.285 | -0.3525 | -0.0001 | 2224 | 2262 | 2483 | 2850 | 3525 | 1 | YES | YES | YES | NO | NO | 0.1269 | 7880220.6462 | FAIL
06110403 | BNL mods | 100 | -0.209 | -0.2136 | -0.2288 | -0.2398 | -0.2588 | -0.0003 | 2090 | 2136 | 2288 | 2398 | 2588 | 3 | YES | YES | YES | YES | YES | 0.093168 | 10733298.9868 | PASS
06110404 | BNL mods | 100 | -0.2173 | -0.2209 | -0.251 | -0.307 | -0.4225 | -0.0017 | 2173 | 2209 | 2510 | 3070 | 4225 | 17 | NO | YES | YES | NO | NO | 0.1521 | 6574621.9592 | FAIL
06110405 | BNL mods | 100 | -0.2081 | -0.2176 | -0.2336 | -0.2431 | -0.2741 | 0.0001 | 2081 | 2176 | 2336 | 2431 | 2741 | -1 | YES | YES | YES | YES | NO | 0.098676 | 10134176.4968 | FAIL
06110406 | BNL mods | 100 | -0.2198 | -0.2283 | -0.2357 | -0.2431 | -0.2611 | 0.0001 | 2198 | 2283 | 2357 | 2431 | 2611 | -1 | YES | YES | YES | YES | YES | 0.093996 | 10638750.5851 | PASS
05250401 | BNL mods | 100 | -0.2353 | -0.2466 | -0.2613 | -0.2802 | -0.3168 | -0.0001 | 2353 | 2466 | 2613 | 2802 | 3168 | 1 | YES | YES | YES | YES | NO | 0.11405 | 8768237.9349 | FAIL
06110407 | BNL mods | 100 | -0.1761 | -0.2173 | -0.2275 | -0.2315 | -0.2366 | 0 | 1761 | 2173 | 2275 | 2315 | 2366 | 0 | YES | NO | YES | YES | YES | 0.085176 | 11740396.3558 | FAIL
06110408 | BNL mods | 100 | -0.2171 | -0.2238 | -0.2357 | -0.2403 | -0.2537 | -0.0001 | 2171 | 2238 | 2357 | 2403 | 2537 | 1 | YES | YES | YES | YES | YES | 0.091332 | 10949064.9499 | PASS
06110409 | BNL mods | 100 | -0.2213 | -0.2317 | -0.2347 | -0.2369 | -0.2447 | -0.0002 | 2213 | 2317 | 2347 | 2369 | 2447 | 2 | YES | YES | YES | YES | YES | 0.088092 | 11351768.6055 | PASS
06110410 | BNL mods | 100 | -0.1699 | -0.2219 | -0.2372 | -0.2402 | -0.2453 | 0 | 1699 | 2219 | 2372 | 2402 | 2453 | 0 | YES | NO | YES | YES | YES | 0.088308 | 11324002.3554 | FAIL
06110411 | BNL mods | 100 | -0.2245 | -0.2339 | -0.2356 | -0.2403 | -0.246 | -0.0003 | 2245 | 2339 | 2356 | 2403 | 2460 | 3 | YES | YES | YES | YES | YES | 0.08856 | 11291779.5845 | PASS
06110401 | Prototype | 100 | -0.2008 | -0.21 | -0.2136 | -0.2192 | -0.2354 | -0.0001 | 2008 | 2100 | 2136 | 2192 | 2354 | 1 | YES | YES | YES | YES | YES | 0.084744 | 11800245.4451 | PASS
06140401 | New Prod. | 100 | -0.2135 | -0.2245 | -0.2291 | -0.2356 | -0.2455 | -0.0004 | 2135 | 2245 | 2291 | 2356 | 2455 | 4 | YES | YES | YES | YES | YES | 0.08838 | 11314777.0989 | PASS
06150401 | New Prod. | 100 | -0.2203 | -0.2321 | -0.237 | -0.2424 | -0.2527 | -0.0006 | 2203 | 2321 | 2370 | 2424 | 2527 | 6 | YES | YES | YES | YES | YES | 0.090972 | 10992393.2639 | PASS
06150402 | New Prod. | 100 | -0.2164 | -0.2293 | -0.234 | -0.2394 | -0.2487 | -0.0004 | 2164 | 2293 | 2340 | 2394 | 2487 | 4 | YES | YES | YES | YES | YES | 0.089532 | 11169190.9038 | PASS
06150403 | New Prod. | 100 | -0.207 | -0.2188 | -0.2232 | -0.2292 | -0.2393 | 0 | 2070 | 2188 | 2232 | 2292 | 2393 | 0 | YES | YES | YES | YES | YES | 0.086148 | 11607930.5381 | PASS
06150404 | New Prod. | 100 | -0.2124 | -0.2238 | -0.229 | -0.235 | -0.2455 | -0.0006 | 2124 | 2238 | 2290 | 2350 | 2455 | 6 | YES | YES | YES | YES | YES | 0.08838 | 11314777.0989 | PASS
06210401 | New Prod. | 100 | -0.2112 | -0.2245 | -0.238 | -0.2549 | -0.2831 | -0.0002 | 2112 | 2245 | 2380 | 2549 | 2831 | 2 | YES | YES | YES | YES | NO | 0.10192 | 9812002.0409 | FAIL
06210402 | New Prod. | 100 | -0.2224 | -0.2314 | -0.2359 | -0.2397 | -0.2478 | 0 | 2224 | 2314 | 2359 | 2397 | 2478 | 0 | YES | YES | YES | YES | YES | 0.089208 | 11209756.9725 | PASS
06210403 | New Prod. | 100 | -0.1728 | -0.2202 | -0.2332 | -0.2346 | -0.242 | -0.0002 | 1728 | 2202 | 2332 | 2346 | 2420 | 2 | YES | NO | YES | YES | YES | 0.08712 | 11478420.5693 | FAIL
06220401 | New Prod. | 100 | -0.2328 | -0.241 | -0.2428 | -0.2461 | -0.2539 | -0.0001 | 2328 | 2410 | 2428 | 2461 | 2539 | 1 | YES | YES | YES | YES | YES | 0.091404 | 10940440.2433 | PASS
06220402 | New Prod. | 100 | -0.226 | -0.2337 | -0.2401 | -0.2442 | -0.2531 | 0 | 2260 | 2337 | 2401 | 2442 | 2531 | 0 | YES | YES | YES | YES | YES | 0.091116 | 10975020.8525 | PASS
06220403 | New Prod. | 100 | -0.2177 | -0.2349 | -0.2459 | -0.2604 | -0.2847 | 0 | 2177 | 2349 | 2459 | 2604 | 2847 | 0 | YES | YES | YES | YES | YES | 0.10249 | 9756859.0719 | PASS
06220404 | New Prod. | 100 | -0.2227 | -0.2292 | -0.2562 | -0.3176 | -0.4349 | 0 | 2227 | 2292 | 2562 | 3176 | 4349 | 0 | YES | YES | YES | NO | NO | 0.15656 | 6387164.3545 | FAIL
07130401 | BNL mods | 100 | -0.2238 | -0.2365 | -0.24 | -0.2435 | -0.2518 | -0.0001 | 2238 | 2365 | 2400 | 2435 | 2518 | 1 | YES | YES | YES | YES | YES | 0.090648 | 11031682.9936 | PASS
07130402 | BNL mods | 100 | -0.2566 | -0.2697 | -0.2764 | -0.2853 | -0.307 | 0 | 2566 | 2697 | 2764 | 2853 | 3070 | 0 | YES | YES | YES | YES | YES | 0.11052 | 9048136.084 | PASS
07130403 | BNL mods | 100 | -0.2196 | -0.229 | -0.2339 | -0.24 | -0.2512 | -0.0002 | 2196 | 2290 | 2339 | 2400 | 2512 | 2 | YES | YES | YES | YES | YES | 0.090432 | 11058032.5548 | PASS
07150401 | BNL mods | 100 | -0.2007 | -0.2061 | -0.2176 | -0.2234 | -0.2363 | 0 | 2007 | 2061 | 2176 | 2234 | 2363 | 0 | YES | YES | YES | YES | YES | 0.085068 | 11755301.641 | PASS
07150402 | BNL mods | 100 | -0.173 | -0.2059 | -0.2157 | -0.2192 | -0.2295 | -0.0003 | 1730 | 2059 | 2157 | 2192 | 2295 | 3 | YES | NO | YES | YES | YES | 0.08262 | 12103606.8748 | FAIL
07150403 | BNL mods | 100 | -0.2178 | -0.2342 | -0.2487 | -0.2704 | -0.3052 | 0 | 2178 | 2342 | 2487 | 2704 | 3052 | 0 | YES | YES | YES | YES | NO | 0.10987 | 9101499.9272 | FAIL
07150404 | BNL mods | 100 | -0.2171 | -0.2314 | -0.2458 | -0.2653 | -0.2982 | 0 | 2171 | 2314 | 2458 | 2653 | 2982 | 0 | YES | YES | YES | YES | NO | 0.10735 | 9315150.1602 | FAIL
07150405 | BNL mods | 100 | -0.0759 | -0.1542 | -0.1742 | -0.1758 | -0.1768 | 0 | 759 | 1542 | 1742 | 1758 | 1768 | 0 | YES | NO | YES | YES | YES | 0.063648 | 15711412.7702 | FAIL
07150406 | BNL mods | 100 | -0.2138 | -0.2292 | -0.2469 | -0.2702 | -0.3105 | 0 | 2138 | 2292 | 2469 | 2702 | 3105 | 0 | YES | YES | YES | YES | NO | 0.11178 | 8946144.2118 | FAIL
07150407 | BNL mods | 100 | -0.1679 | -0.2018 | -0.2115 | -0.2151 | -0.2261 | -0.0001 | 1679 | 2018 | 2115 | 2151 | 2261 | 1 | YES | NO | YES | YES | YES | 0.081396 | 12285616.0008 | FAIL
07150408 | BNL mods | 100 | -0.215 | -0.2285 | -0.2337 | -0.2355 | -0.2397 | 0 | 2150 | 2285 | 2337 | 2355 | 2397 | 0 | YES | YES | YES | YES | YES | 0.086292 | 11588559.7738 | PASS
07200401 | New Prod. | 100 | -0.2303 | -0.2391 | -0.2441 | -0.2527 | -0.2666 | -0.0002 | 2303 | 2391 | 2441 | 2527 | 2666 | 2 | YES | YES | YES | YES | YES | 0.095976 | 10419271.4845 | PASS
07200402 | New Prod. | 100 | -0.2265 | -0.2362 | -0.2401 | -0.2443 | -0.2537 | -0.0002 | 2265 | 2362 | 2401 | 2443 | 2537 | 2 | YES | YES | YES | YES | YES | 0.091332 | 10949064.9499 | PASS
07200403 | New Prod. | 100 | -0.2202 | -0.2293 | -0.2354 | -0.2448 | -0.261 | -0.0004 | 2202 | 2293 | 2354 | 2448 | 2610 | 4 | YES | YES | YES | YES | YES | 0.09396 | 10642826.7348 | PASS
07200404 | New Prod. | 100 | -0.2218 | -0.2309 | -0.2354 | -0.2416 | -0.253 | -0.0004 | 2218 | 2309 | 2354 | 2416 | 2530 | 4 | YES | YES | YES | YES | YES | 0.09108 | 10979358.8054 | PASS
07200405 | New Prod. | 100 | -0.2139 | -0.2224 | -0.2352 | -0.2581 | -0.2979 | -0.0002 | 2139 | 2224 | 2352 | 2581 | 2979 | 2 | YES | YES | YES | YES | NO | 0.10724 | 9324530.9761 | FAIL
07200406 | New Prod. | 100 | -0.2096 | -0.2169 | -0.2253 | -0.2379 | -0.2599 | -0.0004 | 2096 | 2169 | 2253 | 2379 | 2599 | 4 | YES | YES | YES | YES | YES | 0.093564 | 10687871.4035 | PASS
07200407 | New Prod. | 100 | -0.2288 | -0.2402 | -0.2446 | -0.2523 | -0.2647 | -0.0006 | 2288 | 2402 | 2446 | 2523 | 2647 | 6 | YES | YES | YES | YES | YES | 0.095292 | 10494060.3618 | PASS
07200408 | New Prod. | 100 | -0.2193 | -0.2277 | -0.2325 | -0.2386 | -0.2521 | -0.0001 | 2193 | 2277 | 2325 | 2386 | 2521 | 1 | YES | YES | YES | YES | YES | 0.090756 | 11018555.247 | PASS
07210401 | New Prod. | 100 | -0.2086 | -0.2168 | -0.2292 | -0.2497 | -0.2847 | -0.0006 | 2086 | 2168 | 2292 | 2497 | 2847 | 6 | YES | YES | YES | YES | NO | 0.10249 | 9756859.0719 | FAIL
07210402 | New Prod. | 100 | -0.2168 | -0.2253 | -0.2392 | -0.2648 | -0.3094 | -0.0006 | 2168 | 2253 | 2392 | 2648 | 3094 | 6 | YES | YES | YES | NO | NO | 0.11138 | 8977950.1544 | FAIL
07210403 | BNL mods | 100 | -0.2556 | -0.2627 | -0.2712 | -0.2799 | -0.3014 | -0.0001 | 2556 | 2627 | 2712 | 2799 | 3014 | 1 | YES | YES | YES | YES | YES | 0.1085 | 9216250.0922 | PASS
07210404 | BNL mods | 100 | -0.2319 | -0.2423 | -0.2512 | -0.2678 | -0.2975 | -0.0002 | 2319 | 2423 | 2512 | 2678 | 2975 | 2 | YES | YES | YES | YES | NO | 0.1071 | 9337068.1606 | FAIL
07210405 | BNL mods | 100 | -0.2306 | -0.2365 | -0.2484 | -0.268 | -0.3038 | -0.0003 | 2306 | 2365 | 2484 | 2680 | 3038 | 3 | YES | YES | YES | YES | NO | 0.10937 | 9143442.3232 | FAIL
07210406 | BNL mods | 100 | -0.2287 | -0.2369 | -0.2391 | -0.2424 | -0.2504 | 0 | 2287 | 2369 | 2391 | 2424 | 2504 | 0 | YES | YES | YES | YES | YES | 0.090144 | 11093361.7323 | PASS
07210407 | BNL mods | 100 | -0.2218 | -0.2256 | -0.25 | -0.295 | -0.3797 | -0.0002 | 2218 | 2256 | 2500 | 2950 | 3797 | 2 | YES | YES | YES | NO | NO | 0.13669 | 7315717.0866 | FAIL
07220401 | BNL mods | 100 | -0.2258 | -0.2334 | -0.2449 | -0.264 | -0.3003 | -0.0002 | 2258 | 2334 | 2449 | 2640 | 3003 | 2 | YES | YES | YES | YES | NO | 0.10811 | 9250009.25 | FAIL
07220402 | BNL mods | 100 | -0.261 | -0.2758 | -0.2809 | -0.2897 | -0.3106 | -0.0003 | 2610 | 2758 | 2809 | 2897 | 3106 | 3 | YES | YES | YES | YES | YES | 0.11182 | 8943263.9336 | PASS
07290401 | BNL mods | 100 | -0.2552 | -0.2689 | -0.2748 | -0.2842 | -0.3041 | -0.0001 | 2552 | 2689 | 2748 | 2842 | 3041 | 1 | YES | YES | YES | YES | YES | 0.10948 | 9134422.1565 | PASS
07290402 | BNL mods | 100 | -0.258 | -0.2649 | -0.2752 | -0.2836 | -0.3042 | 0 | 2580 | 2649 | 2752 | 2836 | 3042 | 0 | YES | YES | YES | YES | YES | 0.10951 | 9131419.3878 | PASS
07290403 | BNL mods | 100 | -0.2199 | -0.2257 | -0.2642 | -0.3602 | -0.5681 | -0.0001 | 2199 | 2257 | 2642 | 3602 | 5681 | 1 | YES | YES | YES | NO | NO | 0.20452 | 4889592.9903 | FAIL
07290404 | BNL mods | 100 | -0.2323 | -0.2378 | -0.2573 | -0.2904 | -0.3515 | -0.0001 | 2323 | 2378 | 2573 | 2904 | 3515 | 1 | YES | YES | YES | NO | NO | 0.12654 | 7902639.4816 | FAIL
07290405 | BNL mods | 100 | -0.2623 | -0.2699 | -0.2766 | -0.2845 | -0.3031 | -0.0001 | 2623 | 2699 | 2766 | 2845 | 3031 | 1 | YES | YES | YES | YES | YES | 0.10912 | 9164558.8181 | PASS
07290406 | BNL mods | 100 | -0.225 | -0.2339 | -0.2524 | -0.2905 | -0.3649 | -0.0001 | 2250 | 2339 | 2524 | 2905 | 3649 | 1 | YES | YES | YES | NO | NO | 0.13136 | 7612435.6749 | FAIL
07290407 | BNL mods | 100 | -0.2162 | -0.2206 | -0.2287 | -0.2335 | -0.2434 | -0.0003 | 2162 | 2206 | 2287 | 2335 | 2434 | 3 | YES | YES | YES | YES | YES | 0.087624 | 11412398.4297 | PASS
07290408 | BNL mods | 100 | -0.2227 | -0.228 | -0.241 | -0.2583 | -0.2885 | 0 | 2227 | 2280 | 2410 | 2583 | 2885 | 0 | YES | YES | YES | YES | NO | 0.10386 | 9628345.8502 | FAIL
07300401 | BNL mods | 100 | -0.2185 | -0.2271 | -0.2673 | -0.3668 | -0.5847 | -0.0002 | 2185 | 2271 | 2673 | 3668 | 5847 | 2 | YES | YES | YES | NO | NO | 0.21049 | 4750774.3762 | FAIL
07300402 | BNL mods | 100 | -0.1722 | -0.2049 | -0.2158 | -0.2191 | -0.2291 | -0.0005 | 1722 | 2049 | 2158 | 2191 | 2291 | 5 | YES | NO | YES | YES | YES | 0.082476 | 12124739.3181 | FAIL
07300403 | BNL mods | 100 | -0.2212 | -0.2349 | -0.24 | -0.242 | -0.2465 | -0.0001 | 2212 | 2349 | 2400 | 2420 | 2465 | 1 | YES | YES | YES | YES | YES | 0.08874 | 11268875.3662 | PASS
07300404 | BNL mods | 100 | -0.002 | -0.0027 | -0.0121 | -0.0117 | -0.0103 | -0.0006 | 20 | 27 | 121 | 117 | 103 | 6 | YES | NO | NO | YES | YES | 0.003708 | 269687162.891 | FAIL
07300405 | BNL mods | 100 | -0.1751 | -0.2182 | -0.2298 | -0.2318 | -0.2378 | -0.0008 | 1751 | 2182 | 2298 | 2318 | 2378 | 8 | YES | NO | YES | YES | YES | 0.085608 | 11681151.2943 | FAIL
07300406 | BNL mods | 100 | -0.226 | -0.2358 | -0.2426 | -0.2464 | -0.2554 | -0.0001 | 2260 | 2358 | 2426 | 2464 | 2554 | 1 | YES | YES | YES | YES | YES | 0.091944 | 10876185.5042 | PASS
07300407 | BNL mods | 100 | -0.2572 | -0.2711 | -0.2764 | -0.2852 | -0.3057 | 0 | 2572 | 2711 | 2764 | 2852 | 3057 | 0 | YES | YES | YES | YES | YES | 0.11005 | 9086613.6008 | PASS
07300408 | BNL mods | 100 | -0.228 | -0.2357 | -0.2535 | -0.2923 | -0.3666 | -0.0002 | 2280 | 2357 | 2535 | 2923 | 3666 | 2 | YES | YES | YES | NO | NO | 0.13198 | 7577135.2367 | FAIL
08030401 | BNL mods | 100 | -0.2073 | -0.2133 | -0.2384 | -0.2808 | -0.3542 | -0.0003 | 2073 | 2133 | 2384 | 2808 | 3542 | 3 | YES | YES | YES | NO | NO | 0.12751 | 7842399.1467 | FAIL
08030402 | BNL mods | 100 | -0.2305 | -0.2438 | -0.247 | -0.2501 | -0.2578 | -0.0003 | 2305 | 2438 | 2470 | 2501 | 2578 | 3 | YES | YES | YES | YES | YES | 0.092808 | 10774933.1954 | PASS
08030403 | BNL mods | 100 | -0.1759 | -0.2236 | -0.237 | -0.239 | -0.2458 | -0.0001 | 1759 | 2236 | 2370 | 2390 | 2458 | 1 | YES | NO | YES | YES | YES | 0.088488 | 11300967.3628 | FAIL
08030404 | BNL mods | 100 | -0.2238 | -0.238 | -0.2412 | -0.2436 | -0.2488 | -0.0003 | 2238 | 2380 | 2412 | 2436 | 2488 | 3 | YES | YES | YES | YES | YES | 0.089568 | 11164701.6792 | PASS
08040401 | BNL mods | 100 | -0.2338 | -0.2425 | -0.249 | -0.26 | -0.2787 | -0.0002 | 2338 | 2425 | 2490 | 2600 | 2787 | 2 | YES | YES | YES | YES | YES | 0.10033 | 9966909.8593 | PASS
08040402 | BNL mods | 100 | -0.2557 | -0.2618 | -0.2703 | -0.2784 | -0.2984 | -0.0001 | 2557 | 2618 | 2703 | 2784 | 2984 | 1 | YES | YES | YES | YES | YES | 0.10742 | 9308906.762 | PASS
08040403 | BNL mods | 100 | -0.0756 | -0.1536 | -0.1734 | -0.175 | -0.1755 | -0.0003 | 756 | 1536 | 1734 | 1750 | 1755 | 3 | YES | NO | YES | YES | YES | 0.06318 | 15827793.6056 | FAIL
08040404 | BNL mods | 100 | -0.2342 | -0.2448 | -0.2581 | -0.2653 | -0.2805 | -0.0001 | 2342 | 2448 | 2581 | 2653 | 2805 | 1 | YES | YES | YES | YES | YES | 0.10098 | 9902951.0794 | PASS
08260401 | New Prod. | 100 | -0.2232 | -0.236 | -0.2408 | -0.2463 | -0.2567 | -0.0002 | 2232 | 2360 | 2408 | 2463 | 2567 | 2 | YES | YES | YES | YES | YES | 0.092412 | 10821105.4841 | PASS
08260402 | New Prod. | 100 | -0.2236 | -0.2388 | -0.2436 | -0.2488 | -0.2587 | -0.0007 | 2236 | 2388 | 2436 | 2488 | 2587 | 7 | YES | YES | YES | YES | YES | 0.093132 | 10737447.9234 | PASS
08260403 | New Prod. | 100 | -0.2237 | -0.2389 | -0.2442 | -0.2521 | -0.2661 | -0.0009 | 2237 | 2389 | 2442 | 2521 | 2661 | 9 | YES | YES | YES | YES | YES | 0.095796 | 10438849.2213 | PASS
08260404 | New Prod. | 100 | -0.2205 | -0.2338 | -0.2381 | -0.2435 | -0.2535 | -0.0002 | 2205 | 2338 | 2381 | 2435 | 2535 | 2 | YES | YES | YES | YES | YES | 0.09126 | 10957703.2654 | PASS
08260405 | New Prod. | 100 | -0.2236 | -0.2372 | -0.2424 | -0.2475 | -0.2574 | -0.0003 | 2236 | 2372 | 2424 | 2475 | 2574 | 3 | YES | YES | YES | YES | YES | 0.092664 | 10791677.4583 | PASS
08260406 | New Prod. | 100 | -0.2266 | -0.239 | -0.2442 | -0.2485 | -0.2595 | -0.0004 | 2266 | 2390 | 2442 | 2485 | 2595 | 4 | YES | YES | YES | YES | YES | 0.09342 | 10704345.9645 | PASS
08260407 | New Prod. | 100 | -0.2215 | -0.2372 | -0.2443 | -0.249 | -0.2568 | -0.0002 | 2215 | 2372 | 2443 | 2490 | 2568 | 2 | YES | YES | YES | YES | YES | 0.092448 | 10816891.658 | PASS