## C50 (CEBAF Cryomodule Rework) Vertical Cavity Testing Traveler

Abstract- Vertical cavity testing at 2K of re-worked 5-cell cavities for CEBAF cryomodules. Cavities can be tested singly or as part of a hermetic pair.

| Document No. | C50-CM-CST-CAV-VTRF |
| --- | --- |
| JLAB DCS # |  |
| Author | J. Ozelis |
| Date | 5/13/05 |
| Revision # | 1.0 |
| Approvals | Name: J. Mammosser Date: | Name: J. Preble Date: | Name: J. Ozelis Date: | Name: J. Mammosser Date: |
|  | Approver | Project Engineer | Technical Facility Rep. | User Reviewer |

| References |
| --- |
| Drawings |  |  |  |  |
| Documents | VTA SOP |  |  |  |
| Resources |  |  |  |  |
| Static Docs |  |  |  |  |

Revision Note

Initial release

| Part Description | Serial No. | Vendor |
| --- | --- | --- |
| Reworked CEBAF 5-cell cavity | [[C50CavID]] <<C50CavID>> | Accel/Interatom |
| Note any special handling, processing, or off-normal conditions associated with this cavity before this test. | [[Special Handling Comments ]] <<Text>> |

| Step No. | Instructions | Data Input |
| --- | --- | --- |
| 1 | Is this cavity part of a hermetic pair? If “yes”, note its position on the test stand (top/bottom). | [[Part of a pair?]] <<Check=Yes>> [[Cavity position]] {{Top, Bottom}} <<Radio>> |
| 2 | Has the cavity been baked? If so, indicate baking temperature and duration. | [[Baked?]] <<Check=Yes>> [[Bake Time]] <<Time>> [[Bake Temp (deg C)]] <<Float>> |
| 3 | Enter the LabView file name. | [[LabView file name]] <<Text>> |
| 4 | Record Test Date, Dewar No., Top Plate ID, and Operator(s) | [[Date]] <<Date>> [[Dewar no.]] <<Integer>> [[Top Plate ID] <<Text>> [[Test Operator1]] <<Text>> [[Test Operator2]] <<Text>> |
| 5 | Record cavity vacuum pressure, if so instrumented. | [[Cavity Vacuum ]] <<SciNot>> (mbar) |
| 6 | Record Dewar helium bath liquid level, temperature, pressure. Do not continue unless dewar LHe level is greater than/equal to 210cm for a pair test, or 145cm for a single cavity (bottom) test. | [[Dewar He liquid level]] <<Float>> (cm) [[Dewar Temperature]] <<Float>> (K) [[Dewar pressure ]] <<Float>> (torr) |
| 7 | Zero power meters and calibrate cables at cavity fundamental frequency (be sure to go off-resonance when calibrating cables from dewar top plate to cavity). | [[Power meters zeroed?]] <<Check=Yes>> [[Cable calibration OK?]] <<Check=Yes>> |
| 8 | Perform low power network analyzer measurements (measure the five cavity mode frequencies). Record the cavity mode frequencies at the right. | [[pi ]] <<Float>> (MHz) [[4/5pi]] <<Float>> (MHz) [[3/5pi]] <<Float>> (MHz) [[2/5pi]] <<Float>> (MHz) [[1/5pi]] <<Float>> (MHz) |
| 9 | Determine cavity coupling – Procedure: 1497 MHz Vertical Testing Procedure | [[Cavity Coupling]]{{Overcoupled, Undercoupled}} <<Radio>> |
| 11 | Determine the value of Q ext2 by performing decay measurements at low gradient (typically < 2-3 MV/m), and record 5 “good” sets of data. Make sure that the cavity Qo is constant in the gradient range used, and that there is no field emission (observed radiation). Procedure: 1497 MHz Vertical Testing Procedure | Enter data in table below. |
| Eacc (MV/m) | Qo | Qext1 | Qext2 | Q ext2 error% | Rad (mR/hr) |
| [[Eacc1]] <<Float>> | [[Qo1]]<<SciNot>> | [[Qextin1]]<<SciNot>> | [[Qextfp1]]<<SciNot>> | [[error1]] <<Float>> | [[rad1]] <<Float>> |
| [[Eacc2]] <<Float>> | [[Qo2]]<<SciNot>> | [[Qextin2]]<<SciNot>> | [[Qextfp2]]<<SciNot>> | [[error2]] <<Float>> | [[rad2]] <<Float>> |
| [[Eacc3]] <<Float>> | [[Qo3]]<<SciNot>> | [[Qextin3]]<<SciNot>> | [[Qextfp3]]<<SciNot>> | [[error3]] <<Float>> | [[rad3]] <<Float>> |
| [[Eacc4]] <<Float>> | [[Qo4]]<<SciNot>> | [[Qextin4]]<<SciNot>> | [[Qextfp4]]<<SciNot>> | [[error4]] <<Float>> | [[rad4]] <<Float>> |
| [[Eacc5]] <<Float>> | [[Qo5]]<<SciNot>> | [[Qextin5]]<<SciNot>> | [[Qextfp5]]<<SciNot>> | [[error5]] <<Float>> | [[rad5]] <<Float>> |
| 12 | Record Qext2 chosen for CW measurements and the % error in this value. | [[Qext for testing]] <<SciNot>> [[Qext error]] <<Float>> |
| 13 | Process the cavity and generate a Qo-vs-Eacc curve – 1497 MHz Vertical Testing Procedure |  |
| 14 | Record maximum gradient achieved, Qo value at maximum gradient, gradient at Qo spec (8 x 109), onset of field emission, and highest radiation inside dewar lid. | [[Emax]] <<Float>> (MV/m) [[Qo@Emax]] <<SciNot>> [[E @ Qo=8 x 109)]] <<Float>> (MV/m) [[FE Onset]] <<Float>> (MV/m) [[Radmax]] <<Float>> (mR/hr) |
| 15 | Record Qo at Eacc = ~2 and 12.5 MV/m. | [[Qo @ 2 MV/m]] <<SciNot>> [[Qo @ 12.5 MV/m]] <<SciNot>> |
| 16 | Record performance limitation. | [[Performance Limit]]{{RF Power, FE, Quench (non FE), Cable, Operator, Other (MP)}}<<Radio>> |
| 17 | Was cavity He or RF pulse processed ? Record pertinent information at the right. | [[He Processing]]? <<Check=Yes>> [[He Process Notes]] <<Comment>> [[Pulse Processing]]? <<Check=Yes>> [[Pulse Process Notes]] <<Comment>> |
| 18 | Performance notes - record information about cavity performance, limitations, and other observations. | [[Performance notes]] <<Comment>> |
| 19 | Upload the raw data file. | [[RF Test - raw data]] <<FileUpload>> |
| 20 | Start warm-up of cavity. |  |
| 21 | Upload processed (Excel) data file. | [[RF Test - processed data]] <<FileUpload>> |
| 22 | Upload processed Q0-vs-Eacc curve/data (in PDF format). | [[Q0-vs-Eacc Data]] <<FileUpload>> |
| 23 | Upload processed Rad-vs-Eacc curve/data (in PDF format). | [[Rad vs Eacc Data]] <<FileUpload>> |
| 24 | Cavity qualified for use in C50 cryomodule (pair) assembly? | [[Cavity meets specs ?]] Yes <<Check=Yes>> |