“Ring-of-Welding” and Calorimeter Noise Characteristics

# D0Note 4286 draft

##### October 31, 2003

## H. T. Diehl,[1] M. Arov,[2] A. Askew,[3] D. Bauer,[4] 
Y. Scheglov,[5] M. Cooke,[3] D. Edmunds,[6] J. Najdzion,[1] 
S. Sengupta [7]

[1] Fermi National Accelerator Laboratory, Batavia, IL 60510

[2] Northern Illinois University, DeKalb, IL 60115

[3] Rice University, Houston, TX 77005

[4] Indiana University, Bloomington, IN 47405

[5] Petersburg Nuclear Physics Institute, Russia

[6] Michigan State University, E. Lansing, MI  48824

[7] Florida State University, Tallahassee, FL 32306

### Abstract

The evening shift on July 29, 2003 encountered strong noise in the calorimeter trigger and precision readout.  With help from the Cryo Operator and experts lingering in the vicinity, the source of the noise was identified as a welding machine operated on DAB3. This note provides some details of the discovery of the source of the noise, what it looked like in the control room monitoring tools and in the reconstructed data, and some implications about the detector.  It also provides common examples of calorimeter noise not associated with the welding.

[1] Introduction: The Story

The evening shift on July 29, 2003 started off smoothly.  We had inherited store 2828 from the previous shift with a luminosity approximately 24e30/cm2/s.  Global run number 179331 was proceeding smoothly, producing good quality data at normal trigger rates.

By 17:00 we had noticed calorimeter hot cells in the L1 trigger examine plots. The noise seemed to be concentrated across all phi at eta = -11 and -12 according to the “hot cell finder”. The GM plots didn’t show particular excess energy in any histograms.  Logbook notes [1] from the previous shift mention some calorimeter noise in run 179330. They attributed it to a malfunctioning BLS board and noted that the board noise was not seen in the L1 calorimeter trigger Examine.  We continued to attribute the problem to the BLS but, confused about how the noise could be in the trigger, began paging calorimeter experts in search of advice.   We ended run 179331 at 17:20, reloaded pedestals and started run 179332.

By 18:11 the noise had gotten worse. It appeared in several control room monitors. The vertex Examine showed an unusual number of events with no L3 tracks.  The monitor that shows the L1 trigger rate for each trigger (Trigmon) showed some unusually sharp spikes in the rate. Figure [1] shows the trigger rate (Hz) for two low ET threshold EM triggers. The first is CEM(2,6)_ncu. The second is CEM(2,3)CEM(1,9)_ncu. One can see the sharp, brief spikes in the rate.

Fig. [1]. Trigger rate for two low ET EM triggers.  The vertical axis is the rate (hz). The horizontal axis is the time, in minutes, since (then) present. The gap centered on –47 minutes is the gap between runs 179331 and 179332.

Shortly after that we obtained new calorimeter trigger Exmine plots. They showed hot EM triggers concentrated at approximately eta = -7 and at no particular phi. However, the hadronic triggers seemed unaffected. Figure [2] shows calorimeter trigger Examine plots. By this time the control room event display monitor was demonstrating frequent events with tall stripes of calorimeter noise at about the same place in the calorimeter as the trigger noise and across all phi. There was large missing ET as well.

Fig. [2]. Selections from the Trigger Examine in run 179332. Note the hot spot in EM Trig eta and the lack of a corresponding hot spot in HAD Trig eta.

At about this time an expert (D.E.) arrived in the control room and he informed the shift crew that the noise was a long-standing serious problem called “Ring-of-Fire”. He told the shift crew that the problem had appeared on many occasions during Run II but that the cause was unknown to any experts.  It would eventually go away after a few hours.  The shift decided that it was worthwhile to put some effort into uncovering the source of the problem because the data being collected appeared useless.

The shift contacted the D0 cryo operator  (J.N.) and inquired about unusual conditions in the calorimeter controls which might be the source of the “Ring-of-Fire”. He arrived in the control room a few minutes later, prepared to show the shift crew how to understand the calorimeter cryogenic monitoring tools. “However,” he said, “You might want to know that there is welding going-on in DAB3”.  Light dawns over Marblehead!  By this time it was 18:58 and the welder was due to leave by 19:00.  It happened he had turned off his machine and was preparing to go when we found him via the building page.

We arranged a quick test. He fired the welding machine back up. We watched the L1 trigger rate display shown in part in Fig. [1].  Whenever he squeezed the welding machine trigger the spikes in the L1 rate instantly appeared. Then disappeared.  It was the initial trigger that caused the noise, not the subsequent welding arc. We thanked the welder. He left. We started a new run, number 179333. We had no further problems that evening.

We renamed the effect “Ring-of-Welding”.

Subsequent tests verified the observation [2].  Welding was discontinued for the time being at D0.  The hypothesis was developed that the building ground was not isolated from the detector ground. The “Ring-of-Fire” transmission path was demonstrated to be through the calorimeter liquid-Argon temperature monitoring probes [3]. During the fall 2003 shutdown, it was determined that the two grounds were not separated [4]. Investigations are continuing today to locate the short. It is not known at this time whether-or-not the “Ring-of-Welding” noise would have occurred if the grounds were isolated.  I suppose we will find out in time.  The appendix lists the dates and times of welding that occurred at D0 during 2003 [5].

[2] Ring-of-Welding in RECO

We analyzed run 179332 with RecoCert version p14.03.01 [6].  We found hot regions in the calorimeter corresponding to that which we saw during data taking.  Figure [3] shows calorimeter hits in run 179332 (black) compared to 179333 (red) showing the “Ring-of-Welding” hits.  Figure [4] shows the eta distribution of  “all” and “good” s-cone electrons in run 179332  (black) compared to run 179333 (red).  One notes the spike at eta approximately -1.2, in correspondence with the hits, in the plot showing “all” electrons.  The spike is gone when one applies electron selection criteria.  The distribution was flat in phi in both cases, with and without the electron selection.  Figures [5] and [6] show the distribution of  “all” jets and “good” jets in those two runs.  Figure [7] shows the missing transverse energy in those two runs.  The effect of the “Ring-of-Welding” is evident in run 179332.

Figure [3]. IETA (top-left) and IPHI (top-right) of calorimeter cells in runs 179332 (black) and 179333 (red).  Note that the “Ring-of-Welding” run has excess hit cells in the EC’s but that the distribution is flat in phi.  Cell energy in two dimension (bottom-left).  Note that at IETA approximately –12 there is a light-brown colored excess spanning phi.

Figure [4]. The eta distribution of All simple-cone electrons (left) and good simple-cone electrons (right) in runs 179332 (black) and 179333 (red).

Figure [5]. All Jets found with simple cone of radius 0.7 in run 179332 (black) and run 179333 (red). The “Ring-of-Welding” run has more jets at high-pT particularly clustered at |eta| ~ 1.2. The EMF of those jets is smaller than normal. Perhaps the noise was in the hadronic sections as well as EM, but that those cells weren’t part of the hadronic trigger.

Figure [6]. “Good” Jets found with simple cone of radius 0.7 in run 179332 (black) and run 179333 (red). The “Ring-of-Welding” run has approximately the same number of jets at high-pT and the cluster at |eta| ~ 1.2 is gone. The EMF of those jets is more normal. The eta distribution continues to show a difference.

Figure [7]. Missing ET total and components in run 179332 (black) and run 179333 (red). The “Ring-of-Welding” run has more events with high MET, particularly in the tail and many may be off the plot. The strongest indication of the difference is that the good run has more events at low MET, for these histograms are normalized to each other.

[3] Examples of Other Calorimeter Noise Sources

We have used RecoCert to investigate calorimeter noise that has occurred at other times during RunII.  In this section we show two other examples: a bad ADC in run 175645, and uncalibrated pedestals after a BLS power supply was changed prior to run 180301.

[3.1] Run 175645 – a bad ADC

Run 175645 was recorded on April 12, 2003. The CalMuo shifter noted [7] the large number of hits in a particular part of the calorimeter and that there were an unexpected number of events with missing ET larger than 60 GeV. Figure [8] shows a Lego plot event display picture recorded by the shifter. The hot cells stand in a line of constant phi and the missing ET is large.

Figure [8]. Event display (Lego) plot from run 175645. The hot cells are evident as is the large missing transverse energy (yellow spike).

Figure [9]. All calorimeter cells RecoCert histograms from run 175645. Notice the lump in the “PT Sum” and “E Sum” figures in the top row. These are due to the energy deposited in the hot spot at phi ~ 4.4 to 4.6, corresponding to a 16-channel ADC.

Figure [10]. The missing transverse energy distribution in run 175645. The lump centered on ~ 30 GeV in the x-component is the effect of the bad ADC. Presumably there’s a similar lump offscale someplace on the y-component. The beginning of the lump can be seen in the total MET plot at 100 GeV.

We ran RecoCert to check the effect of the line of hot cells. Figure [9] shows the hits in the central calorimeter.  Notice the lump in the “PT Sum” and “E Sum” figures in the top row. These are due to the energy deposited in the hot spot at phi ~ 4.4 to 4.6, corresponding [8] to a single 16-channel ADC.  An ADC is 8-pairs of IETA channels spanning more than one row of IPHI (6 pairs of cells per row). All of the cells of EM3 had the same large energy.  Figure [10] shows the missing transverse energy distribution.  N. Parua replaced the ADC during the next available access, which occurred on that same day [9]. The problem was solved.

[3.2] Run 180301 – Pedestals need updating after BLS supply exchange

A calorimeter BLS supply failed on August 19, 2003 [10] and was replaced that evening [11].  Run 180301 was recorded on August 20, 2003 after the BLS supply was replaced but before a stable set of new pedestals were available [12] or before a new calibration was done [13]. We studied this run with RecoCert and compared the run with 180329, taken after the next calibration [13]. Figure [11] shows all central calorimeter cells for run 180301 and run 180329. The hot spots are evident in the IETA and IPHI plots.  Figure [12] shows “all” jets in run 180301 (black) vs. run 180329 (red). The eta and phi plots and the correlated eta-phi plot show an excess of jets corresponding to the excess of hits shown in the previous figure.  Study of the “good” jets histograms indicated that the jet selection criteria could not remove all of the effects of the extra noise. Substantial jet excesses were seen in the same phi and eta distributions.  Figure [13] shows the missing transverse energy distributions. There was a 5 GeV shift in the uncalibrated run.

Figure [11]. Calorimeter all central cells for run 180301 (black) and run 180329 (red). The hot spots are evident in the IETA and IPHI plots.

Figure [12]. All jets in run 180301 (black) vs. run 180329 (red). The eta and phi plots and the correlated eta-phi plot show an excess of jets corresponding to the excess of hits shown in the previous figure.

Figure [13]. Missing transverse energy and components in run 180301 (black) and run 180329 (red).

[4] Summary and Discussion

The evening shift on July 29, 2003 discovered the cause of severe noise in the calorimeter and trigger named “Ring-of-Fire”. We renamed it “Ring-of-Welding”. We characterize it and some of its effects on the reconstructed events in this note.  This kind of calorimeter noise was not supposed to be able to occur because it was thought that the detector and building grounds were isolated from each other.  They are not and investigations have been underway in the shutdown this fall.

“Ring-of-Welding” is a serious form of calorimeter noise, but it is not the only one. We have showed evidence of two other kinds of calorimeter noise in this note. One is the result of a bad ADC.  The other the result of a delay in obtaining calorimeter pedestals after a BLS supply was changed. While the effects aren’t the same as in “Ring-of-Welding” data they are serious enough to question the data’s quality.

We expect that this is not a comprehensive list of causes of calorimeter noise. We know and appreciate that calorimeter experts are working hard to identify and suppress them. 
A worthwhile contribution aimed at assisting them might be to make a systematic quantification or classification of “bad runs” as scored in the “Run Quality Database” [15].  Perhaps the causes can be accounted-for in the logbook or by experts.

#### Appendix: Welding at DZero in 2003

NOTE:

Times shown are approximate. Some notes were recorded from memory days after the fact, and may be inaccurate.  Actually welding usually not continuous for periods delineated.  Set-up, leak checking and superinsulation activities are necessarily interspersed with weld times.

Date	    	Times							Location

7/29/03:  	1530-1900 including “Ring of Fire Test”   		 Weld Shop

7/28/03:  	1530-1800							“

7/24/03:	1530-1800	N-305						“

7/23/03:	1530-1800	N-305						“

7/17/03:	1530-1845	Separator Field joints			Refrigerator Area

7/16/03:	1530-1845		“					“

7/15/03:	1530-1845		“					“

7/14/03:	1530-1845		“					“

7/8/03:		1530-2000	Separator Field joints, inner lines		“

7/7/03:		1530-2000	Separator Field joints, inner lines		“

7/3/03:		1530-2020		“		inner lines		“

7/2/03		1530-2000		“		inner lines		“

7/1/03:		1530-2000	Subassemblies and field joints	Weld shop & Refer Areas

6/30/03:	1530-2000	Subassemblies & field joints			“	“

6/26/03:	1530-2000	Subassemblies				Weld shop

6/25/03:	1530-2000	Subassemblies				Weld shop

6/24/03:	1530-2000	Subassemblies				Weld shop

6/23/03:	1530-2000	Subassemblies				Weld shop

6/4/03:		1300-1530? (half day could have been morning)	Weld shop

4/28/03:	0730-1530	Separator subassembly		Weld shop

4/14/03:	0730-1530		“					“

4/11/03:	0730-1530		“					“

3/18/03: 	?		Weld relief tree to separator vessel		“

3/3/03:		0730-1530	VLPC Test  LN2 Rtn U-tube			“

2/28/03	?(45 min of welding )						“

2/17/03:	0730-1530	VLPC	Test					“

1/22/03:	0730-1530	VLPC Test Phase sep. jumper		“

1/17/03:	0730-1530	VLPC Test					“

1/16/03:	0730-1530	VLPC Test					“

1/15/03:	0730-1530	VLPC Test					“

1/14/03:	0730-1530	VLPC Test					“

1/6/03:		0730-1530	VPLC Test					“

#### Acknowledgements

We thank the experts, particularly Dean Schamberger, Robert Zitoun, Nirmalya Parua,  V. Shary, Laurent Duflot and others for their work in maintaining and operating the calorimeter and Harry Melanson and others for writing and maintaining RecoCert.

B. Vachon was co-daq-expert on shift. Her nagging the shift to understand the problem was important in getting to the cause of the noise.

References

[1] D. Kaefer, CALMUO Entry 227292, D0 Control Room Logbook, July 29, 2003.

[2] R. Rucinski, Communication to operations email list on or about July 31, 2003. The tests were carried out by B. Barger, D. Edmunds, D. Schamberger and a welder. The results were independent of welding machine and location of the machine in the building.

[3] D. Schamberger, public and private communications.

[4] D. Denisov, D. Schamberger, and R. Zitoun, public and private communications.

[5] The complete contents of the appendix comes from B. Barger via R. Rucinski.

[6] RecoCert is a product written so as to look at the quality of the output of the reconstructed data. It’s maintained by Harry Melanson and is available for use by anyone interested.

[7] L. Duflot, CALMUO Logbook entries 193929, 193947, 193968, and 193978.

[8] D. Schamberger and R. Zitoun, private communication.

[9] L. Duflot, CALMUO Logbook entry 194033.

[10] V. Shary, CALMUO Logbook entry 233940.

[11] V. Shary, CALMUO Logbook entry 233993.

[12] L. Duflot, CALMUO Logbook entry 234170.

[13] V. Shary, CALMUO Logbook entry 234300.

[14] V. Shary, CALMUO Logbook entry 234393.

[15] G. Bernardi keeps track of JET/MET run quality in the “Run Quality Database” located at URL http://d0db.fnal.gov/qualitygrabber/qualQueries.html.