# STANDARD OPERATING PROCEDURE - 4006

PREPARING AND REMOVING LOADS FROM STERILIZERS

PURPOSE:  To establish a standard operating procedure that identifies processes that will ensure proper item placement and removal from the sterilization carriage.

POLICY:  Loading procedure for the sterilizer must allow for free circulation of sterilant around each package and item.

RESPONSIBILITY:  SPD Technicians

PROCEDURE:

1.  Loading:

a.   The sterilizer carts and baskets must be loaded in an orderly manner.

b.   Adequate space must be provided so that the sterilant can freely circulate.  A

vertical hand should be able to easily pass between the individual packages.

- Small items should be placed in a wire basket or similar container to keep

items from moving around during the sterilization cycle.  The container must

allow for free circulation of the sterilant.

- All textile packages are placed on edge to promote passage of steam through the pack.

- Materials capable of holding water, such as basins and solid bottom trays, must be positioned on edge, faced in the same direction to allow for condensation elimination.

- In a mixed load, linen items should be placed on top shelf and metal items on bottom shelf.

- Packaging systems such as peel packs are to be placed on edge, paper to plastic and properly spaced to allow for free circulation of the sterilant.

- Peel packs may be used to isolate small items in container systems.

- All perforated sterilization containers must be placed flat on the cart to assure effective sterilization.

- Items placed on the sterilization cart must not come into contact with the sterilizer wall.  Allow at least three inches between the sterilizer chamber and the top most package of the load.

2.  Unloading:

- Items that have been removed from the sterilizer should not be handled during the cooling process.

- During the cooling process, the sterilization carriage should be placed in a low traffic area where no air conditioning or other cold vents are in close proximity.

- Warm items must never be transferred from the carriage to cold metal racks or shelving.

- Items must be handled as little as possible to avoid the increased risk of contamination.

- Items removed from the cart after cooling should be inspected for any compromise in packaging.

- Lot control numbers and expiration dates are then applied. The lot control number is a six-digit number with the first number being the sterilizer number, the next three numbers are the Julian date and the last two digits are the cycle number.  Shelf life date – 30 days or one year.

3.  Other considerations:

#### Steam Sterilization

- Sterilization of items such as reusable specialty needles, reusable

suction instruments, etc. will be more effective by adding moisture or injecting water into the lumen.

- Items for steam sterilization requiring lubrication must be lubricated with an approved water soluble instrument lubricant.

Ethylene Oxide (EtO)

- Items for EtO must be completely dried.

#### Sterrad

- Items that are made of cellulose materials such as paper, cloth, wood, and cotton will not be processed in the Sterrad.

- Observe estrictions on lumen size and length for successful processing.

- Items must be completely dried.

- Packaging must be of an approved container or wrapping material.

- Flexible endoscopes and implantable items are not to be processed in Sterrad.

References:  AAMI-SSA-1992, 4.7, 4.72

VA Handbook 7176, Supply, Processing and Distribution Operational

Requirements

Manufacturer’s Guidelines