Draft Guidance for

Industry and FDA Staff

Class II Special Controls Guidance Document: Bone Sonometers

Draft Guidance

This guidance document is being distributed for comment purposes only.

Document issued on: [OSM will insert release date of FR Notice on posting]

Comments and suggestions regarding this draft document should be submitted within 90 days of publication in the Federal Register of the notice announcing the availability of the draft guidance.  Submit written comments to the Division of Dockets Management (HFA-305), Food and Drug Administration, 5630 Fishers Lane, rm. 1061, Rockville, MD 20852.  Alternatively, electronic comments may be submitted to http://www.fda.gov/dockets/ecomments.  All comments should be identified with the docket number listed in the notice of availability that publishes in the Federal Register.

For questions regarding this document contact Robert A. Phillips, 301-594-1212, ext. 130.

When final this document will supersede Bone Sonometer PMA Applications; Final Guidance for Industry and FDA, issued June 21, 2001.

U.S. Department of Health and Human Services

Food and Drug Administration

Center for Devices and Radiological Health

Radiological Devices Branch
Division of Reproductive, Abdominal, and Radiological Devices

Office of Device Evaluation

Preface

Additional Copies

Additional copies are available from the Internet at: http://www.fda.gov/cdrh/[specific address], or to receive this document via your fax machine, call the CDRH Facts-On-Demand system at 800-899-0381 or 301-827-0111 from a touch-tone telephone.  Press 1 to enter the system.  At the second voice prompt, press 1 to order a document.  Enter the document number (1547) followed by the pound sign (#).  Follow the remaining voice prompts to complete your request.

Table of Contents

1.	Introduction	1

2.	Background	2

3.	The Content and Format of an Abbreviated 510(k) Submission	2

4.	Scope	5

5.	Device Description	6

6.	Risks to Health	7

7.	Electrical Safety	8

8.	Electromagnetic Compatibility	8

9.	Acoustic Intensity	9

10.	Non-Clinical Testing	9

11.	Clinical Testing	10

12.	Labeling	14

Draft Guidance for Industry and FDA Staff

Class II Special Controls Guidance Document: Bone Sonometers

This draft guidance, when finalized, will represent the Food and Drug Administration's (FDA's) current thinking on this topic.  It does not create or confer any rights for or on any person and does not operate to bind FDA or the public.  You can use an alternative approach if the approach satisfies the requirements of the applicable statutes and regulations.  If you want to discuss an alternative approach, contact the FDA staff responsible for implementing this guidance.  If you cannot identify the appropriate FDA staff, call the appropriate number listed on the title page of this guidance.

# Introduction

This draft guidance document was developed as a special controls guidance to support the reclassification of bone sonometers into class II (special controls).  The device transmits ultrasound energy into the human body to measure the acoustic properties of bone that indicate overall bone health and fracture risk.

This draft guidance will be issued in conjunction with a Federal Register notice announcing the proposal to reclassify this device type.  This guidance is issued for comment purposes only.  If a final rule to reclassify this device type is not issued, this guidance document will not be issued as a special control.

Following the effective date of a final rule reclassifying the device, any firm submitting a premarket notification submission (510(k)) for a bone sonometer will need to address the issues covered in the special control guidance.  However, the firm need only show that its device meets the recommendations of the guidance or in some other way provides equivalent assurances of safety and effectiveness.

When final this document will supersede the guidance entitled, Bone Sonometer PMA Applications, issued June 21, 2001 (http://www.fda.gov/cdrh/ode/guidance/1377.html).

The Least Burdensome Approach

This draft guidance document reflects our careful review of what we believe are the relevant issues related to bone sonometers and what we believe would be the least burdensome way of addressing these issues.  If you have comments on whether there is a less burdensome approach, however, please submit your comments as indicated on the cover of this document.

FDA's guidance documents, including this guidance, do not establish legally enforceable responsibilities.  Instead, guidances describe the Agency's current thinking on a topic and should be viewed only as recommendations, unless specific regulatory or statutory requirements are cited.  The use of the word should in Agency guidances means that something is suggested or recommended, but not required.

# Background

FDA believes that special controls, when combined with the general controls, will be sufficient to provide reasonable assurance of the safety and effectiveness of bone sonometers.  Thus, a manufacturer who intends to market a device of this generic type should (1) conform to the general controls of the Federal Food, Drug, and Cosmetic Act (the act), including the premarket notification requirements described in 21 CFR 807 Subpart E, (2) address the specific risks to health associated with bone sonometers identified in this guidance, and (3) obtain a substantial equivalence determination from FDA prior to marketing the device.

This special control guidance document identifies the classification regulation and product code for the bone sonometer (Please refer to Section 4.  Scope).  In addition, other sections of this special control guidance document list the risks to health identified by FDA and describe measures that, if followed by manufacturers and combined with the general controls, will generally address the risks associated with these bone sonometers and lead to a timely 510(k) review.  This document supplements other FDA documents regarding the specific content requirements of a premarket notification submission.  You should also refer to 21 CFR 807.87 and “How to Prepare a 510(k) Submission” on FDA’s Device Advice at http://www.fda.gov/cdrh/devadvice/314.html.

As described in the guidance entitled, The New 510(k) Paradigm - Alternate Approaches to Demonstrating Substantial Equivalence in Premarket Notifications; Final Guidance, http://www.fda.gov/cdrh/ode/parad510.html, a manufacturer may submit a Traditional 510(k) or has the option of submitting either an Abbreviated 510(k) or a Special 510(k).  FDA believes an Abbreviated 510(k) provides the least burdensome means of demonstrating substantial equivalence for a new device, particularly once FDA issues a class II special controls guidance document.  Manufacturers considering certain modifications to their own cleared devices may lessen the regulatory burden by submitting a Special 510(k).

# The Content and Format of an Abbreviated 510(k) Submission

An Abbreviated 510(k) submission must include the required elements identified in 21 CFR 807.87, including the proposed labeling for the device sufficient to describe the device, its intended use, and the directions for its use.  In an Abbreviated 510(k), FDA may consider the contents of a summary report to be appropriate supporting data within the meaning of 21 CFR 807.87(f) or (g); therefore, we recommend that you include a summary report.  The report should describe how this special control guidance document was used during the device development and testing and should briefly describe the methods or tests used and a summary of the test data or description of the acceptance criteria applied to address the risks identified in this document, as well as any additional risks specific to your device.  This section suggests information to fulfill some of the requirements of section 807.87 as well as some other items that we recommend you include in an Abbreviated 510(k).

Coversheet

The coversheet should prominently identify the submission as an Abbreviated 510(k) and cite the title of this special controls guidance document.

Proposed labeling

Proposed labeling should be sufficient to describe the device, its intended use, and the directions for its use.  (Please refer to Section 11.  Labeling for specific information that should be included in the labeling for devices of the type covered by this guidance document.)

Summary report

We recommend that the summary report contain:

Description of the device and its intended use

We recommend that you describe the performance specifications and, when appropriate, include detailed, labeled drawings of the device.  (Please refer to Section 5.  Device Description for specific information that we recommend you include in the device description for devices of the type covered by this guidance document.)  You should also submit an “indications for use” enclosure.

Description of device design requirements

We recommend that you include a brief description of the device design requirements.

Identification of the risk analysis method

We recommend that you identify the risk analysis method(s) you used to assess the risk profile, in general, as well as the specific device’s design and the results of this analysis.  (Please refer to Section 6.  Risks to Health for the risks to health generally associated with the use of this device that FDA has identified.)

Discussion of the device characteristics

We recommend that you discuss the device characteristics that address the risks identified in this class II special controls guidance document, as well as any additional risks identified in your risk analysis.

Description of the performance aspects

We recommend that you include a brief description of the test method(s) you have used or intend to use to address each performance aspect identified in Sections 7 - 10 of this class II special controls guidance document.  If you follow a suggested test method, you may cite the method rather than describing it.  If you modify a suggested test method, you may cite the method but should provide sufficient information to explain the nature of and reason for the modification.  For each test, you may either (1) briefly present the data resulting from the test in clear and concise form, such as a table, or (2) describe the acceptance criteria that you will apply to your test results.  (See also 21 CFR 820.30, Subpart C - Design Controls for the Quality System Regulation.)

Reliance on standards

If any part of the device design or testing relies on a recognized standard, we recommend that you include either:

- a statement that testing will be conducted and meet specified acceptance criteria before the device is marketed or

- a declaration of conformity to the standard.

Because a declaration of conformity is based on results from testing, we believe you cannot properly submit a declaration of conformity until you have completed the testing the standard describes.  For more information, please refer to section 514(c)(1)(B) of the Act and the FDA guidance, Use of Standards in Substantial Equivalence Determinations; Final Guidance for Industry and FDA.http://www.fda.gov/cdrh/ode/guidance/1131.html.

If it is not clear how you have addressed the risks identified by FDA or additional risks identified through your risk analysis, we may request additional information about aspects of the device’s performance characteristics.  We may also request additional information if we need it to assess the adequacy of your acceptance criteria.  (Under 21 CFR 807.87(l), we may request any additional information that is necessary to reach a determination regarding substantial equivalence.)

As an alternative to submitting an Abbreviated 510(k), you can submit a Traditional 510(k) that provides all of the information and data required under 21 CFR 807.87 and described in this guidance.  A Traditional 510(k) should include all of your methods, data, acceptance criteria, and conclusions.  Manufacturers considering certain modifications to their own cleared devices should consider submitting Special 510(k)s.

The general discussion above applies to any device subject to a special controls guidance document.  The following is a specific discussion of how you should apply this special controls guidance document to a premarket notification submission for a bone sonometer.

# Scope

The scope of this document is limited to the device described below.

21 CFR 892.1180, class II [product code to be designated, if a final rule is published].

A bone sonometer is a device that transmits ultrasound energy into the human body to measure acoustic properties of bone that indicate overall bone health and fracture risk.  The primary components of the device are a voltage generator, a transmitting transducer, a receiving transducer, and hardware and software for reception and processing of the received ultrasonic signal.

This guidance addresses the following indications:

- determining the possible presence of osteoporosis and assessing fracture risk

- monitoring bone changes over time

- assessing non-age-related bone loss.

# Device Description

We recommend that you identify your device, by the regulation and product code described in section 4.  Scope, and provide the information discussed below.

## Sonometer

### We recommend that you provide a complete description of the device, including:

### indications for use

### physical description of the device, including how transducers are positioned on the body and anatomical site(s) scanned

### engineering or block diagrams showing the major components

### block diagram of the data analysis program

### calibration standards used (built-in and external).

## Hardware

### We recommend that you provide the following specifications:

### transducer diameter, resonant frequency, bandwidth, focusing properties

### A/D converter sampling rate and bit depth.

## Software

We recommend that you include the appropriate software documentation as described in the guidance titled, Guidance for the Content of Premarket Submissions for Software Contained in Medical Devices at www.fda.gov/cdrh/ode/guidance/337.html.  As discussed in that guidance, we recommend that you identify the “level of concern” (minor, moderate, or major) associated with your device and provide documentation consistent with that level.  Generally, we consider “minor level of concern” appropriate for the bone sonometers addressed in this class II special controls guidance document.

## Beam

### We recommend that you describe the beam (e.g., size, spectra, intensity).  The description should include the average and maximum acoustic intensities measured for the device (e.g., ISPPA.3, ISPTA.3, MI). See also the guidance entitled, Information for Manufacturers Seeking Marketing Clearance of Diagnostic Ultrasound Systems and Transducers, at http://www.fda.gov/cdrh/ode/ulstran.pdf.

## Phantoms

If the device uses phantoms, for example in quality control or maintenance procedures, we recommend that you discuss the design, including contents, of any phantoms used.  We also recommend that you provide a product sheet detailing the technical and descriptive specifications of the phantom(s) and explain how the phantoms are used.

## Accessories

We recommend that you list all accessories that you intend to make available for use with your bone sonometer.  In addition, for each of your accessories, we recommend that you provide:

- a complete description, including photographs or drawings, and construction, including materials

- any material standards met

- any previous clearance.

If an accessory is exempt from the 510(k) requirements of the act, we recommend that you indicate its classification regulation.  If your accessory is not exempt or previously cleared for this use, bundling with your bone sonometer submission may be appropriate.  For information about bundling premarket submissions see the guidance entitled, Bundling Multiple Devices or Multiple Indications in a Single Submission.

# Risks to Health

In the table below, FDA has identified the risks to health generally associated with the use of the device addressed in this document.  The measures recommended to mitigate these identified risks are given in this guidance document, as shown in the table below.  We recommend that you also conduct a risk analysis, befeore submitting your 510(k), to identify any other risks specific to your device and include the results of this analysis in your 510(k).  If you elect to use an alternative approach to address a particular risk identified in this document, or have identified risks additional to those in this document, then you should provide sufficient detail to support the approach you have used to address that risk.

| Identified risk | Recommended mitigation measures (see guidance section below) |
| --- | --- |
| Electrical shock | 7. Electrical Safety |
| Electromagnetic interference | 8. Electromagnetic compatibility |
| Tissue damage | 9. Acoustic Intensity |
| Inaccurate measurement leading to inappropriate diagnosis or therapy | 10. Non-Clinical Testing 11. Clinical Testing 12. Labeling |

# Electrical Safety

We recommend that you evaluate the electrical safety of your device and its ability to function after exposure to environmental handling hazards.  We recommend that you evaluate your device according to one or more of the following standards (or equivalent methods):

- International Electrotechnical Commission (IEC) 60601-1 Medical Electrical Equipment - Part 1: General Requirements for Safety

- Underwriters Laboratory (UL) 2601-1 Amendment 1 Medical Electrical Equipment: General Requirements for Safety

- American National Standards Institute (ANSI)/AAMI ES-1 Safe current limits for electromedical apparatus

- IEC 60529 Degrees of protection provided by enclosures (IP Code) Consolidated Edition

- IEC 60721-4-x TR (Technical Reports).

The features and design of your device will determine which of the above standards or equivalent methods you should use and whether other standards are appropriate in addition to or in place of these.

# Electromagnetic Compatibility

Electromagnetic compatibility (EMC) encompasses both emissions (interference with other electronic devices) and immunity (resistance to interference, which is created by emissions from other electronic devices, with the performance of the device).  We recommend that you evaluate the EMC of your device according to IEC 60601-1- 2 Medical Electrical Equipment -- Part 1: General Requirements for Safety; Electromagnetic Compatibility -- Requirements and Tests (Second Edition, 2001) or equivalent method.

## Emissions

EMC testing should demonstrate that the device will not adversely interfere with the performance of other electronic devices, such as active implantable devices, e.g., pacemakers and defibrillators.  Testing should include radio frequency (RF) electromagnetic, low frequency magnetic, and conducted emissions.

## Immunity

EMC testing should also demonstrate that the device will perform as expected in the presence of other electrical and electronic devices or other sources of electromagnetic disturbance (EMD) in the intended environment of use (immunity).  The device should operate in an acceptable manner (few EMC standards require operation within specification) during and after exposure to various forms of electromagnetic disturbance.  Immunity testing should include:

- electrostatic discharge (ESD)

- radiated RF electromagnetic fields

- electrical fast transients and bursts

- surges

- conducted RF electromagnetic energy

- voltage dips, short interruptions, and voltage variations on power supply input lines

- low-frequency magnetic fields

- quasi-static electric fields.

# Acoustic Intensity

We recommend that you measure the sonometer’s acoustic output intensity.  See “Information for Manufacturers Seeking Marketing Clearance of Diagnostic Ultrasound Systems and Transducers” (http://www.fda.gov/cdrh/ode/ulstran.pdf) for the measurement methods we recommend.

# Non-Clinical Testing

We recommend that you provide:

- descriptions of algorithms for computing the measurements, e.g., broadband ultrasonic attenuation (BUA) and/or speed of sound (SOS)

- results of laboratory testing to establish the accuracy and precision of measured values

- bandwidth used in the calculation of attenuation versus frequency (for BUA measurement).  We recommend that you discuss the algorithm in detail (e.g., log spectral difference, centroid shift).

- specification of marker on the pulse waveform for measuring transit times for SOS computation (e.g., first zero crossing, envelope maximum, first zero crossing prior to the envelope maximum, threshold value)

- explanation of purposes and implementations of any filters applied to the received data

- description of procedures and testing performed to validate quality assurance tests

- analysis of the long-term stability of acoustic properties of test objects

- validation of your recommended method for cleaning and disinfecting the system between patients.

# Clinical Testing

We recommend you conduct clinical studies for generating reference databases (described below) and assessing reproducibility for all bone sonometers.  For these studies, and any other clinical studies you conduct in support of your 510(k), we recommend you provide the information described in this section.

## Study Design

We recommend you describe, in detail, the study design, including the number of sites, any randomization, and control arms.

## Subjects Enrolled

We recommend that you show the number of subjects enrolled in the study, segregated by:

- gender

- ethnic origin

- disease category

- age range

- principal inclusion and exclusion criteria

- unusual inclusion and exclusion criteria.

## Methodology

We recommend you include a detailed description of the methodology you used in gathering clinical data.

## Exclusion Criteria

For all of the clinical studies described in this guidance, we recommend the exclusion criteria listed below.

Current or recent use of bone-active drugs

We recommend that your exclusion criteria include:

- any use of bisphosphonates for more than 1 week

      - calcitonin, within 3 months

      - therapeutic doses (>1000 I.U. daily) of vitamin D, within 6 months

      - estrogens or selective estrogen receptor modulator (SERM) within 6 months

      - therapeutic doses (> 2 mg/day) fluoride within 3 years

      - drugs under research protocols within 2 years

      - unstudied or unapproved drugs.

Presence of metabolic bone disease

We recommend that your exclusion criteria include:

      - hyper- or hypo-parathyroidism within 5 years

      - osteitis deformans (Paget's disease of bone)

      - renal osteodystrophy

      - osteomalacia.

Other diseases or drugs

We recommend that your exclusion criteria include:

      - gastrointestinal malabsorption

      - liver disorders

      - chronic renal disease

      - unstabilized hyper- or hypothyroidism

      - hyper- or hypoadrenocorticism

      - concomitant use of oral corticosteroids, any dose

      - use of oral corticosteroids within past 6 months, if less than equivalent of 7.5 mg daily of prednisone, or use of a greater dosage, within the past year

      - use within past 6 months of anti-seizure drugs, barbiturates, or anticoagulants

      - stroke with total or partial paralysis with residual disability lasting more than 3 months.

### Diagnostic Measurements

### We recommend that you provide:

- rationales and formulas (such as linear combinations) for indexes that are derived from acoustic measurements

- description of the algorithms used to calculate the values presented (T-scores, Z-scores, etc.)

- copies of all pages of test results that are displayed either on a monitor or printed out as hard copy.

Reference Databases

Reference databases are used to calculate T and Z scores.  These scores facilitate diagnostic interpretation of measurements.  Diagnostic measurements on bone are often expressed as T-scores and Z-scores because relative (i.e., dimensionless) values like T-scores and Z-scores are easier for the physician to interpret than absolute values of measurements like bone density, broadband ultrasound attenuation, or speed of sound.

The T-score expresses a measurement relative to a normal reference database of healthy young Caucasian females.  The Z-score expresses a measurement relative to a normal reference database of an individual’s peer group.

Bone sonometers indicated for determining the possible presence of osteoporosis should incorporate a normative reference database for T-scores.  Bone sonometers for other indications should have both a T-score and Z-score reference database.

T-Score

The World Health Organization (WHO) defines osteoporosis, not by absolute values but by T-score (T-score < -2.5).  T-score, as used by WHO in their definition of osteoporosis, is the patient's measurement minus the average measurement of healthy young (between ages 20 and 39) Caucasian females divided by this group's standard deviation.  We believe reference databases of non-Caucasian females or males of any ethnicity are not appropriate for T-score calculation, unless you also develop new diagnostic values (i.e., cutoffs) for osteoporosis or osteopenia, and present the mean and standard deviation for each decade cohort.

We recommend that you develop a normative reference database of at least 300 healthy Caucasian women between the ages of 20 and 39 with at least 150 subjects in each decade.  We recommend that you use at least two geographically separated sites.

Z-Score

Z-score is the patient's measurement minus the average measurement for those in the same age decade, gender, and ethnicity divided by this group's standard deviation.  We believe databases for different ethnicity and genders are appropriate.  We recommend that you provide gender- and ethnic-specific normative databases, across all ages within the range of interest (generally from 20 to 80 or 90), to facilitate interpretation of bone measurements for evaluating secondary (non-age related) causes of low bone mass.

Databases for each ethnicity and gender should consist of measurement results from at least 50 subjects per institution per decade for the range of decades of interest.  We recommend that you use at least two geographically separated sites.  We recommend that you analyze the databases by decade and present the mean and standard deviation of the appropriate bone sonometer measurement for each decade.

If a clinical study is conducted prior to obtaining 510(k) clearance of the device, the study must be conducted under the Investigational Device Exemptions (IDE) regulation, 21 CFR Part 812.  We believe that generally, the bone sonometer addressed by this guidance document is a non-significant risk device, therefore the studies to obtain reference databases are subject to the abbreviated requirements of 21 CFR 812.2(b).  In addition to the requirements of section 21 CFR 812.2(b), sponsors of such trials must comply with the regulations governing institutional review boards (21 CFR Part 56) and informed consent (21 CFR Part 50).

Reproducibility Studies

These studies define the clinical precision of the device.  For monitoring indications, the expected change in a patient's measurement over a given time interval establishes the appropriate level of precision.  For all indications, we recommend that you establish reproducibility with respect to measurement-to-measurement, device-to-device, and operator-to-operator variability.

We recommend that you describe your variability studies and include methods, a line listing of data, and analyses.  We also recommend that you provide a brief rationale supporting the statistical validity of your analyses.  Your variability studies should include both phantom and clinical studies or an explanation why one or the other is not appropriate.

We recommend that you use at least 3 operators and 3 devices to estimate the repositioning, inter-operator, and device-to-device variability for ultrasound measurements.  We recommend you conduct this study on at least 6 pre-menopausal Caucasian women (20 to 45 years of age) and at least 6 post-menopausal Caucasian women (aged 55 to 80 years).  Subjects should be distributed throughout each of these age ranges.  We believe the most efficient design entails each subject being measured in duplicate (with repositioning between measurements) by each combination of operators and devices.

We recommend that you segregate your data into two age groups (below 45 and above 55) and analyze the results separately to obtain the variance components for operator, device, and repositioning, as well as the combined variability.  These in turn should be used to obtain the corresponding coefficients of variation.  If you demonstrate the results are combinable across age groups, then it is appropriate to pool all the data.  We also recommend that you present the combined variability as a fraction of the standard deviation of the young normal reference group or T-score standard deviation (TSD).  The T-score of an individual +/- 2 TSD provides an approximate 95% confidence interval on the measured T-score.

For monitoring indications, we recommend that you specify the frequency of monitoring and show that the overall precision of the device measurement allows for that frequency over a specified time interval.

While, in general, additional clinical studies will not be needed for most bone sonometer devices, FDA may recommend that you collect clinical data to demonstrate the effectiveness for a bone sonometer with any one of the following:

- indications for use dissimilar from indications for use of legally marketed bone sonometers of the same type

- designs dissimilar from designs previously cleared under a premarket notification

- new technology, i.e., technology different from that used in legally marketed bone sonometers.

FDA will always consider alternatives to clinical testing when the proposed alternatives are supported by an adequate scientific rationale.  If you conduct any additional clinical studies, we recommend you describe the design, conduct, and results of the clinical studies and follow the recommendations above for reference databases and reproducibility.

# Labeling

The premarket notification should include labeling in sufficient detail to satisfy the requirements of 21 CFR 807.87(e).  The following suggestions are aimed at assisting you in preparing labeling that satisfies the requirements of 21 CFR 807.87(e).

Indications for Use

We recommend that the indications for use specify the target population and include one or more of the following elements:

- determining the possible presence of osteoporosis and assessing fracture risk

- monitoring bone changes over time

- assessing non-age-related bone loss.

For example:

The device performs a quantitative ultrasound measurement of bone, the results of which can be used in conjunction with other clinical risk factors as an aid to the physician in the diagnosis of osteoporosis and medical conditions leading to reduced bone density, and ultimately in the assessment of fracture risk.

Contraindications

We recommend that the labeling list all contraindications, if any.

Warnings

We recommend that the labeling include warnings against using the device on skin with open wounds or signs of trauma, bleeding, or infection.

Directions for Use

As a prescription device, under 21 CFR 801.109, the device is exempt from having adequate directions for lay use.  Nevertheless, we recommend providing clear and concise instructions for professional users that delineate the technological features of the specific device and how the device is to be used on patients.  Instructions should encourage local/institutional training programs designed to familiarize users with the features of the device and how to use it in a safe and effective manner.

We recommend that the instructions also include:

- a protocol for cleaning and disinfecting between patients, including a list of any accessories needed to clean and disinfect the device and the chemical name of the recommended germicide

- methods for preparing the skin (e.g., cleaning with alcohol pads)

- methods for ensuring good coupling between the ultrasound transducer and the skin surface (e.g., use of coupling gel for dry systems and eliminating bubbles for wet systems)

- recommendations for quality assurance methods

- maximum period of time over which phantoms or test objects may be assumed reliable.

We also recommend that the instructions emphasize the need to use the same healthy young Caucasian female database for the calculation of every patient's T-score, regardless of gender and ethnicity and advise against the use of gender- and ethnic-specific normative data for fracture risk assessment.

Device Description

We recommend that labeling identify the:

- anatomical site(s) scanned

- major components of system

- ultrasound parameter(s) being used

- scanning method.

Precautions

We recommend that the labeling:

- list possible adverse events associated with the use of this device

- caution against the use of different calibration phantoms in monitoring.

Individualization of Treatment

We recommend that labeling advise users that the device is not limited to use in Caucasian women.  For example, “This device may be used to determine the possible presence of osteoporosis and to estimate fracture risk in men and in Caucasian and non-Caucasian women.  You should use the same young normal Caucasian female database to calculate the T-score for all patients, regardless of gender and ethnicity.”

We also recommend that labeling address gender and ethnicity in monitoring bone changes over time, for example, “There are similarly no gender or ethnic restrictions on its use in monitoring bone changes over time or assessing non-age-related bone loss, although for the latter assessment, you should use databases that are matched to the patient's age, gender, and ethnicity.”

Clinical Studies

Information in this section should include a brief summary, i.e., a self-contained description of the design, conduct, and results of the clinical studies for generation of the reference database and assessment of precision.  The summary should include the information described below.

Study Design

We recommend that labeling briefly describe the study design, in particular, by addressing the number of sites, any randomization, control arms, etc.

Subjects Enrolled

We recommend that your labeling show the number of subjects enrolled in the study, segregated by:

- gender

- ethnic origin

- disease category

- age range

- principal and unusual inclusion and exclusion criteria.

Methodology

We recommend that labeling include a concise statement of the methodology used in gathering the effectiveness and safety data.

User Manual, Technical Manual, and Other Labeling

Physician Information

We recommend that the physician labeling discuss:

- osteoporosis

- different methods of measuring bone strength

- error and precision of these measurements

- interpretation of results (e.g., differences in patient classification are possible for two different methods).

- precision of the device

- appropriate monitoring interval, if used for monitoring patients under treatment

- references to publications on the differences of results from the various bone densitometers and sonometers.

We also recommend that the physician labeling describe a quality assurance program for the user that assures the continued proper operation and calibration of the device.  It should include the conditions (e.g., duration of use, component replacement, quality control and calibration procedures) that should be managed during normal use to maintain the safety and effectiveness of the device.

Patient Labeling

We recommend that the patient labeling explain:

- osteoporosis

- why the disease is of concern to women

- the measurement of bone using ultrasound.

We also recommend that the patient labeling explain that ultrasound measurements are different from x-ray bone densitometry (e.g., How does it work?  What will it tell me?  Why is this important?  What other tests are available and how does your device differ?  What should the patient know?).  (See also Guidance on Medical Device Patient Labeling, http://www.fda.gov/cdrh/ohip/guidance/1128.html.)

# Footnotes

Refer to http://www.fda.gov/cdrh/ode/indicate.html for the recommended format.

If FDA makes a substantial equivalence determination based on acceptance criteria, the subject device should be tested and shown to meet these acceptance criteria before being introduced into interstate commerce.  If the finished device does not meet the acceptance criteria and, thus, differs from the device described in the cleared 510(k), FDA recommends that submitters apply the same criteria used to assess modifications to legally marketed devices (21 CFR 807.81(a)(3)) to determine whether marketing of the finished device requires clearance of a new 510(k).

See Required Elements for a Declaration of Conformity to a Recognized Standard (Screening Checklist for All Premarket Notification [510(K)] Submissions), http://www.fda.gov/cdrh/ode/reqrecstand.html.

Under the authority of 21 CFR 1002.50(b), diagnostic ultrasound products are exempt from the reporting requirements of 21 CFR 1002.10 and 1002.12, if a 510(k) has been submitted.  See also “Appendix D Exemption from Reporting Under 21 CFR 1002” in Information for Manufacturers Seeking Marketing Clearance of Diagnostic Ultrasound Systems and Transducers, September 30, 1997 (http://www.fda.gov/cdrh/ode/ulstran.pdf ).  However, these devices remain subject to the reporting requirements identified in 21 CFR 1002.20 Reporting of accidental radiation occurrences, 21 CFR Part 1003  Notification of Defects, and 21 CFR Part 1004 Repurchase, repairs, or replacement of electronic products.

http://www.fda.gov/cdrh/mdufma/guidance/1215.html.

Assessment of Fracture Risk and Its Application to Screening for Postmenopausal Osteoporosis, World Health Organization, WHO Technical Report Series 843, 1994.

The optimal range is 300 to 500 subjects according to Martin et al., Normal Values in Clinical Chemistry, Marcel Dekker, 1975, pages 51-52.

See http://www.fda.gov/oc/ohrt/irbs/devices.html#risk.

Although final labeling is not required for 510(k) clearance, final labeling must comply with the requirements of 21 CFR 801 before a medical device is introduced into interstate commerce.  In addition, final labeling for prescription medical devices must comply with 21 CFR 801.109.  Labeling recommendations in this guidance are consistent with the requirements of part 801.