D-Zero Detector Central Fiber Tracker (CFT) Axial Project

Readout Electronics

D0 Central Tracker Trigger (CTT) Mixer System

Backplane Description & Power Distribution Specification

Date: July 4, 2000

Revision Date: March 27th, 2003

Stefano Rapisarda, Neal Wilcer, John Anderson

Document # ESE-D0-000704

Table of Contents

1	Introduction	2

2	Overview	3

Figure 1, Picture of the mixer subrack.	3

Figure 2, Mixer System backplane	5

3	Physical Attributes of the Backplane	6

3.1	Dimensions	6

3.2	Layer Description	6

Figure 3, Mixer Backplane Stackup	6

3.3	Connectors	6

3.4	Backplane to chassis mounting	6

Figure 4. Front View of Mixer System Backplane	7

Figure 5. Rear View of Mixer System Backplane	7

3.5	DC power connections	7

Figure 6. Backplane and Bus Bar Assembly	8

Figure 7. Picture of Mixer Crate with Power Supplies.	8

Figure 8. Picture of Bus Bar to Vicor Power Supply Connections.	9

Figure 9. Picture of Mixer Subrack AC Terminal Block Connections	9

3.6	AC Distribution	10

Figure 10, Schematic diagram of AC distribution box.	10

Figure 11. Picture of Backplane Bus Bar Assembly	11

Figure 12. Bus Bar Stud	11

Figure 13. Backplane to Bus Bar Assembly	12

Figure 14. Picture of section 1 of Bus Bar Assembly	12

Figure 15. Mechanical Drawing of section 1 of Bus Bar Assembly	13

Figure 16. Mounting Hardware for Bus Bar Assembly sections 1 and 2	13

Figure 17. Picture of section 2 of Bus Bar Assembly	14

Figure 18. Mechanical Drawing of section 2 of Bus Bar Assembly	14

Figure 19. Picture of section 3 of Bus Bar Assembly	15

Figure 20. Mechanical Drawing of section 3 of Bus Bar Assembly	15

3.7	Torque Specs	16

4	Power source	16

5	Power requirements	16

6	Analysis of Current Densities	16

Figure 21. Analysis of Current Densities; Areas of Interest	17

Table 2. Single Leg of Bus Bar section 3 Cross Section	18

Table 5. Bus Bar section 2 Cross Section	18

Table 6. Bus Bar Section 2 to Bus Bar section 1	19

Table 7. Bus Bar Section 1 Cross Section	19

Table 8. Bus Bar section 1 to Spacer (also spacer to Backplane)	19

Table 9. Backplane to Power/GND Plane (thru holes, vias)	19

Figure 22. Side View of Through Hole	20

Table 10. Power/GND Plane to Connector.	23

7	Appendix A - Mixer System Backplane connectors pinout	26

Table 11, Pin/Signal assignments for the backplane top section connector	26

Table 12, Pin/Signal assignments for the backplane bottom section connector.	27

8	References	28

# Introduction

This document describes the mixer system backplane. The mixer system is part of the readout electronics of the D-Zero detector Central Fiber Tracker (CFT) at Fermilab. More information on the experiments performed at Fermilab is available on the laboratory web page:

http://www.fnal.gov/

More information on the D0 Detector is available on:

http://www-d0.fnal.gov/

The designers welcome suggestions and corrections [Ref. 4], which can be addressed directly to the engineer responsible of the project. Contact information is available on the Electronics System Engineering (ESE) web page:

http://www-ese.fnal.gov/

More information and documentation on the Mixer Project are available on:

http://www-ese.fnal.gov/D0_CTT_MIXER/

# Overview

The mixer system [Ref 4] consists of a 21 slots subrack with a custom backplane. Slot 1 is occupied by a custom subrack controller, slot 2 to slot 21 host twenty mixer boards. The mixer system is partitioned in five subsystems of four boards each. The mixer system custom backplane was developed as an interconnect and power distribution bus for the mixer board set. It resembles the 6U VME J1/J2 backplane form, but does not necessarily adhere to VME backplane specifications [Ref. 2].

Figure 1, Picture of the mixer subrack.

Figure 2 shows a sketch of the mixer system backplane electrical interconnections. The backplane schematic is available in a separate document [Ref.4].

The backplane support the following interfaces:

Subrack controller general-purpose bus interface.

Used for configuration download and for remote access to the mixer boards. More information about the Subrack Controller is provided in separate documents [Ref.4, 5e].

Subrack controller slow monitoring serial bus interface.

During data acquisition, due to the proximity of the mixer system to detector components, noise considerations suggest to avoid accessing the mixer system over the "general purpose bus interface". The slow monitoring serial bus provides a way to readback status information from the mixer boards.

Subrack controller 1553 interface.

The subrack controller is remotely accessible through a MIL-STD-1553 serial bus. The MIL-STD-1553 is a networking standard used for integration of military platforms [Ref. 6]. The backplane host the two 1553 triaxial connectors needed by the subrack controller to implement the 1553 interface.

Board to board buses.

Mixer boards need to exchange part of the data they receive over the input links with the adjacent boards belonging to the same mixer subsystem (four boards). This is accomplished with two parallel buses 41 bit wide, one to the next backplane slot to the left and one to the next backplane slot to the right.

Timing signals distribution

Out of a mixer subsystem (group of four boards) the leftmost is used as "master" for distribution of global clock and global frame synchronization (SYNC) signals to the other three boards (slaves). The clock is distributed over three (one for each slave board) point-to-point LVDS connections. The SYNC is distributed over three point-to-point LVTTL connections.

Power supply.

A mixer board is powered through a backplane connection to the 3.3Volts subrack power supply. A small brick power supply is used to provide 5 Volts power needed by the subrack controller (slot 1). Two green LEDs, one for each supply voltage are ON when the backplane is powered. Three polyfuses are used as over-current protection. Each mixer board has its own over-current/over-voltage protection.

Figure 2, Mixer System backplane

# Physical Attributes of the Backplane

## Dimensions

The mixer system backplane is 16.910” [429.514mm] wide by 10.310” [261.874mm] high. The thickness is .192” [4.877mm], it was determined using a 8:1 aspect ratio of backplane thickness compared to the smallest thru hole diameter (0.024” [0.6096mm]). The 8:1 aspect ratio is the maximum allowable to assure even plating in thru holes.

## Layer Description

The backplane is fourteen layers thick. Top and bottom layers are 1 oz. copper (thickness: .0014" [.0356mm]) pours attached to digital ground. Plated areas provide the option to connect these copper pours to the earth ground (chassis). There are five inner signal layers, two +3.3V layers, and five GND layers. Layer IN2 and IN4 have 100Ω differential impedance and are used for clock signals with the differential traces stacked in a broadside-coupled stripline fashion. The remainder of the signal layers has 50 Ω impedance and the signal traces are arranged in symmetrical stripline fashion. Figure 3 shows the layers stackup.

Figure 3, Mixer Backplane Stackup

## Connectors

Backplane connectors used for power distribution to the Mixer Boards and signal lines are a combination of the A and B type 2mm hard metric [Ref. 7]. These are press fit connectors. Each slot is keyed in two places for guiding the mixer boards into the mixer system backplane.

## Backplane to chassis mounting

Standard VME spec backplane mounting is used. Figure 4 shows the front side of the backplane, Figure 5 shows the rear of the backplane.

Figure 4. Front View of Mixer System Backplane

Figure 5. Rear View of Mixer System Backplane

## DC power connections

Power connections meet the “Electrical Design Standards for Electronics to be used in Experiment Apparatus at Fermilab”, rev 4.0, dated October 16 1998 [Ref.3]. The +3.3V supply and the GND return are attached to the backplane via screws attached to the plated copper bus bar assemblies. See Figure 6, Figure 7 and Figure 8.

Figure 6. Backplane and Bus Bar Assembly

Figure 7. Picture of Mixer Crate with Power Supplies.

Figure 8. Picture of Bus Bar to Vicor Power Supply Connections.

Figure 9. Picture of Mixer Subrack AC Terminal Block Connections

## AC Distribution

BiRa Systems [Ref. 8] built the AC distribution box to D0 specifications.  It is a 3U x 19” rack mount unit with circuit breakers on the front panel and AC sockets on the back side.  Three phase AC is supplied to the box, and is then distributed as two phase AC to the power supplies.  A dual receptacle single phase AC connection is also available.  Each AC receptacle is individually controlled/protected with a circuit breaker.  Additionally, the distribution box contains a relay to control all of the AC outputs; this relay is connected to the interlock system and requires 3-24 VDC @ 10mA to enable the AC outputs. The three phases are evenly distributed to nine two-phase outlets to allow for load balancing.

The schematic diagram of the AC distribution box is shown below:

Figure 10, Schematic diagram of AC distribution box.

Figure 11. Picture of Backplane Bus Bar Assembly

Figure 12 shows components of the fastener used to attach the backplane to the bus bar assemblies. The +5V supply is used only for the crate controller (slot# 1) is attached via a molex style connector.

Figure 12. Bus Bar Stud

Figure 13 shows a bottom view of the backplane, fasteners, and section 1 (refer to Figure 6) of the bus bar assembly.

Figure 13. Backplane to Bus Bar Assembly

Figure 14 and Figure 15 show section 1 (refer to Figure 6) of the tin plated copper bus bar assemblies used for the +3.3V and GND.

Figure 14. Picture of section 1 of Bus Bar Assembly

Figure 15. Mechanical Drawing of section 1 of Bus Bar Assembly

Figure 16 shows the hardware for sections 1 and 2 (refer to Figure 6) of the bus bar assemblies used for the +3.3V and GND.

Figure 16. Mounting Hardware for Bus Bar Assembly sections 1 and 2

Figure 17 and Figure 18 show section 2 (refer to Figure 6) of the bus bar assemblies used for the +3.3V and GND.

Figure 17. Picture of section 2 of Bus Bar Assembly

Figure 18. Mechanical Drawing of section 2 of Bus Bar Assembly

Figure 19 and Figure 20 show section 3 (refer to Figure 6) of the bus bar assemblies used for the +3.3V and GND.

Figure 19. Picture of section 3 of Bus Bar Assembly

Figure 20. Mechanical Drawing of section 3 of Bus Bar Assembly

Section 3 of the Bus Bar assembly will be attached to the Vicor power supply using 8-32 silicon bronze nuts and lock washers.

## Torque Specs

10-32 Silicon Bronze screws : 24 in.-lb [0.27651 Kg.-m].

5/16”-18 Phosphor Bronze PEM bolts:

125 in.-lb [1.440156 Kg.-m]. (torque spec for silicon bronze)

1/4”-20 Phosphor Bronze PEM bolts:

70 in.-lb. [0.806487 Kg.-m](torque spec for silicon bronze)

8-32 Silicon Bronze screws: 21 in.-lb [0.241946 Kg.-m].

8-32 Silicon Bronze nuts on Vicor Supply: 10 in.-lb (as per Vicor spec)

# Power source

The +3.3V is supplied by an Vicor Mini PFC power supply [Ref. 20] using three 80 amp modules in parallel for a total of 240 amps @ +3.3V. +5V will be supplied by a small “brick” style supply.

# Power requirements

As of this writing, estimated power requirements are 240 amps @ +3.3V, 2 amps @ +5V. Gnd is common and requires 240 amps.

# Analysis of Current Densities

The areas of interest are shown below in Figure 21. Each area should follow the “no more than 1000 amps/sq.in.” rule. The charts on the following pages breakdown each of these areas. For each chart, areas that conform to the 1000 amps/sq.in. [155 amps/sq cm] are colored green. Areas that don’t are colored yellow. Special formulas are listed where appropriate.

Figure 21. Analysis of Current Densities; Areas of Interest

| Supply Name | From | To | Type | Amps | Area (sq. in.) [sq.cm] | # in parallel | Net Amps/sq in. [Amps/sq.cm] | Measurement Notes for Area |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| +3.3V, GND | Power supply | Power Supply Bus Bar section 3 for +3.3v and GND. | Compression Flat | 80 | 0.1884248 [1.215641] | 1 | 424 [65] | Contact area of nut on power supply (.25 sq.in. [1.61 sq.cm]) – area consumed by power supply stud (.06157522 sq.in. [.3972587 sq.cm]) |

Table 1. Power Supply to Bus Bar section 3

| Supply Name | From | To | Type | Amps | Area (sq in.) [sq.cm] | # in parallel | Net Amps/sq in. [Amps/sq.cm] | Measurement Notes for Area |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| +3.3V, GND | Single leg of section 3 Bus Bar Start | Bus Bar section 3 | Cross Section | 80 | 0.094 [0.60645] | 1 | 851 [131] | Used cross sectional area of Single Leg of Bus Bar section 3. Cross section area (.188"x.500") [1.213cm x 3.226cm]) |

Table 2. Single Leg of Bus Bar section 3 Cross Section

| Supply Name | From | To | Type | Amps | Area (sq in.) [sq.cm] | # in parallel | Net Amps/sq in. [Amps/sq.cm] | Measurement Notes for Area |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| +3.3V, GND | Bus Bar section 3 start | Bus Bar section 3 finish | Cross Section | 240 | 0.419804 [2.7084] | 1 | 571 [88] | Used main area of bus bar section 3 (2.549”x.188”) [6.474 cm x .4775 cm] – area consumed by stud (".188x.316”) [.4775 cm x .8026 cm] |

Table 3. Bus Bar section 3

| Supply Name | From | To | Type | Amps | Area (sq in.) [sq.cm] | # in parallel | Net Amps/sq in. [Amps/sq.cm] | Measurement Notes for Area |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| +3.3V, GND | Surface of Bus Bar section 3 | Surface of Bus Bar section 2 | Compression | 240 | 0.740425 [4.7769] | 1 | 324 [50] | Used contact area of Bus Bar section 2 (.625”x1.357”) [1.5875cmx3.4468cm] - areas consumed by 2 mounting holes (.0784 sq. in.+.0293 sq. in.) [.5058 sq. cm + .189 sq. cm.] |

Table 4. Bus Bar section 3 To Bus Bar section 2

| Supply Name | From | To | Type | Amps | Area (sq in.) [sq.cm] | # in parallel | Net Amps/sq in. [Amps/sq.cm] | Measurement Notes for Area |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| +3.3V, GND | Bus Bar section 2 cross section start | Bus Bar section 2 cross section finish | Cross Section | 240 | 0.38125 [2.45967] | 1 | 629 [97] | Cross sectional area of bus bar section 2 (.610”x.625”) [1.5494 cm x 1.5875 cm] |

Table 5. Bus Bar section 2 Cross Section

| Supply Name | From | To | Type | Amps | Area (sq in.) [sq.cm] | # in parallel | Net Amps/sq in. [Amps/sq.cm] | Measurement Notes for Area |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| +3.3V, GND | Surface of Bus Bar section 2 | Surface of Bus Bar section 1 | Compression | 240 | 0.64505 [4.1616] | 1 | 372 [57] | Area of contact on Bus Bar section 1 (.650”x1.11”) [1.651cm x 2.8194cm] – area of stud hole (.07645 sq. in.) [.4932 sq. cm.] |

Table 6. Bus Bar Section 2 to Bus Bar section 1

| Supply Name | From | To | Type | Amps | Area (sq in.) [sq.cm] | # in parallel | Net Amps/sq in. [Amps/sq.cm] | Measurement Notes for Area |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| +3.3V, GND | Bus Bar Section 1 Start | Bus Bar Section 1 End | Cross Section | 240 | 0.2415 [1.5581] | 1 | 993 [153] | Cross sectional area of bus bar section 1 (.650" x .525") [1.651cm x 1.334cm]) – height of busbar x width of backplane power stud (.525" x .190") [1.334cm x .483cm]). |

Table 7. Bus Bar Section 1 Cross Section

| Supply Name | From | To | Type | Amps | Area (sq in.) [sq.cm] | # in parallel | Net Amps/sq in. [Amps/sq.cm] | Measurement Notes for Area |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| +3.3V, GND | Backplane Bus Bar | Spacer (standoff) | Compression Flat | 240 | 0.16649 [1.074127] | 10 | 144.1 [22.3] | Same spacer as +3.3V. Outside area of spacer ((r2 where 2r = d =.500” [1.27cm]) – Inside area of spacer ((r2 where 2r = d=.195” [.495cm]) |

Table 8. Bus Bar section 1 to Spacer (also spacer to Backplane)

| Supply Name | From | To | Type | Amps | Area (sq in.) [sq.cm] | # in parallel | Net Amps/sq in. [Amps/sq.cm] | Measurement Notes for Area |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| +3.3V, GND | Backplane surface | Internal power plane | PC board vias | 240 | 0.001385 [0.008935] | 10 | 17322.9 [2685.1] | Construction of thru holes as per Figure 22. The "cross sectional area of a via" formula shown below was used. Plating thickness is .001" [.0254mm]. Finished hole size is .190" [4.83mm] for thru hole, and .024" [.61mm] for vias. 10 vias/thru hole |

Table 9. Backplane to Power/GND Plane (thru holes, vias)

Cross Sectional Area of a Via

runplated – radius of unplated thru hole (via)

rplated – radius of plated thru hole (via)

Note that all of the current densities in thru holes and vias are beyond the 1000 amp/sq.in. [155 amps/sq cm] limit. The actual issue here is how much power is lost in these vias, and will the surrounding area be sufficient to dissipate the heat. Special bus bar mounting screw thru holes were created in the backplane to increase the number of current paths. See Figure 22.

Figure 22. Side View of Through Hole

Power loss through the vias and thru holes is calculated using the formula below to determine the resistance of a single via and a single screw hole. Once these resistances are established, the parallel resistance of all of the vias and screw holes is determined. The formula

is then used to determine how much power would be dissipated in the area. We used .100” [2.54mm] for the variable x below. The .100” [2.54mm] is worst case. A temp value of 70(C (158(F) was used for these calculations.

Resistance (() of a via or thru hole

See Figure 23 for description of variables

(    Bulk resistivity of copper   6.787(10-7 ((*in.)

((  Thermal coefficient of resistance  .0039

runplated – radius of unplated thru hole or via (in.)

rplated – radius of plated thru hole or via (in.)

x – separation between contact points (in.)

temp – temperature ((C)

Figure 23. Side View of a Via

Resistance of the .190” [4.826mm] dia. thruhole

Resistance of .024" [0.610mm] dia via

Rtotal of +3.3V and GND thru holes and vias

Determining power loss in the +3.3V and GND vias and thru holes (Note: in all formulas used, GND and +3.3V calculations are the same.)

Following the “power density of a circuit card should be limited to a maximum of  ½ watt per square inch”1 rule,  the areas of each of the vias and thru holes are calculated, the total area for thru holes and vias is found, and the wattage per square inch determined. For these measurements, the formula for surface area of a cylinder was used. d is the diameter of the thru hole or via, and x is the distance from the backplane outer layer to the plane layer (.100” [2.54mm] worst case).

For the .190” [4.826mm] dia. thru hole

For the .024” [0.610mm] dia. thru hole

As shown before, each thru hole has 10 vias attached to it, and there are 10 thru holes per power supply. So for each supply,

+3.3V and GND total area

Using the power loss figures we calculated earlier, determine the power loss for unit of area (power density) for the vias and thru holes for each supply.

+3.3V and GND thru holes and vias

All of these power densities fall well within .500 Watt/sq.in. [.0775Watt/sq.cm] specification.

Table 10. Power/GND Plane to Connector.

| Supply Name | From | To | Type | Amps | Area (sq in.) | # in parallel | Net Amps/Sq in. | Measurement Notes for Area |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Supply Name | From | To | Type | Amps | Area (sq in.) | # in parallel | Net Amps/Sq in. | Measurement Notes for Area |
| +3.3V | Thru holes and vias | Connector pins | Power/GND plane or copper pour. Sheet current | 240 | 0.045657 | 2 | 2628.29358 | Total current divided into cross sectional area of copper sheet (.0027" thick), over entire width of backplane (16.91"). Flow is vertical |
| GND | Thru holes and vias | Connector pins | Power/GND plane or copper pour. Sheet current | 240 | 0.045657 | 5 | 1051.317432 | Total current divided into cross sectional area of copper sheet (.0027" thick), over entire width of backplane (16.91"). Flow is vertical |

Like the vias and thru holes listed in the previous chart, the power and gnd plane layers do not conform to the 1000 amps/sq.in. rule. Instead, the amount of power dissipated in the power and gnd layers will be calculated to determine if the backplane can handle the heat loss. The formula used for power dissipation is W=I2R, where R is the resistance of the plane layers.

The construction of the backplane is such that each bus bar thru hole feeds the equivalent of two slots (10 thru holes per bus bar for 20 slots). For the resistance calculations below, d0 is one bus bar thru hole and di is a connector thru hole. Since each slot has multiple di, we define the equivalent diameter value as:

Where n is the total number of connector pins in two slots that carry a particular voltage or gnd. This is defined as a section.

The value used for separation between contact points (x) is always the worst case instance.

Once the resistance for a section is calculated, the resistance for the entire plane is calculated by dividing the single section resistance by the total number of bus bar thru holes, in this case ten. This number is again divided by the number of times the plane occurs in the backplane to achieve the total resistance of the planes used for each voltage.

Calculating resistance (() of a power or ground plane with one source and multiple destinations

(    Bulk resistivity of copper   6.787(10-7 (in.* ()

((  Thermal coefficient of resistance  .0039

d0 – diameter of 0th contact point (in.)

di – diameter of ind contact point (in.)

t – thickness of plane (in.)

x – separation between contact points (in.)

temp – temperature ((C)

Resistance calculations for +3.3V plane(s)

Resistance calculations for GND plane(s)

Power loss in +3.3V plane(s)

Power loss in GND plane(s)

Following the “power density of a circuit card should be limited to a maximum of  ½ watt per square inch”1 rule,  the areas of the planes are calculated, and the wattage per square inch determined.

For the +3.3V plane

For the GND plane

Worse case of the total power density is the watts/sq.in. sum of the 2 plane layers power densities since the 2 layers are stacked on top of each other.

This value of 91.265 mWatt/sq.in. falls well below the 500 mWatt/sq.in. (77.5mWatt/sq.cm) guideline.

# Appendix A - Mixer System Backplane connectors pinout

| Pin # | Row ‘z’ (bottom shield) | Row ‘a’ | Row ‘b’ | Row ‘c’ | Row ‘d’ | Row ‘e’ | Row ‘f’ (top shield) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | GND | INLFT2 | GND | INLFT1 | GND | OUTRT1 |  |
| 2 |  | GND | INLFT3 | GND | OUTRT2 | GND | GND |
| 3 | GND | INLFT4 | GND | OUTRT3 | GND | OUTRT4 |  |
| 4 |  | GND | INLFT5 | GND | OUTRT5 | GND | GND |
| 5 | GND | INLFT7 | GND | INLFT6 | GND | OUTRT6 |  |
| 6 |  | GND | INLFT8 | GND | OUTRT7 | GND | GND |
| 7 | GND | INLFT9 | GND | OUTRT8 | GND | OUTRT9 |  |
| 8 |  | GND | INLFT10 | GND | OUTRT10 | GND | GND |
| 9 | GND | INLFT12 | GND | INLFT11 | GND | OUTRT11 |  |
| 10 |  | GND | INLFT13 | GND | OUTRT12 | GND | GND |
| 11 | GND | INLFT14 | GND | OUTRT13 | GND | OUTRT14 |  |
|  | polarizing and guiding peg |
| 12 | GND | GND | INLFT15 | GND | OUTRT15 | GND |  |
| 13 |  | INLFT17 | GND | INLFT16 | GND | OUTRT16 | GND |
| 14 | GND | GND | INLFT18 | GND | OUTRT17 | GND |  |
| 15 |  | INLFT19 | GND | OUTRT18 | GND | OUTRT19 | GND |
| 16 | GND | GND | INLFT20 | GND | OUTRT20 | GND |  |
| 17 |  | INLFT22 | GND | INLFT21 | GND | OUTRT21 | GND |
| 18 | GND | GND | INLFT23 | GND | OUTRT22 | GND |  |
| 19 |  | INLFT24 | GND | OUTRT23 | GND | OUTRT24 | GND |
| 20 | GND | GND | INLFT25 | GND | OUTRT25 | GND |  |
| 21 |  | INLFT27 | GND | INLFT26 | GND | OUTRT26 | GND |
| 22 | GND | GND | INLFT28 | GND | OUTRT27 | GND |  |
| 1 |  | INLFT29 | GND | OUTRT28 | GND | OUTRT29 | GND |
| 2 | GND | GND | INLFT30 | GND | OUTRT30 | GND |  |
| 3 |  | INLFT32 | GND | INLFT31 | GND | OUTRT31 | GND |
| 4 | GND | GND | INLFT33 | GND | OUTRT32 | GND |  |
| 5 |  | INLFT34 | GND | OUTRT33 | GND | OUTRT34 | GND |
| 6 | GND | GND | INLFT35 | GND | OUTRT35 | GND |  |
| 7 |  | INLFT37 | GND | INLFT36 | GND | OUTRT36 | GND |
| 8 | GND | GND | INLFT38 | GND | OUTRT37 | GND |  |
| 9 |  | INLFT39 | GND | OUTRT38 | GND | OUTRT39 | GND |
| 10 | GND | GND | INLFT40 | GND | OUTRT40 | GND |  |
| 11 |  | GND | GND | INLFT41 | GND | OUTRT41 | GND |
| 12 | GND | MOD_ID4 | MOD_ID3 | MOD_ID2 | MOD_ID1 | MOD_ID0 |  |
| 13 |  | +IN_CLK | GND | +OUT_CLK0 | +OUT_CLK1 | +OUT_CLK2 | GND |
| 14 | GND | -IN_CLK | GND | -OUT_CLK0 | -OUT_CLK1 | -OUT_CLK2 |  |
| 15 |  | GND | GND | GND | GND | GND | GND |
| 16 | GND | STAT_CLK | GND | STAT_RST | GND | STAT |  |
| 17 |  | Power 3.3V | Power 3.3V | Power 3.3V | Power 3.3V | Power 3.3V | GND |
| 18 | GND | Power 3.3V | Power 3.3V | Power 3.3V | Power 3.3V | Power 3.3V |  |
| 19 |  | Power 3.3V | Power 3.3V | Power 3.3V | Power 3.3V | Power 3.3V | GND |

Table 11, Pin/Signal assignments for the backplane top section connector

| Pin # | Row ‘z’ (bottom shield) | Row ‘a’ | Row ‘b’ | Row ‘c’ | Row ‘d’ | Row ‘e’ | Row ‘f’ (top shield) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | GND | VADDR7 | VWRITE | VDAT6 | VDAT5 | VDAT2 |  |
| 2 |  | VADDR15 | *SYSRST | VDAT7 | VDAT0 | VDAT3 | GND |
| 3 | GND | VADDR12 | VAS | VSYSCLK | VDAT1 | VDAT4 |  |
| 4 |  | VADDR9 | *VDTACK | *VBERR | VADDR13 | VADDR5 | GND |
| 5 | GND | VADDR11 | VADDR2 | VADDR10 | VADDR0 | VADDR3 |  |
| 6 |  | VADDR8 | VADDR6 | VADDR1 | VADDR14 | VADDR4 | GND |
| 7 | GND | Power 3.3V | Power 3.3V | Power 3.3V | Power 3.3V | Power 3.3V |  |
| 8 |  | Power 3.3V | Power 3.3V | Power 3.3V | Power 3.3V | Power 3.3V | GND |
| 9 | GND | OUTLFT2 | GND | OUTLFT1 | GND | INRT1 |  |
| 10 |  | GND | OUTLFT3 | GND | INRT2 | GND | GND |
| 11 | GND | OUTLFT4 | GND | INRT3 | GND | INRT4 |  |
|  | polarizing and guiding peg |
| 12 | GND | GND | OUTLFT5 | GND | INRT5 | GND |  |
| 13 |  | OUTLFT7 | GND | OUTLFT6 | GND | INRT6 | GND |
| 14 | GND | GND | OUTLFT8 | GND | INRT7 | GND |  |
| 15 |  | OUTLFT9 | GND | INRT8 | GND | INRT9 | GND |
| 16 | GND | GND | OUTLFT10 | GND | INRT10 | GND |  |
| 17 |  | OUTLFT12 | GND | OUTLFT11 | GND | INRT11 | GND |
| 18 | GND | GND | OUTLFT13 | GND | INRT12 | GND |  |
| 19 |  | OUTLFT14 | GND | INRT13 | GND | INRT14 | GND |
| 20 | GND | GND | OUTLFT15 | GND | INRT15 | GND |  |
| 21 |  | OUTLFT17 | GND | OUTLFT16 | GND | INRT16 | GND |
| 22 | GND | GND | OUTLFT18 | GND | INRT17 | GND |  |
| 1 |  | OUTLFT19 | GND | INRT18 | GND | INRT19 | GND |
| 2 | GND | GND | OUTLFT20 | GND | INRT20 | GND |  |
| 3 |  | OUTLFT22 | GND | OUTLFT21 | GND | INRT21 | GND |
| 4 | GND | GND | OUTLFT23 | GND | INRT22 | GND |  |
| 5 |  | OUTLFT24 | GND | INRT23 | GND | INRT24 | GND |
| 6 | GND | GND | OUTLFT25 | GND | INRT25 | GND |  |
| 7 |  | OUTLFT27 | GND | OUTLFT26 | GND | INRT26 | GND |
| 8 | GND | GND | OUTLFT28 | GND | INRT27 | GND |  |
| 9 |  | OUTLFT29 | GND | INRT28 | GND | INRT29 | GND |
| 10 | GND | GND | OUTLFT30 | GND | INRT30 | GND |  |
| 11 |  | OUTLFT32 | GND | OUTLFT31 | GND | INRT31 | GND |
| 12 | GND | GND | OUTLFT33 | GND | INRT32 | GND |  |
| 13 |  | OUTLFT34 | GND | INRT33 | GND | INRT34 | GND |
| 14 | GND | GND | OUTLFT35 | GND | INRT35 | GND |  |
| 15 |  | OUTLFT37 | GND | OUTLFT36 | GND | INRT36 | GND |
| 16 | GND | GND | OUTLFT38 | GND | INRT37 | GND |  |
| 17 |  | OUTLFT39 | GND | INRT38 | GND | INRT39 | GND |
| 18 | GND | GND | OUTLFT40 | GND | INRT40 | GND |  |
| 19 |  | GND | GND | OUTLFT41 | GND | INRT41 | GND |

Table 12, Pin/Signal assignments for the backplane bottom section connector.

# References

[1] Howard Johnson, Martin Graham, 
“High Speed Digital Design”, 1993, Prentice Hall PTR. ISBN 0-13-395724-1.

[2] Wade D.Peterson, VMEbus International Trade Association, “The VMEbus Handbook”, Fourth edition 1997. ISBN 1-885731-08-6.

[3] “Electrical Design Standards for Electronics to be used in Experiment Apparatus at Fermilab”, Rev. 4.0 10/16/98

[4] D0 CTT Mixer System documentation. Information available on the Internet:
http://www-ese.fnal.gov/D0_CTT_Mixer/

[5] Jamieson T. Olsen,

  - "Digital Front End Mother Board register map", Engineering Note 2000-01-28a.

  - "DFE and Mixer Board Registers ", Engineering Note 2000-01-28a.

  - "DFE backplane address/data bus protocol", Engineering Note 2000-03-28a.

  - "DFE slow monitoring custom serial bus specification", 
 Engineering Note 2000-01-31a.

  - Digital Front End Controller (DFEc) Board:
http://d0server1.fnal.gov/users/jamieson/www/projects/dfec/index.html

Information available on the internet: http://d0server1.fnal.gov/users/jamieson/www/index.html

[6] MIL-STD-1553 is a military standard that defines the electrical and protocol characteristics for a data bus. Information on the 1553 standard is available on the internet from several sources, one of them is: 
http://www.1553-mil-std.com/

[7] Amp. Web site: http://www.amp.com/

[8] BiRa Systems. Web site: http://www.bira.com/

[9] Cosel USA Inc. Web site: http://www.coselusa.com/

[10] CTS Corporation. Web site: http://www.ctscorp.com/

[11] Harting. Web site: http://www.harting.com/

[12] Littlefuse. Web site: http://www.littlefuse.com/

[13] McMaster-Carr. Web site: http://www.mcmaster.com/

[14] Molex GC/Waldom. Web site: http://www.molex.com/

[15] Panasonic. 
Web site: http://www.panasonic.com/industrial/components/components.html

[16] Raychem. Web site: http://www.raychem.com/

[17] Penn-Engineering. Web site: http://www.penn-eng.com/

[18] Toko. Web site: http://www.toko.com/

[19] Trompeter Electronics. Web site: http://www.trompeter.com/

[20] Vicor. Web site: http://www.vicr.com/

# Footnotes

Derived from “High Speed Digital Design, A Handbook of Black Magic”. Appendix C, pg. 413 Resistance of a circuit trace.

1 From “Electrical Design Standards for Electronics to be used in Experiment Apparatus at Fermilab”, Rev. 4.0 10/16/98, pg. 9

Derived from “High Speed Digital Design, A Handbook of Black Magic”. Appendix C, pg. 414 Resistance of a power or ground plane

1 From “Electrical Design Standards for Electronics to be used in Experiment Apparatus at Fermilab”, Rev. 4.0 10/16/98, pg. 9