##### QADATA Summary – January 2004

The Inorganic Blind Sample Project (IBSP) reviews 108 inorganic analytical lines at the National Water Quality Laboratory (NWQL) and 79 inorganic analytical lines at the Ocala Water Quality and Research Laboratory (OWQRL) using double-blind, quality-assurance samples.  The QADATA Summary is a short summary of only those analytical lines for which the quality-assurance charts indicate or have indicated a bias and/or higher than expected variability.  Overall, the vast majority of the analyses at the NWQL and OWQRL demonstrate acceptable variability and little or no bias.

The data reviewed is for the time period 06/01/03 through 01/15/04; however, longer-range charts are shown when they add clarity to the problem.  The attached quality-assurance charts contain data from samples that have been subject to the same laboratory handling procedures as the samples that are submitted by the USGS district offices, and the data have the same level of quality control review as the data released to the districts.  Thus, any factors that affect the districts’ data should also affect the data shown in the IBSP quality-assurance charts.  Samples for which the results are outside of control limits 
(+/-2 f-pseudosigma) are reanalyzed, but the values are not updated.  Data from samples that are known to have deteriorated or are known to have been affected by a problem at IBSP are not displayed on the charts.

In this report, the “Problem” is determined and written by IBSP.  The “Problem” is then sent to the laboratories that then get two weeks to respond to the issue by writing a “Lab Comment” and a “Corrective Action.”  The labs’ comments are incorporated into the QADATA Summary as received.  Comments from previous QADATA Summaries are included in this report when they help provide a clearer perspective of the overall concern.

Current charts (updated twice weekly) may be accessed at http://bqs.usgs.gov/bsp/FY03_04charts.htm.  Laboratory and district personnel can also make their own QADATA retrieval to parallel their project’s sampling period by using the QADATA Application online at http://bqs.usgs.gov/bsp (select QADATA Application button).  Charts pertaining specifically to this report can be accessed electronically by following the underlined hyperlinks embedded within the report and associated with each analysis discussed.

Tedmund M. Struzeski

struzesk@usgs.gov

Table of Contents

Note:  All analytes are hyperlinked to their respective positions within the report.

NWQL

Analyte 	Page

Aluminum, Whole-Water Recoverable, ICP-MS, Labcode 2372, Parmcode 1105E . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	3

Arsenic, Filtered, GFAA, Labcode 2160, Parmcode 1000C . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 	3

Beryllium, Whole-Water Recoverable, ICP, Labcode 2353, Parmcode 1012C . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	4

Cadmium, Filtered , ICP-MS, Labcode 1788, Parmcode 1025G . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	4

Cadmium, Filtered , GFAA, Labcode 1554, Parmcode 1025F . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	4

Cadmium, Whole-Water Recoverable, GFAA, Labcode 1555, Parmcode 1027F . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	5

Chromium, Whole-Water Recoverable, GFAA, Labcode 1937, Parmcode 1034E . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 	6

Copper, Filtered , GFAA, Labcode 1558, Parmcode 1040F . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 	 6

Iron, Whole Water Recoverable, ICP, Labcode 2359, Parmcode 1045C . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  	7

Manganese, Filtered, ICP, Labcode 0648, Parmcode 1056C . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	7

Magnesium, Filtered, ICP, Labcode 0663, Parmcode 0925C . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	8

Sodium, Filtered, ICP, Labcode 0675, Parmcode 0930C . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	8

Potassium, Filtered, ICP, Labcode 2773, Parmcode 0935C . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 	8

OWQRL

Analyte

Ammonia, Filtered, Colorimetric, Labcode 1980, Parmcode 0608H . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	9

Phosphorus, Filtered, Colorimetric, Labcode 1983, Parmcode 0666D . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	9

Phosphorus, Whole-Water Recoverable, Colorimetric, Labcode 1984, Parmcode 0665D  . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 	10

Antimony, Whole-Water Recoverable, GFAA, Labcode 3122, Parmcode 1097(1)  . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 	11

Arsenic, Whole-Water Recoverable, GFAA, Labcode 3119, Parmcode 1002(1)  . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 	11

Cadmium, Filtered, GFAA, Labcode 1554, Parmcode 1025F . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 	12

Chromium, Filtered, GFAA, Labcode 1936, Parmcode 1030(1) . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	13

Selenium, Whole-Water Recoverable, GFAA, Labcode 3124, Parmcode 1147(0)  . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 	13

Alkalinity, Whole-Water Recoverable, Titrimetric, Labcode 0070, Parmcode 90410A  . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	14

Alkalinity, Filtered, Titrimetric, Labcode 3160, Parmcode 29801(0) . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 	14

Table 1

NWQL PERFORMANCE ON IBSP SAMPLES

Note:  Click the blue, underlined text in the left column to access the control chart for the associated parameter.

| Trace Metal Determinations | Problem | NWQL Lab Comment | Corrective Action |
| --- | --- | --- | --- |
| Aluminum, Whole-Water Recoverable ICP-MS labcode 2372 parmcode 1105E | Data indicate a large positive bias for part of November 2003. Although all biased data are from the same reference material, other methods for filtered aluminum (NWQL and OWQRL) indicate no apparent problems with this reference material. | The in-bottle digestion process (INBD) was experiencing an aluminum contamination during this time period. The HCl-acid dispenser broke and was replaced with a temporary dispenser while another one was ordered. The dispenser was rinsed several times and there was no evidence of contamination. Over a short time period, however, the dispenser started “freezing up” and was difficult to use. It was determined later that it was not adequate for use with acids and the contamination was due to this dispenser. Three INBD sets were affected by this contamination. In addition, small concentrations (above the LRL) of Boron, Copper, Strontium and Lead showed up in the set blanks, however, Aluminum showed the largest contamination. The large Aluminum positive bias for part of November 2003 was due to this contamination. | As soon as the source of contamination was determined, the dispenser was replaced with an appropriate dispenser and the acid bottle was disposed of. The contamination for all five elements has been eliminated. The INBD process is monitored closely for contamination by the use of INBD prep blanks. Whole-Water Recoverable Samples from these three sets that were affected by contamination from any of the five elements were given an “E” when a value was detected at less than 10X the contamination in the blank. |
| Arsenic, Filtered GFAA labcode 2160 parmcode 1000C | Data since October 2003 indicate a slight positive bias (7%). Since October there are 11 values (all results) greater than their median value. | New standards were made in November of 2003 including Continuing Calibration Verifications (CCV) and Third-Party Check standards (TPC). Both the CCV’s and TPC’s for arsenic are showing a positive bias on QC charts. | New calibration standards will be remade and verified as of February 6th, 2004. The element will continue to be monitored closely for a bias. |
| Beryllium, Whole-Water Recoverable ICP labcode 2353 parmcode 1012C | New ICP instrumentation came online around July 2003. Prior to July 2003, the data show no bias. Since July 2003, data demonstrate a slight negative bias (8%). | New standards were made in November 2003. Since then, QC charts for CCV’s and TPC’s show a slight positive bias. The six BQS samples were rerun and showed a slight positive bias. Before the standards were made in November, Beryllium was showing a slightly high bias and the standards were nearing their expiration date, thus, necessitating the need for new calibration standards. | New calibration standard will be remade and verified by February 2nd, 2004. The element will continue to be monitored closely for a bias. |
| Cadmium, Filtered ICP-MS labcode 1788 parmcode 1025G | Data for the last six months (June 2003 through December) demonstrate a slight positive bias (6%) with all 48 values greater than their median value. This bias is not evident for whole-water recoverable cadmium by ICP-MS (labcode 2376). | The bias is not evident on NWQL QC charts. | New calibration standards were made and verified for use at the end of December 2003. We will continue to monitor internal QC and IBSP charts for bias and variability. If the TPC standards indicate a bias, calibration standards will be remade and verified. |
| Cadmium, Filtered GFAA labcode 1554 parmcode 1025F | A positive bias is still evident; however, the overall bias is less than 5% for the time-frame October 2003 to mid-December. Analysis should be closely monitored until the next QADATA Summary is distributed. | The laboratory has and will continue to closely monitor this analysis. |  |
| Previous comments for cadmium, whole-water recoverable, GFAA, labcode 1554, parmcode 1025F |
|  | A positive bias is demonstrated (8.0%). This same pattern is shown for cadmium, whole-water recoverable by GFAA (labcode 1555). | The bias is not evident on NWQL QC charts. | New calibration standards were made and verified for use at the end of July. NIST-traceable third party check standards (TPC) were implemented in September. There is not enough data collected from the TPC standards for evaluation. We will continue to monitor internal QC and IBSP charts for bias and variability. If the TPC standards indicate a bias, calibration standards will be remade and verified. |
| Cadmium, Whole-Water Recoverable GFAA labcode 1555 parmcode 1027F | As of July 2003 the positive bias has improved (<2%); however, the data now exhibit high variability. Since September there are three out of 11 points (27%) outside of +/-6 f-pseudosigma. | For one of the samples, a blank was accidentally sent to the LIMS instead of a sample and a zero went in the database. There is no obvious explanation as to why the other two points are outside +/- 6 f-pseudosigma. The QC surrounding these samples was in control and the instrument was running within specifications. | To help alleviate human error when putting data into the LIMS, the GFAA section has implemented the use of an additional QC check sheet that is included in all data packets. This sheet lists all samples that have to be rerun due to QC failures, dilutions, possible mispours, and bottles that cannot be found. The analyst initials each sample when the action is performed in LIMS. This should help alleviate the problem of incorrect data going into LIMS. In addition, the section will closely monitor internal QC and IBSP charts for bias and variability. |
| Previous comments for cadmium, filtered, GFAA, labcode 1555, parmcode 1027F |
| From QADATA Summary November 2003 | A positive bias is still evident; however, the bias has decreased from 7.8 to 5.8%. This same pattern is shown for cadmium, filtered, by GFAA (labcode 1554). | The bias is not evident on NWQL QC charts. | New calibration standards were made and verified for use at the end of July. NIST-traceable third party check standards (TPC) were implemented in September. There is not enough data collected from the TPC standards for evaluation. We will continue to monitor internal QC and IBSP charts for bias and variability. If the TPC standards indicate a bias, calibration standards will be remade and verified. |
| From QADATA Summary September 2003 | The results for February 2003 through May 2003 demonstrate a positive bias (7.8%). The results for the same reference materials analyzed at the OWQRL do not exhibit a bias. | Began using new standards July 9, 2003. Online QC charts prior to July 9 show some bias, but bias has decreased since new standards were made. | Improvement is noted since standards were made. NWQL will continue to monitor data quality. |
| From QADATA Summary July 2003 | As of January 2003, a high bias has developed and variability increased. | NWQL QC charts show variable but acceptable data. | Instrument has recently been optimized, and the NWQL will continue to monitor. |
| Chromium, Whole-Water Recoverable GFAA labcode 1937 parmcode 1034E | The bias shown previously is not evident in the most recently received data. No corrective action necessary. |  |  |
| Previous comments for chromium, whole-water recoverable, GFAA, labcode 1937, parmcode 1034E |
| From QADATA Summary November 2003 | A negative bias (15.0%) has developed starting July 2003. | NWQL also identifies a chromium whole water negative bias. New intermediate stocks were prepared in July. Dissolved and whole water calibration standards were made from the fresh stocks in July. The QC charts for dissolved chromium do not show a bias. | New whole water recoverable calibration standards were made and verified October 20th. NIST-traceable third party check standards (TPC) were implemented in September. As of yet, we do not have enough data to evaluate bias and variability with the TPC. We continue to monitor internal QC and IBSP charts for bias and variability after these changes. |
| Copper, Filtered GFAA labcode 1558 parmcode 1040F | From June 2003 through late-October, the data exhibit a positive bias (11%). Data received since November show no bias. No corrective action necessary. |  |  |
| Iron, Whole Water Recoverable ICP labcode 2359 parmcode 1045C | The high variability previously demonstrated is not evident in data received as of September 2003. No corrective action necessary. |  |  |
| Previous comments for iron, whole-water recoverable, ICP, labcode 2359, parmcode 1045C |
| From QADATA Summary November 2003 | Since August 2003, the data indicate high variability. Three of 13 values (23%) are greater than 6 f-pseudosigmas from the median. | One of the three results outside of 6 f-pseudosigma from the MPV was due to a sample login mix-up. The sample has been relogged and is pending analysis. The other two outliers were due to carryover where the previous samples had Fe concentrations greater than 20,000 ug/L. | The ICP-OES instrumentation has the ability to monitor elemental concentration during the rinse. To prevent carryover in the future, the software will be set to monitor Fe concentrations greater than 250 ug/L and increase the rinse time if necessary. In addition to instrument/software corrections, analyst and reviewers will be instructed to watch for carryover when performing the secondary data review. The data packets will also be rechecked to look for any other samples the may have been affected by carryover and if necessary samples will be rerun and updated. |
| Manganese, Filtered ICP labcode 0648 parmcode 1056C | From mid-September 2003 through November, the data exhibit a positive bias (7%). Data received since December suggest there is no longer a bias. No corrective action necessary. |  |  |

| Major Determinations | Problem | NWQL Lab Comment | Corrective Action |
| --- | --- | --- | --- |
| Magnesium, Filtered ICP labcode 0663 parmcode 00925C | From mid-September through October 2003, the data exhibit a positive bias and high variability. Starting late-November, the data show little bias (<2%) and all data are within +/- 1.3 f-pseudosigmas of the median. No corrective action necessary. |  |  |
| Sodium, Filtered ICP labcode 0675 parmcode 00930C | The data display an overall positive bias (6%) from October 2003 through mid-December. Twenty-one of 22 results during this time period are greater than their median value. Higher than usual variability is demonstrated September through mid-December. Data received since mid-December show improved variability and no bias. Analysis should be closely monitored until the next QADATA Summary is distributed. | The laboratory has and will continue to closely monitor this analysis. |  |
| Potassium, Filtered ICP labcode 2773 parmcode 00935C | The data display an overall positive bias (7%) from October 2003 through early-December. All results during this time period (22 values) are greater than their median value. Data received since mid-December show no bias. Analysis should be closely monitored until the next QADATA Summary is distributed. | The laboratory has and will continue to closely monitor this analysis. |  |

###### Table 2

OWQRL PERFORMANCE ON IBSP SAMPLES

Note:  Click the blue, underlined text in the left column to access the control chart for the associated parameter.

| Nutrient Determinations | Problem | OWQRL Lab Comment | Corrective Action |
| --- | --- | --- | --- |
| Ammonia, Filtered Colorimetric Low-level labcode 1980 parmcode 0608H | Data are biased negative starting November 2003. Although all biased data are from the same reference material (N79), this material is not showing a bias for other ammonia methods at the OWQRL or the NWQL. Analysis should be closely monitored until the next QADATA Summary is distributed. | Our internal control charts for this parameter also showed a problem with this SRS during the same time period as the blind samples. | New reagents were made. New data shows that the bias no longer evident. This line will be monitored to be sure that the bias has been corrected. |
| Phosphorus, Filtered Colorimetric labcode 1983 parmcode 0666D | The negative bias has improved. Since October 2003, the overall negative bias is <2%. There is no bias indicated after December 2003. No corrective action necessary. |  |  |
| Previous comments for phosphorus, filtered, labcode 1983, parmcode 0666D |
| From QADATA Summary November 2003 | From May through October 2003, a negative bias is demonstrated (9.8%). The same pattern is shown for phosphorus, whole-water recoverable. The NWQL is not showing this bias. | The bias for this parameter and the total phosphorus, which is run at the same time, seems to be worse for low level samples than for high level. The average bias for all blind samples above 0.1 mg/L is -3.1% for total and dissolved phosphorus. The bias in the lower part of the calibration is also evident in our internal control charts. | New stock solutions were made on 10/2/2003. All tubing has been changed and all annual maintenance was performed ahead of schedule. Preliminary results have shown that this might have helped eliminate the bias. We will also be checking calibration procedures. |
| From QADATA Summary September 2003 | All data received since May 1, 2003, indicate a negative bias. The NWQL is not showing this bias. | Most of the data points exhibiting the negative bias are a commercially obtained sample that has also exhibited a negative bias when run at the NWQL. Other blind samples are also showing a negative bias, though not as much as the commercial one. Whole-water recoverable phosphorus, which is analyzed on the same instrument at the same time, is not showing this same bias. The same commercially obtained sample is not submitted for Total Phosphorus. | The cause of the bias is currently being investigated. This line is being monitored closely. |
| Phosphorus, Whole-Water Recoverable Colorimetric labcode 1984 parmcode 0665D | The negative bias has improved. Since October 2003, the overall negative bias is <4%. There is no bias indicated after December 2003. No corrective action necessary. |  |  |
| Previous comments for phosphorus, whole-water recoverable, labcode 1984, parmcode 0665D |
| From QADATA Summary November 2003 | From June through October 2003, a negative bias is demonstrated (10.8%). The same pattern is shown for phosphorus, filtered. The NWQL is not showing this bias. | The bias for this parameter and the total phosphorus, which is run at the same time, seems to be worse for low level samples than for high level. The average bias for all blind samples above 0.1 mg/L is -3.1% for total and dissolved phosphorus. The bias in the lower part of the calibration is also evident in our internal control charts. | New stock solutions were made on 10/2/2003. All tubing has been changed and all annual maintenance was performed ahead of schedule. Preliminary results have shown that this might have helped eliminate the bias. We will also be checking calibration procedures. |

| Trace Metal Determinations | Problem | OWQRL Lab Comment | Corrective Action |
| --- | --- | --- | --- |
| Antimony, Whole-Water Recoverable GFAA labcode 3122 parmcode 1097(1) | From September through October 2003, a negative bias is demonstrated. All data since November are within one f-pseudosigma of the median. No corrective action necessary. |  |  |
| Arsenic, Whole-Water Recoverable GFAA labcode 3119 parmcode 1002(1) | The data still indicate high variability. Two of the last seven values received (29%) are outside of +/-2 f-pseudosigma. | Service agreement discontinued in April has been renewed. | The GFAA was recently serviced by the manufacturer and several problems corrected. It will be monitored closely to be sure that problems with the instrument do not recur. |
| Previous comments for arsenic, whole-water recoverable, labcode 3119, parmcode 1002(1) |
| From QADATA Summary Nov 2003 | High variability is still demonstrated. Three of the last seven values received (43%) are outside of +/-2 f-pseudosigma. | Due to factors beyond our control and the uncertainty of this facility’s continued existence, certain cost reduction policies were initiated in April, 2003. This may have led to undetected instrument deficiencies over time in the trace metals section. | The cost reduction policies have been discontinued for FY04. Samples will be retested to see if any biases are detected. |
| From QADATA Summary Sept 2003 | The most recently received data indicate on-going problems with variability and a positive bias. Since October 2002, five out of 19 points (26%) are outside of +/-2 f-pseudosigma with 3 of 19 (16%) outside of +/-3 f- pseudosigma. | Problems with low energy from the Arsenic lamp have led to increased variability and decreased sensitivity. | All of our lamp power supplies for the graphite furnace have been evaluated and those that need to be replaced have been identified. The instrument is also currently being run in the single rather than multi-element mode to help increase sensitivity and decrease the variability. |
| From QADATA Summary July 2003 | The two new values received since the last QADATA Summary are both within one f-pseudosigma of the median. Analysis should be closely monitored until the next QADATA Summary is distributed. | The problems with the GFAA at low wavelengths have been corrected. | We will continue to monitor the GFAA closely to be sure that the problems do not reappear. |
| From QADATA Summary May 2003 | The variability is beyond acceptable levels. IBSP recommends that the source of the problem be determined before any more analyses of environmental samples are performed. Since October 2002, four out of 13 points (31%) are outside of +/-2 f-pseudosigma with 3 of 13 (23%) outside of +/-3 f-pseudosigma. | We will be running a new Demonstration of Capability for Arsenic before any more samples are analyzed. | More service work was done to correct problems with the GFAA. Past samples have been checked by other techniques (ICP, ICP/MS) to make sure that values reported are correct. |
| From QADATA Summary March 2003 | High variability is still exhibited. Since late October 2002, three out of 9 points (33%) are outside of +/-2 f-pseudosigma. | A problem with the beam combiner on the SIMAA graphite furnace may have led to some erratic results. | Service was performed on the AA to correct the problem. Samples are being rerun to see if any data needs to be corrected. |
| From QADATA Summary January 2003 | Increased variability starting late October 2002. | Problem with matrix modifier. | Amount of matrix modifier used was adjusted to give more consistent results. |
| Cadmium, Filtered GFAA labcode 1554 parmcode 1025(F) | High variability demonstrated during October 2003. The three most recent values are within one f-pseudosigma of the median. No corrective action necessary. |  |  |
| Chromium, Filtered GFAA labcode 1936 parmcode 1030(I) | High variability demonstrated throughout September and October 2003. Although all these data are from the same reference material, this material is not showing a problem for whole-water recoverable chromium (labcode 3019) or for chromium at the NWQL. The three most recent values are within one f-pseudosigma of the median. Analysis should be closely monitored until the next QADATA Summary is distributed. | Service agreement discontinued in April has been renewed. | The GFAA was recently serviced by the manufacturer and several problems corrected. It will be monitored closely to be sure that problems with the instrument do not recur. |
| Previous comments for chromium, filtered, labcode 1936, parmcode 1030(1) |
| From QADATA Summary Nov 2003 | The three most recent values are all outside +/-2 f-pseudosigma indicating high variability. These values are all from the same reference material which is used for chromium, whole-water recoverable (OWQRL and NWQL) and chromium, filtered (NWQL) with no apparent problems. | Due to factors beyond our control and the uncertainty of this facility’s continued existence, certain cost reduction policies were initiated in April, 2003. This may have led to undetected instrument deficiencies over time in the trace metals section. | The cost reduction policies have been discontinued for FY04. Samples will be retested to see if any biases are detected. |
| Selenium, Whole-Water Recoverable GFAA labcode 3124 parmcode 1147(0) | The variability has improved. Although data are now biased slightly negative, all data received since mid-September 2003 are within one f-pseudosigma of the median. Analysis should be closely monitored until the next QADATA Summary is distributed. | Service agreement discontinued in April has been renewed. | The GFAA was recently serviced by the manufacturer and several problems corrected. It will be monitored closely to be sure that problems with the instrument do not recur. |
| Previous comments for selenium, whole-water recoverable, labcode 3124, parmcode 1147(0) |
| From QADATA Summary Nov 2003 | The variability has recently increased. Two of the last six values received (33%) are outside of +/-2 f-pseudosigma. | Due to factors beyond our control and the uncertainty of this facility’s continued existence, certain cost reduction policies were initiated in April, 2003. This may have led to undetected instrument deficiencies over time in the trace metals section. | The cost reduction policies have been discontinued for FY04. Samples will be retested to see if any biases are detected. |

| Major Determinations | Problem | OWQRL Lab Comment | Corrective Action |
| --- | --- | --- | --- |
| Alkalinity, Whole-Water Recoverable Titrimetric labcode 0070 parmcode 90410A | Corrective action taken by the OWQRL appears to have fixed previous data anomalies. Data since August 2003 show no bias and acceptable variability. No corrective action necessary. |  |  |
| Previous comments for alkalinity, filtered, labcode 0070, parmcode 90410(A) |
| From QADATA Summary Nov 2003 | Lower-level reference materials exhibit a large positive bias (up to 38%) with the bias decreasing as the alkalinity increases. The bias started around the beginning of July 2003. Past laboratory comments suggest that perhaps the autotitrator speed is set too fast for low-level samples. | A new analyst was being trained in July. The protocol for setting the titrator speed was not explained properly at first, leading to the bias shown on the chart for July for the low level samples. | After the bias was detected, the analyst was retrained. No low level blind samples for alkalinity have been received since the end of July. |
| Alkalinity, Filtered Titrimetric labcode 3160 parmcode 29801(0) | Corrective action taken by the OWQRL appears to have fixed previous data anomalies. Data since August 2003 show little bias and acceptable variability. No corrective action necessary. |  |  |
| Previous comments for alkalinity, whole-water recoverable, labcode 3160, parmcode 29801(0) |
| From QADATA Summary Nov 2003 | Lower-level reference materials exhibit a positive bias with the bias decreasing as the alkalinity increases. The bias started around the beginning of July 2003. Past laboratory comments suggest that perhaps the autotitrator speed is set too fast for low-level samples. | A new analyst was being trained in July. The protocol for setting the titrator speed was not explained properly at first, leading to the bias shown on the chart for July for the low level samples. | After the bias was detected, the analyst was retrained. No low level blind samples for alkalinity have been received since the end of July. |