##### QADATA Summary – November 2003

The Inorganic Blind Sample Project (IBSP) reviews 108 inorganic analytical lines at the National Water Quality Laboratory (NWQL) and 79 inorganic analytical lines at the Ocala Water Quality and Research Laboratory (OWQRL) using double-blind, quality-assurance samples.  The QADATA Summary is a short summary of only those analytical lines for which the quality-assurance charts indicate or have indicated a bias and/or higher than expected variability.  Overall, the vast majority of the analyses at the NWQL and OWQRL demonstrate acceptable variability and little or no bias.

The data reviewed is for the time period 04/01/03 through 10/28/03; however, longer-range charts are shown when they add clarity to the problem.  The attached quality-assurance charts contain data from samples that have been subject to the same laboratory handling procedures as the samples that are submitted by the USGS district offices, and the data have the same level of quality control-review as the data released to the districts.  Thus, any factors that affect the districts’ data should also affect the data shown in the IBSP quality-assurance charts.  Samples for which the results are outside of control limits 
(+/-2 f-pseudosigma) are reanalyzed, but the values are not updated.  Data from samples that are known to have deteriorated or are known to have been affected by a problem at IBSP are not displayed on the charts.

In this report, the “Problem” is determined and written by IBSP.  The “Problem” is then sent to the laboratories that then get two weeks to respond to the issue by writing a “Lab Comment” and a “Corrective Action.”  The labs’ comments are incorporated into the QADATA Summary as received.  Comments from previous QADATA Summaries are included in this report when they help provide a clearer perspective of the overall concern.

Charts containing all data received as of 11/17/03 may be accessed at http://bqs.usgs.gov/bsp/FY03_04charts.htm.  Laboratory and district personnel can also make their own QADATA retrieval to parallel their project’s sampling period by using the QADATA Application online at http://bqs.usgs.gov/bsp (select QADATA Application button).  Charts pertaining specifically to this report can be accessed electronically by following the blue, underlined hyperlinks embedded within the report and associated with each analysis discussed.

Tedmund M. Struzeski

struzesk@usgs.gov

Table of Contents

Note:  All analytes are hyperlinked to their respective positions within the report.

NWQL

Analyte 	Page

Orthophosphate, Filtered, Colorimetric, Labcode 1974,  Parmcode 0671H . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	3

Barium, Whole-Water Recoverable, ICP, Labcode 2352, Parmcode 1007C . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	4

Cadmium, Whole-Water Recoverable, GFAA, Labcode 1555, Parmcode 1027F . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 	5

Cadmium, Filtered, GFAA, Labcode 1554, Parmcode 1025F . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 	5

Chromium, Whole-Water Recoverable, GFAA, Labcode 1937, Parmcode 1034E . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 	6

Iron, Whole Water Recoverable, ICP, Labcode 2359, Parmcode 1045C . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 	6

OWQRL

Analyte

Phosphorus, Filtered, Colorimetric, Labcode 1983, Parmcode 0666D . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	7

Phosphorus, Whole-Water Recoverable, Colorimetric, Labcode 1984, Parmcode 0665D  . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 	7

Phosphorus, Whole-Water Recoverable, EPA, Colorimetric, Labcode 1993, Parmcode 0665E  . . . . . . . . . . . . . . . . . . . .. . . . . . . . . . . . . . . 	7

Arsenic, Whole-Water Recoverable, GFAA, Labcode 3119, Parmcode 1002(1)  . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	8

Chromium, Filtered, GFAA, Labcode 1936, Parmcode 1030(1) . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	9

Selenium, Whole-Water Recoverable, GFAA, Labcode 3124, Parmcode 1147(0)  . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	9

Alkalinity, Whole-Water Recoverable, Titrimetric, Labcode 0070, Parmcode 90410A  . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	9

Alkalinity, Filtered, Titrimetric, Labcode 3160, Parmcode 29801(0) . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	9

Table 1

NWQL PERFORMANCE ON IBSP SAMPLES

Note:  Click the blue, underlined text in the left column to access the control chart for the associated parameter.

| Nutrient Determinations | Problem | NWQL Lab Comment | Corrective Action |
| --- | --- | --- | --- |
| Orthophosphate, Filtered Colorimetric labcode 1974 parmcode 0671H | NS10 is still showing a negative bias (-7.9%). The OWQRL shows a -2.6 % bias for this reference material. The overall negative bias demonstrated by this analytical line is minimal (-2.2%) and is due exclusively to data from NS10. | NS10 still exhibits a negative bias that is not representative of other QC samples with a similar concentration. A stronger cleaning protocol for the analytical line instituted in August at first seemed to improve the performance for NS10, but later samples still exhibit the negative bias observed for some time. | The NWQL will continue to seek causes for the negative bias for NS10. We will continue to monitor performance for NS10 and other certified reference samples at a similar concentration. |
| Previous comments for orthophosphate, filtered, labcode 1974, parmcode 0671H |
| From QADATA Summary September 2003 | Results for NS10 are still demonstrating a negative bias with about one-half of the NS10 data outside of -2 f-pseudosigma. The OWQRL is not demonstrating a negative bias of this magnitude and has no data outside of -2 f-pseudosigma. | A new vial of the NIST certified reference sample purchased by the NWQL was opened and a sample at a concentration similar to NS10 was prepared and analyzed. Values were slightly higher than the calculated mean value. Samples made at three concentrations from a new commercially prepared certified reference material were also analyzed. The values for these samples were close to but slightly lower than the calculated mean. | The Nutrients Unit will continue to monitor IBSP control charts and will analyze the new QC samples as well as on-line NIST QC samples. We will continue to look for other causes for the low bias for NS10. |

| Trace Metal Determinations | Problem | NWQL Lab Comment | Corrective Action |
| --- | --- | --- | --- |
| Barium, Whole-Water Recoverable ICP labcode 2352 parmcode 1007C | There are not enough data from the new ICPs (four values at this time) to make a bias assessment. Analysis should be closely monitored until the next QADATA Summary is distributed. | The NWQL will continue to monitor WWR barium QC on the new instruments to determine if a bias continues. | No corrective action needed at this time, however internal QC including NIST-traceable third party check standards will be used for monitoring bias and variability. |
| Previous comments for barium, whole-water recoverable, ICP, labcode 2352, parmcode 1007C |
| From QADATA Summary September 2003 | The negative bias (-5.2% for WY03) has still not improved. | New ICP instrumentation went online beginning in August. There is no IBSP data provided from this time to date. Online QC data for these new instruments indicate no appreciable bias at these concentrations. A Rapi-Note and an NWQL Technical Memorandum will be issued in the near future detailing the use of the new ICPs for Whole-Water Recoverable analyses. | The NWQL will continue to monitor WWR barium QC on the new instruments to determine if a bias continues. |
| From QADATA Summary July 2003 | The historical negative bias has not improved and may be getting worse. It is again recommended that the samples showing the biased results be reanalyzed using one of the new ICPs and then the data compared. | The NWQL agrees with IBSP that the barium bias is getting worse. New ICP instrumentation has come online for the dissolved line. | The NWQL will continue to monitor the barium WWR line on the different instrument to see if there is an improvement. The ICP WWR method has been moved to a different instrument that has had recent maintenance. We will reanalyze IBSP samples on this instrumentation. We plan to have the new Perkin Elmer ICP instrumentation online this summer. |
| From QADATA Summary May 2003 | The historical negative bias persists. It is recommended that the samples showing the biased results be reanalyzed using one of the new ICPs and then the data compared. | Online QC charts show that the negative bias increased significantly during mid-March. Several troubleshooting measures were checked on the instrument including: plasma power, pump efficiency, cleaning of nebulizer and spray chamber, dark current test, all of which showed no indication of a problem. The only test that showed a significant problem was that the Purged Optical Path, which needed adjusting. | The NWQL will reanalyze samples on the new ICPs to compare the data, once we have assessed the WWR prove-out data for the new instruments. We will continue to monitor online QC charts to see if the negative bias has improved as a result of the optical path adjustment. |
| Cadmium, Whole-Water Recoverable GFAA labcode 1555 parmcode 1027F | A positive bias is still evident, however, the bias has decreased from 7.8 to 5.8%. This same pattern is shown for cadmium, filtered, by GFAA. | The bias is not evident on NWQL QC charts. | New calibration standards were made and verified for use at the end of July. NIST-traceable third party check standards (TPC) were implemented in September. There is not enough data collected from the TPC standards for evaluation. We will continue to monitor internal QC and IBSP charts for bias and variability. If the TPC standards indicate a bias, calibration standards will be remade and verified. |
| Previous comments for cadmium, whole-water recoverable, GFAA, labcode 1555, parmcode 1027F |
| From QADATA Summary September 2003 | The results for February 2003 through May 2003 demonstrate a positive bias (7.8%). The results for the same reference materials analyzed at the OWQRL do not exhibit a bias. | Began using new standards July 9, 2003. Online QC charts prior to July 9 show some bias, but bias has decreased since new standards were made. | Improvement is noted since standards were made. NWQL will continue to monitor data quality. |
| From QADATA Summary July 2003 | As of January 2003, a high bias has developed and variability increased. | NWQL QC charts show variable but acceptable data. | Instrument has recently been optimized, and the NWQL will continue to monitor. |
| Cadmium, Filtered GFAA labcode 1554 parmcode 1025F | A positive bias is demonstrated (8.0%). This same pattern is shown for cadmium, whole-water recoverable by GFAA. | The bias is not evident on NWQL QC charts. | New calibration standards were made and verified for use at the end of July. NIST-traceable third party check standards (TPC) were implemented in September. There is not enough data collected from the TPC standards for evaluation. We will continue to monitor internal QC and IBSP charts for bias and variability. If the TPC standards indicate a bias, calibration standards will be remade and verified. |
| Chromium, Whole-Water Recoverable GFAA labcode 1937 parmcode 1034E | A negative bias (-15.0%) has developed starting July 2003. | NWQL also identifies a chromium whole water negative bias. New intermediate stocks were prepared in July. Dissolved and whole water calibration standards were made from the fresh stocks in July. The QC charts for dissolved chromium do not show a bias. | New whole water recoverable calibration standards were made and verified October 20th. NIST-traceable third party check standards (TPC) were implemented in September. As of yet, we do not have enough data to evaluate bias and variability with the TPC. We continue to monitor internal QC and IBSP charts for bias and variability after these changes. |
| Iron, Whole Water Recoverable ICP labcode 2359 parmcode 1045C | Since August 2003, the data indicate high variability. Three of 13 values (23%) are greater than 6 f-pseudosigmas from the median. | One of the three results outside of 6 f-pseudosigma from the MPV was due to a sample login mixup. The sample has been relogged and is pending analysis. The other two outliers were due to carryover where the previous samples had Fe concentrations greater than 20,000 ug/L. | The ICP-OES instrumentation has the ability to monitor elemental concentration during the rinse. To prevent carryover in the future, the software will be set to monitor Fe concentrations greater than 250 ug/L and increase the rinse time if necessary. In addition to instrument/software corrections, analyst and reviewers will be instructed to watch for carryover when performing the secondary data review. The data packets will also be rechecked to look for any other samples the may have been affected by carryover and if necessary samples will be rerun and updated. |

###### Table 2

OWQRL PERFORMANCE ON IBSP SAMPLES

Note:  Click the blue, underlined text in the left column to access the control chart for the associated parameter.

| Nutrient Determinations | Problem | OWQRL Lab Comment | Corrective Action |
| --- | --- | --- | --- |
| Phosphorus, Filtered Colorimetric labcode 1983 parmcode 0666D | From May through October 2003, a negative bias is demonstrated (-9.8%). The same pattern is shown for phosphorus, whole-water recoverable. The NWQL is not showing this bias. | The bias for this parameter and the total phosphorus, which is run at the same time, seems to be worse for low level samples than for high level. The average bias for all blind samples above 0.1 mg/L is -3.1% for total and dissolved phosphorus. The bias in the lower part of the calibration is also evident in our internal control charts. | New stock solutions were made on 10/2/2003. All tubing has been changed and all annual maintenance was performed ahead of schedule. Preliminary results have shown that this might have helped eliminate the bias. We will also be checking calibration procedures. |
| Previous comments for phosphorus, filtered, labcode 1983, parmcode 0666D |
| From QADATA Summary September 2003 | All data received since May 1, 2003, indicate a negative bias. The NWQL is not showing this bias. | Most of the data points exhibiting the negative bias are a commercially obtained sample that has also exhibited a negative bias when run at the NWQL. Other blind samples are also showing a negative bias, though not as much as the commercial one. Whole-water recoverable phosphorus, which is analyzed on the same instrument at the same time, is not showing this same bias. The same commercially obtained sample is not submitted for Total Phosphorus. | The cause of the bias is currently being investigated. This line is being monitored closely. |
| Phosphorus, Whole-Water Recoverable Colorimetric labcode 1984 parmcode 0665D | From June through October 2003, a negative bias is demonstrated (-10.8%). The same pattern is shown for phosphorus, filtered. The NWQL is not showing this bias. | See comment for Phosphorus, Filtered (labcode 1983, parmcode 0666D). | See comment for Phosphorus, Filtered (labcode 1983, parmcode 0666D). |
| Phosphorus, Whole-Water Recoverable, EPA Colorimetric labcode 1993 parmcode 0665E | From May through October 2003, a negative bias is demonstrated (-10.2%). The same pattern is shown for phosphorus, filtered. The NWQL is not showing this bias. | See comment for Phosphorus, Filtered (labcode 1983, parmcode 0666D). | See comment for Phosphorus, Filtered (labcode 1983, parmcode 0666D). |

| Trace Metal Determinations | Problem | OWQRL Lab Comment | Corrective Action |
| --- | --- | --- | --- |
| Arsenic, Whole-Water Recoverable GFAA labcode 3119 parmcode 1002(1) | High variability is still demonstrated. Three of the last seven values received (43%) are outside of +/-2 SD. | Due to factors beyond our control and the uncertainty of this facility’s continued existence, certain cost reduction policies were initiated in April, 2003. This may have led to undetected instrument deficiencies over time in the trace metals section. | The cost reduction policies have been discontinued for FY04. Samples will be retested to see if any biases are detected. |
| Previous comments for arsenic, whole-water recoverable, labcode 3119, parmcode 1002(1) |
| From QADATA Summary Sept 2003 | The most recently received data indicate on-going problems with variability and a positive bias. Since October 2002, five out of 19 points (26%) are outside of +/-2 SD with 3 of 19 (16%) outside of +/-3 SD. | Problems with low energy from the Arsenic lamp have led to increased variability and decreased sensitivity. | All of our lamp power supplies for the graphite furnace have been evaluated and those that need to be replaced have been identified. The instrument is also currently being run in the single rather than multi-element mode to help increase sensitivity and decrease the variability. |
| From QADATA Summary July 2003 | The two new values received since the last QADATA Summary are both within one f-pseudosigma of the median. Analysis should be closely monitored until the next QADATA Summary is distributed. | The problems with the GFAA at low wavelengths have been corrected. | We will continue to monitor the GFAA closely to be sure that the problems do not reappear. |
| From QADATA Summary May 2003 | The variability is beyond acceptable levels. IBSP recommends that the source of the problem be determined before any more analyses of environmental samples are performed. Since October 2002, four out of 13 points (31%) are outside of +/-2 SD with 3 of 13 (23%) outside of +/-3 SD. | We will be running a new Demonstration of Capability for Arsenic before any more samples are analyzed. | More service work was done to correct problems with the GFAA. Past samples have been checked by other techniques (ICP, ICP/MS) to make sure that values reported are correct. |
| From QADATA Summary March 2003 | High variability is still exhibited. Since late October 2002, three out of 9 points (33%) are outside of +/-2 SD. | A problem with the beam combiner on the SIMAA graphite furnace may have led to some erratic results. | Service was performed on the AA to correct the problem. Samples are being rerun to see if any data needs to be corrected. |
| From QADATA Summary January 2003 | Increased variability starting late October 2002. | Problem with matrix modifier. | Amount of matrix modifier used was adjusted to give more consistent results. |
| Chromium, Filtered GFAA labcode 1936 parmcode 1030(I) | The three most recent values are all outside +/-2 SD indicating high variability. These values are all from the same reference material which is used for chromium, whole-water recoverable (OWQRL and NWQL) and chromium, filtered (NWQL) with no apparent problems. | See comment for Arsenic, WWR (labcode 3119, parmcode 1002[1]). | See comment for Arsenic, WWR (labcode 3119, parmcode 1002[1]). |
| Selenium, Whole-Water Recoverable GFAA labcode 3124 parmcode 1147(0) | The variability has recently increased. Two of the last six values received (33%) are outside of +/-2 SD. | See comment for Arsenic, WWR (labcode 3119, parmcode 1002[1]). | See comment for Arsenic, WWR (labcode 3119, parmcode 1002[1]). |

| Major Determinations | Problem | OWQRL Lab Comment | Corrective Action |
| --- | --- | --- | --- |
| Alkalinity, Whole-Water Recoverable Titrimetric labcode 0070 parmcode 90410A | Lower-level reference materials exhibit a large positive bias (up to 38%) with the bias decreasing as the alkalinity increases. The bias started around the beginning of July 2003. Past laboratory comments suggest that perhaps the autotitrator speed is set too fast for low-level samples. | A new analyst was being trained in July. The protocol for setting the titrator speed was not explained properly at first, leading to the bias shown on the chart for July for the low level samples. | After the bias was detected, the analyst was retrained. No low level blind samples for alkalinity have been received since the end of July. |
| Alkalinity, Filtered Titrimetric labcode 3160 parmcode 29801(0) | Lower-level reference materials exhibit a positive bias with the bias decreasing as the alkalinity increases. The bias started around the beginning of July 2003. Past laboratory comments suggest that perhaps the autotitrator speed is set too fast for low-level samples. | See above comment for Alkalinity, WWR (labcode 0070, parmcode 90410A). | See above comment for Alkalinity, WWR (labcode 0070, parmcode 90410A). |